Keck Adaptive Optics Note 714

Keck Next Generation Adaptive Optics:

AO Control System Design Document

Doug Morrison, Kevin Tsubota, Jimmy Johnson

May 4, 2010

Revision History

| Revision | Date | Author(s) | Remarks |
| --- | --- | --- | --- |
| 1.0 |  | Douglas Morrison Kevin Tsubota | Initial release. |

Table of Contents

1	Introduction	7

2	Glossary	7

3	Referenced Documents	8

4	Document Overview	8

4.1	Subsystem Template	8

4.1.1	Description	9

4.1.2	Interfaces	9

4.1.3	Control Loops	10

4.1.4	Devices	10

5	Controls Infrastructure	10

5.1	Framework	11

5.2	Common Services	13

5.2.1	Connection Service	13

5.2.2	Event Service	14

5.2.3	Health Monitoring	15

5.2.4	Logging	15

5.2.5	Archiving	15

5.2.6	Alarming	16

5.2.7	Configuration	16

5.3	Component Architecture	16

5.3.1	Controllers	17

5.3.2	Composites	20

5.3.3	Device Control Interface	20

5.4	Low Level Motion Device Control	23

5.4.1	Introduction	23

5.4.2	Discrete Commanded Devices	24

5.4.3	Non-tracking Motion Devices	25

5.4.4	Tracking Motion Devices	26

5.5	Low Level Non-Motion Device Control	27

5.5.1	Introduction	27

5.5.2	Remote Power Control	28

5.5.3	Camera / Detector Control	29

5.5.4	Video	29

5.5.5	RS-232 and USB	29

6	AO System	30

6.1	Introduction	30

6.2	Wide Field Relay	30

6.2.1	Description	30

6.2.2	Control Loops	31

6.2.3	Devices	32

6.3	Acquisition System	33

6.3.1	Description	33

6.3.2	Interfaces	33

6.3.3	Control Loops	34

6.3.4	Devices	35

6.4	LGS Wavefront Sensors	35

6.4.1	Description	35

6.4.2	Interfaces	36

6.4.3	Control Loops	38

6.4.4	Devices	41

6.5	Low Order Wavefront Sensors	41

6.5.1	Description	42

6.5.2	Interfaces	42

6.5.3	Control Loops	44

6.5.4	Devices	45

6.6	Narrow Field Relay	45

6.6.1	Description	45

6.6.2	Control Loops	46

6.6.3	Devices	47

6.7	NGS Wavefront Sensor	47

6.7.1	Description	47

6.7.2	Interfaces	48

6.7.3	Control Loops	50

6.7.4	Devices	50

6.8	Atmospheric Dispersion Corrector	50

6.8.1	Description	51

6.8.2	Interfaces	51

6.8.3	Control Loops	52

6.8.4	Devices	53

6.9	RTC Interface	53

6.9.1	Description	53

6.9.2	Interfaces	54

6.9.3	Control Loops	56

6.9.4	Devices	57

6.10	Offloading	57

6.10.1	Description	57

6.10.2	Interfaces	58

6.10.3	Control Loops	59

6.10.4	Devices	61

6.11	Calibration Simulation	61

6.11.1	Description	61

6.11.2	Interfaces	62

6.11.3	Devices	63

6.12	AO Enclosure	63

6.12.1	Description	63

6.12.2	Interfaces	64

6.12.3	Control Loops	64

6.12.4	Devices	65

6.13	AO Control Sequencer	66

6.13.1	Description	66

6.13.2	Interfaces	69

6.13.3	Control Loops	74

6.13.4	Devices	74

7	Laser System	75

7.1	Introduction	75

7.2	Laser Switchyard	75

7.2.1	Description	75

7.2.2	Interfaces	76

7.2.3	Control Loops	78

7.2.4	Devices	79

7.3	Laser Beam Generation System	80

7.3.1	Description	80

7.3.2	Interfaces	81

7.3.3	Control Loops	82

7.3.4	Devices	84

7.4	Laser Safety System Interface	85

7.4.1	Description	85

7.4.2	Interfaces	86

7.4.3	Control Loops	88

7.4.4	Devices	88

7.5	Laser Service Enclosure	88

7.5.1	Description	88

7.5.2	Interfaces	89

7.5.3	Control Loops	89

7.5.4	Devices	89

7.6	Laser Interface	89

7.6.1	Description	90

7.6.2	Interfaces	90

7.6.3	Control Loops	91

7.6.4	Devices	92

7.7	Laser Control Sequencer	92

7.7.1	Description	92

7.7.2	Interfaces	95

7.7.3	Control Loops	97

7.7.4	Devices	98

8	Data Server	98

9	Risk Analysis	99

10	Issues Requiring Resolution During DD Phase	99

A1.	Functional Requirements	100

List of Figures

Figure 1: A Block diagram of the NGAO system viewed abstractly as a distributed controls system	10

Figure 2: A block diagram of the NGAO controls infrastructure	11

Figure 3: Middleware communication with multiple devices.	12

Figure 4: Generic KCSF Device Interface Layer	28

Figure 5: AO Control Block Diagram	30

Figure 6: Wide Field Relay Collaboration Diagram	31

Figure 7: Image Rotator Control Loop	32

Figure 8: Field Rotation Offset Control Loop	32

Figure 9: Acquisition Subsystem Collaboration Diagram	33

Figure 10: LGS WFS Collaboration Diagram	36

Figure 11: WFS Focus Control Loop	39

Figure 12: Tomography LGS WFS Offloading	40

Figure 13: Patrolling LGS WFS Offloading	40

Figure 14: Low-Order WFS Collaboration Diagram	42

Figure 15: Narrow Field Relay Collaboration Diagram	46

Figure 16: ADC Tip-Tilt Error Correction Control Loop	46

Figure 17: ADC Tip-Tilt Error Correction Control Loop	47

Figure 18: NGS WFS Subsystem Collaboration Diagram	48

Figure 19: ADC Collaboration Diagram	51

Figure 20: ADC Tracking Control Loop	53

Figure 21: RTC Interface Block Diagram	54

Figure 22: External Offloading Collaboration Diagram	58

Figure 23: Telescope Offload Control Loop	60

Figure 24: DAR Offset Control Loop	60

Figure 25: Calibration Subsystem Collaboration Diagram	61

Figure 26: AO Enclosure Collaboration Diagram	64

Figure 27: AO Control Sequencer Collaboration Diagram	67

Figure 28: AO Sequencer State Diagram	68

Figure 29 Laser Control Block Diagram	75

Figure 30: Laser Switchyard Collaboration Diagram	76

Figure 31: Laser Steering Control Loop	78

Figure 32: Laser Polarization Control Loop	79

Figure 33: Beam Generation System Collaboration Diagram	80

Figure 34: Laser Focus Tracking Control Loop	83

Figure 35: Asterism Rotator Tracking Control Loop	83

Figure 36: Flexure Compensation Control Loop	83

Figure 37: Laser Safety System Collaboration Diagram	86

Figure 38: Laser Service Enclosure Collaboration Diagram	88

Figure 39: Laser Interface Collaboration Diagram	90

Figure 40: Laser Control Sequencer Diagram	92

Figure 41: Laser Sequencer State Diagram	94

Figure 42: Asterism Validation Control Loop	97

# Introduction

This document describes the system design for the Non-Real-Time Controls portion of the NGAO system for the Preliminary Design Phase.

We begin by discussing the layout and format of the document, and cover some of the common terms that will be used throughout. An overview of the distributed controls architecture is discussed, which includes development of control components and use of common services. Section 5 covers the software design for the motion control hardware system, followed by the software controls architecture design in sections 6 and 7. The design document closes with a summary of existing risks and detailed design phase goals. An abbreviated form of the identified software functional requirements is listed in the appendix.

# Glossary

The following terms and acronyms will be used throughout this document:

| Acronym | Definition |
| --- | --- |
| AO | Adaptive Optics |
| API | Application Programming Interface |
| ATST | Advanced Technology Solar Telescope |
| BGS | Beam Generation System |
| BTO | Beam Train Optics |
| CBD | Component Based Development |
| CCM | Container Component Model |
| DAR | Differential Atmospheric Refraction |
| DCS | Drive Control System |
| DDS | Data Distribution Service Middleware |
| DM | Deformable Mirror |
| FSM | Field Steering Mirrors |
| ICE | Internet Communication Engine Middleware |
| IF | Interferometer |
| IFS | Integral Field Spectrograph |
| KAON | Keck Adaptive Optics Note |
| KAT | Keck Angle Tracker |
| KCSF | Keck Common Service Framework |
| LGS | Laser Guide Star |
| LOWFS | Low Order Wavefront Sensor |
| LLT | Laser Launch Telescope |
| LSS | Laser Safety System |
| LTO | Laser Transport Optics |
| MEMS | Micro-Electro Mechanical Systems |
| MSCS | Multi-System Command Sequencer |
| NFR | Narrow Field Relay |
| NGAO | Next Generation Adaptive Optics |
| NGS | Natural Guide Star |
| OA | Observing Assistant |
| OOCD | Observing Operations Concept Document |
| PA | Position Angle |
| PSF | Point Spread Function |
| RTC | Real-Time Controller |
| TT | Tip-Tilt |
| TTFA | Tip-Tilt Focus and Astigmatism (LOWFS) |
| TWFS | Truth Wavefront Sensor |
| USB | Universal Serial Bus |
| VA | Vertical Angle |
| WFR | Wide Field Relay |
| WFS | Wavefront Sensor |

# Referenced Documents

KAON 550: System Configurations Spreadsheet

KAON 569: Non-Real-Time-Controls System Design Report

KAON 668: Device Control Architecture

KAON 671: Control System SW Architecture: Container Component Model

KAON 672: Control System SW Architecture: Connection Service

KAON 673: Control System SW Architecture: Logging Service

KAON 676: Control System SW Architecture: Configuration Service

KAON 677: Control System SW Architecture: Alarm Sub-System and Service

KAON 679: Control System SW Architecture

KAON 682: NGAO Master Device List

KAON 705: AO-Level Control Loops

KAON 715: NGAO Preliminary Motion Control Design

KAON 735: Main AO Facility Sequence

KAON 736: Main LGS Facility Sequence

KAON 739: NGAO Calibration Methods and Procedures

# Document Overview

This document is divided into five main sections:

- Controls Architecture Overview – This section presents an overview of the NGAO distributed controls architecture as discussed in KAON 679: Controls System Software Architecture, focusing primarily on the component-container distributed model and the KCSF common services.

- Motion Control Software Design – Discusses the motion requirements, controller types, and architecture for low-level device control.

- AO System Software Design – Details the down-link control process of the AO system.

- Laser System Software Design – Details the up-link control process of the LGS system.

- Risks and Issues – Summarizes the remaining risks and issues that need to be resolved.

Each of the software design sections will give an overview of each subsystem as well as list their expected inputs and available outputs to the system. The AO and Laser software design sections will also detail the device hierarchies managed by each software controller, as well as the active tracking and offloading control loops. The following template is used to detail each controls subsystem.

## Subsystem Template

Each subsystem of the NGAO controls architecture will be described in detail, and follow a consistent template as described here.

### Description

An overview of each subsystem is given by describing its responsibilities, component composition and external collaborators. Each subsystem is typically accompanied by a simplified collaboration diagram showing the command-response relationship between components. The diagrams use the following notation and symbols:

|  | A green collaboration item represents an NGAO controller that directly interfaces with and commands a hardware device. The component is identified by a descriptive name and its software class implementation. |
| --- | --- |
|  | A white collaboration item represents a component with strictly software dependencies. This can be a software task or mid-level composite. The component is identified by a descriptive name and its software class implementation. |
|  | A blue collaboration item represents one of the main NGAO subsystems. This composite implementation manages a number of software and hardware components. The object is identified by a descriptive name and its software class implementation. |
|  | A collaboration item that represents an NGAO component external to the depicted system. This component can be a device controller, mid-level software component, or high-level subsystem, and will follow the same color coding described previously. |
|  | A circle with a dashed line represents unspecific or otherwise arbitrary input sources to a component (e.g. an operator tool or script). |
|  | A collaboration item depicted as a stack represents multiple instances of a single component (the count defined by the connecting line). This component can be a device controller, mid-level software component, or high-level subsystem, and will follow the same color coding described previously. |
|  | A dashed lightning bolt represents a telemetry input to a component. The information being passed is described by the named arrow. |
|  | A solid line between components represents a peer-to-peer connection. The arrow defines the direction of the commands. It is assumed the client component can and will receive information as a return parameter or through state telemetry. |
|  | Defines multiplicity for the encompassed components. This notation exists primarily to reduce duplication of grouped elements in the diagram. |

### Interfaces

The Interfaces section details the mechanisms provided by the subsystem for input and output with external components.

#### Inputs

The primary mechanism for commanding and interfacing with a subsystem is provided through the KCSF command interface. Methods are available to get and set attributes, execute commands, and monitor changes in component state. This section details all of the available actions recognized by the subsystem.

#### Outputs

NGAO subsystems will provide status and device information in the form of KCSF events. Any KCSF component can subscribe to a subsystem's published events. This section details each of the subsystem's event streams.

### Control Loops

A subsystem or component may implement asynchronous background control tasks. These tasks could be responsible for device tracking, error offloading, and health monitoring. This section details all of the background tasks that will be performed by the subsystem and its controllers.

### Devices

A subsystem is a logical organization of a set of hardware devices and their associated software controllers. This section will detail each of the devices managed by the subsystem and how they are used to satisfy the observing use-cases and requirements.

# Controls Infrastructure

The entire NGAO system can be viewed as a distributed controls system, in which many different components communicate with each other in various client/server or master/slave relationships using possibly several communication networks and protocols. We have represented a high-level abstraction of the NGAO system in the block diagram shown below in Figure 1.

Figure 1: A Block diagram of the NGAO system viewed abstractly as a distributed controls system

Each block in the diagram represents a set of control functions unique to a particular system or subsystem that can logically be grouped together. The grouping is an abstraction, as the various functions represented by each block are actually to be implemented using multiple software and hardware components that may also be distributed.

The software architecture we have chosen to build the NGAO controls system is a framework based on the Advance Technology Solar Telescope’s (ATST) Common Services Framework (CSF), described in the next section. We have dubbed the architecture the “Keck Common Services Framework”, or KCSF. The KCSF provides the base infrastructure needed to develop a distributed control system.

In a distributed system, components will need to communicate with each other, perform certain actions, and work together in a pre-defined way to form a larger system. The framework enables this by providing common communications, common data types, common commands and common tools throughout. The solution uses a component based development model that supports methods, attributes and properties. The framework provides a kernel API which hides all the details of network access and provides object browsing, discovery and common services.

Figure 2: A block diagram of the NGAO controls infrastructure

This diagram depicts the controls system as a software hierarchy. At the top level are the main interfaces to the various subsystems (some of these interfaces are referred to as sequences in the diagram). All user commands to the subsystems pass through these top level interfaces. The middle level of the hierarchy represents an abstraction of more complex lower level control tasks, namely the basic control functions for that subsystem. Users do not access the system at this level except for engineering or troubleshooting purposes. Finally, at the bottom level of the hierarchy are the devices controlled by the control system themselves. The telescope, LTCS, instrument, and atmospheric profile are outside of the direct control of the NGAO system.

## Framework

We have chosen to use a Component-Based Development (CBD) model. This model focuses on removing dependencies between objects by designing components with strict and well defined interfaces. As part of this design a narrow interface model will be employed to reduce component interdependencies. A narrow interface essentially reduces and restricts a components public signature to a well-known and expected set of general methods, thus allowing otherwise unfamiliar clients to interact with a distributed component. Specific functionality is defined through a set of configurable parameters passed to the component interface. This allows developers to add or remove functionality from a component with minimal impact to the existing infrastructure and design. The goal of the component development process is to be able to isolate the functional implementation of a task within a single plug-and-play object that can be modified or swapped with an alternative implementation without impacting the rest of the system. A user should be able to replace a component that satisfies a similar use-case without creating or breaking any dependencies in the system.  Refer to KAON 671 for more details.

Supporting the component based model will be a data communications infrastructure that is adopted from the ATST Common Services Framework.  Much of the implementation of the communications infrastructure is based upon services and support features found in third-party communications middleware packages. However, the framework isolates the dependence on third-party middleware from the rest of the KCSF software system so that replacing the middleware is always a viable option. We have successfully prototyped three different communication middleware packages (ICE and two versions of DDS). In addition, different middleware can be transparently adapted for different services if they prove to be better in one particular area (i.e., middleware can be mixed and matched as needed). The connection service that encapsulates the chosen middleware essentially forms a software bus to which components can connect. This is illustrated in Figure 3: Middleware communication with multiple devices., which shows many components attached to an abstract software bus. The components all use a simple abstract interface (the software bus) to communicate with each other and are oblivious to the underlying communications mechanism.

Figure 3: Middleware communication with multiple devices.

The framework takes care of the following:

- Automatically registering all components when they are created so that they become known on the bus.

- Automatically un-registering all components when they are destroyed so that they are removed from the bus.

- Allows any component to be referenced by its fully qualified name, regardless of its location.

- Allows operations to be performed on a remote component as though it were local to the caller.

- Transparently handles network problems such as packet drops etc.

Moreover, the communications infrastructure must be capable of high throughput, varied data size and diverse message types, using different messaging paradigms. KCSF supports this by allowing different communication middleware to be used without affecting the application developer and by using a middleware that supports:

- Peer-to-peer command messaging, allowing arbitrarily complex messages to be sent directly from one component to another.

- Publish-subscribe event messaging, allowing the generation and reception of messages by components without regard to the intended message recipients or sources.

- Simple connection support. All that is needed for an application to establish a peer-to-peer connection to another application is the name of the target application. The communications infrastructure will locate the target application (possibly starting it if it is not running).

- Heartbeat monitoring. Applications are watched and an alarm is raised if an application unexpectedly stops responding.

- Distributed systems. Distributed systems are easily supported, allowing seamless integration of third party modules.

By combining the communication middleware with component-based development, services and a technical architecture, a number of benefits can be realized:

- System details are hidden from the developer and user. This allows application developers to concentrate on applications, not the underlying infrastructure and allows system developers to make infrastructure changes with out affecting deployed applications.

- By using a narrow interface, components can be updated as needed without having to worry about version mismatch on interfaces or methods. The KCSF API is very simple and is designed such that all third party communication middleware is completely isolated from the API. All major services are provided through standard interfaces promoting simplicity and consistency. Dependencies between objects are eliminated, resulting in improved ease of maintenance. As a result components may be swapped out with alternate implementations without impacting the rest of the system. Users can replace a component with another that satisfies a similar use-case without creating or breaking any dependencies in the system. The component model makes it very easy to use components from different development groups in an application.

- The technical architecture takes care of deploying components. Components can easily be mixed and matched and even be added or removed during run time, as can services, effectively changing the system function on the fly. The implementation of the CCM provided by KCSF allows components to be installed, uninstalled, started, stopped or updated without bringing the entire system down. For deployment, the system simply needs to be configured by the engineer to tell it which components to deploy in which containers and on what machines.

Once a component is deployed, the API provides access to its internal state as well as how it is connected to other entities. Parts of the applications can be stopped to debug a particular problem, or diagnostic components can be brought in. Instead of staring at hundreds of lines of logging output and enduring long reboot times, applications can often be quickly debugged using a live command shell. Any entity can be tracked to see its current health and if it is registered.

## Common Services

Common services are infrastructure utilities that provide a black box implementation of typical, repetitive, or convenient functionality need by a software system. KCSF Common services are the backbone of controls architecture, and provide an abstract layer between the controls system and the underlying middleware infrastructure. As a distributed architecture, KCSF offers services to find and connect to remote processes, communicate between individual components in a system, and interface with other system outside of the KCSF infrastructure. Additional services are provided to manage software configurations; monitor the health and status of the system; respond to alarms; and archive system data and events.

The following sections give an overview of the KCSF common services. Additional details can be found in the Common Services section of KAON 679: NGAO Software Architecture. Any NGAO specific functionality or considerations related to the use of the common service will be discussed here.

### Connection Service

KCSF provides a utility to find and create peer-to-peer connections to remote components in a distributed system through the Connection Service. This service is implemented as a black box to hide the details of the underlying middleware infrastructure, and provide an abstract, simple interface to clients. The following functionality is provided by the Connection Service:

- Register – A distributed component is associated with a unique name and is made public to the distributed control system. Once registered other components can find this object across the network by its unique name.

- Remove – An object is removed from the registry. The object name is no longer a valid endpoint on the network.

- Connect – Allows a component to find and connect to other components in the system based on their unique registered names. This creates a direct peer-to-peer connection through a proxy interface.

- Disconnect – A peer-to-peer connection with a remote component is terminated. The proxy interface no longer has a reference to the remote client.

The registration and removal process are handled automatically by the KCSF middleware: developers do not need to explicitly perform these tasks in their code. A component developer only needs to be concerned with opening and closing connections to distributed components to perform the functions of the controls system.

NGAO will not require any special implementation or any additional functionality that is not currently available with the standard KCSF Connection Service.

### Event Service

The majority of broadcast and non-command messaging is satisfied using events. Events are based upon the publish-subscribe design pattern, and allow a component to make information available to the entire distributed system without requiring an explicit connection between clients. Events are designed to provide fast peer-less communication channels using the capabilities of the underlying middleware. KCSF provides access to the event system through the Event Service.

The KCSF Event Service has the following properties:

- An event stream (a.k.a topic) represents a many-to-many mapping: events may be posted to the topic from more than one source and received by zero or more targets. Typically, however, most event topics will have a single source.

- Events posted by a single source into an event topic are received by all targets in the same order as they were posted.

- Delivery of events to one subscriber cannot be blocked by the actions of another subscriber.

- Events are not queued by the service. A "late" subscriber will not see earlier events.

- The event service does not drop events. A published event will be delivered to all subscribers.

- Event topics are identified by a unique name.

- The Event Service supports arbitrary topic names.

- Events are automatically tagged with the source and a timestamp.

The Event Service provides a simple API to interface with event topics:

- Subscribe – Start listening to events published on a particular topic.

- Unsubscribe – Stop listening to events on a particular topic.

- Post – Publish data to a topic.

Events are received by a component by attaching a callback to a subscription. The event service, upon receipt of an event, invokes this callback in a separate thread. All events received from the same subscription use the same thread: delivery order is preserved within the callback processing. If events are received faster than the callback can process them, the unprocessed events are locally queued within the event service. This is a potential problem, but represents a trade-off of mutually exclusive goals. Component developers are encouraged to write callbacks that process quickly. Numerous approaches are available to handle the case where the required action cannot be performed quickly - the best approach to use is dependent upon the nature of the specific task and is thus the responsibility of the component developer.

### Health Monitoring

The ability to quickly ascertain device health throughout the NGAO control system is very important. KCSF supports this by providing a health monitoring service. The health of each component is monitored by its parent container. The various health states for a component are: good, ill, bad, and unknown, and are defined as follows (in order of worsening health):

- Good: No problems have been detected by the component, it is fully operational.

- Ill: Problems have been detected, but they do not prevent observing. Data quality, however, may be affected. It may also be the case that operation of the component will fail soon if corrective action is not taken. The component is partially operational.

- Bad: Severe problems have been detected. The component is unable to operate correctly. Corrective action is required.

- Unknown: The component is not responding. It may or may not be operating. This health value is not set by the component (obviously) but may be set by the health service.

The health service automatically posts an event showing changes to the component health and logs a warning on worsening health and a note on improving health. When a health condition worsens to bad or unknown the log message severity switches from warning to severe.

All components define an abstract performCheckHealth method, which must be implemented by the component developer. This method is responsible for determining the health of the component and returning its health status. For most NGAO controllers the health status will be based on feedback from the device or motion controller, and will typically be in either a good or bad state. Higher level composite components and sequencers will have to calculate health based on the overall status of their individual controllers. For these components determining health will be based on the importance of specific devices over others in the control system, where many failure scenarios can be categorized as either bad or ill.

### Logging

Logging is the capability of writing messages to create a record of system events. Typically the content of a log message isn’t intended for immediate review or action, and is instead intended to trace the execution of tasks and provide feedback of system performance. Log messages can typically be written to a variety of output streams for temporary viewing or long-term preservation. KCSF provides logging functionality through the Log Service.

The Log Service offers clients the ability to write messages to one of three types of outputs: console, file, and database. A log message is categorized by a specific log level, indicating the type of message and its importance. The KCSF Log Service defines a total of seven log levels ranging from a simple trace message to an emergency system event. When a message is issued the Log Service will automatically append timestamp and source information and format the message for consistency. The Log Service provides convenience methods for each of the defined log levels to simplify component development. The KCSF Log Service is described in detail in KAON 673: NGAO Software Architecture: Logging Service.

In addition to the Log Service, KCSF also provides a graphical log viewer. This tool allows users to inspect persistent log files saved to a database. Logs can be searched by specifying a range of options, including date, log levels and sources. All log events that fit the search options will be displayed in a simple table organized by date-time, level and source.

The KCSF Log Service and viewer will be used unchanged for NGAO.

### Archiving

The Archiving service allows users to capture and preserve any telemetry or diagnostic data publicly available in the system. The client would tell the Archiver the component and specific attributes that the user would like to write to the Data Server (a relational database). As long as the component is active the values of the attributes will be written to the archive storage. Two write models are currently being explored for archiving: push and pull.

In the push model the archiver is responsible for creating attribute monitors, and listening for events on the requested archive streams. When an event occurs the data is written to the database. In the pull model the archiver will periodically command the components to write their attribute data by trigger a class action. The component would set the data in an attribute list and send it to the archiving service where it will be written to the database.

Data can be retrieved from the archive store using the KCSF provided management tools. A GUI will be available that will allow users to browse and search archive data, and export it to a desired format based on date, component, or run ID.

The Archiving Service and tools must be developed for KCSF.

### Alarming

Alarms are reports of component failure (software or hardware) or other abnormal behavior within the system. Alarms occur asynchronously in a random fashion. Some alarm conditions may clear themselves and others may require operator intervention. The KCSF Alarm Service provides capabilities for components to set and clear alarms, and provides a managing system to detect the occurrence of alarms and initiate the appropriate system response. The Alarm Service is part of the larger NGAO alarm system.

The Alarm System is described in detail in KAON 677: NGAO Software Architecture: Alarm System

### Configuration

The configuration service is just one part of the larger configuration system in KCSF, which in general provides the ability to manage the configurable properties of each component in the NGAO system. It consists of a configuration database, the associated tools to manage the database, and the configuration service, which provides a client interface to each component to request its own configuration information from the database.

During initialization, each component calls the configuration service to access its specific configuration data. There is a specific method for Container Managers, Containers, and Components that accept a fully-qualified name used to lookup the information relevant to each of these types. The Configuration Service utilizes attribute lists to return data to the invoking object. The details of this process are handled automatically by the parent container with no action required on the part of the user.

## Component Architecture

The CCM defines a distinct separation between the technical and functional requirements of a task. The model is based on the utilization of two distinct software modules: containers and components. Containers provide a structural or logical organization to software objects (components), and are responsible for their lifecycle and management. This is illustrated in Error: Reference source not found.

A container will create components, start them running, shut them down, and remove them from the system. Containers provide a uniform method of system management and allow component developers to ignore the majority of the technical requirements of individual objects. Containers, component interfaces and abstract implementations are provided as part of the software infrastructure.

A component is a piece of software that implements the functional requirements of the system and conforms to a well defined interface or set of interfaces. The interface abstracts away the actual implementation of a component, and from the perspective of the system, there is no difference between one implementation and another. In this way, a developer can swap out multiple implementations of a component for a specific use case without affecting the system as a whole. This is most evident when you compare the maintenance and upgrade of a component-based and a non-component-based system.

Implementation of the NGAO Controls system will focus on the design and development of component subclasses. KCSF provides two component implementations that can be used to build a controls system. They are the controller and composite base classes. From these classes a controls system can derive all of the device and software control functionality required by the system.

### Controllers

Controllers are the lowest level software component in a control system and interface directly with the motion control and device hardware. Controllers provide functionality to set or get properties; execute synchronous or asynchronous tasks; and monitor state with callbacks. The controller development process is focused on overloading the stubbed perform methods to recognize and execute actions and tasks issued by clients. The actual receipt and dispatching of actions is handled by the controller’s technical architecture, and is completely invisible to the developer.

The NGAO controller design will be based on a state driven architecture. Actions will be used to issue transition between the different run-time states of the component. The state machines for each controller will vary but typically follow the standard state pattern for motion controllers: home, initialize, slew, inpos, track, halt, and shutdown. State machines will be built using the State Machine Compiler (SMC) utility. This tool allows developers to define states and transitions by creating a dispatching wrapper that executes tasks on the driver class (in this case the Controller). A low-level state controller will be created for each device type (as outlined in section Low Level Motion Device Control). The developer will only need to override the task implementations in the subclasses to gain the desired functionality.

The following sections detail the generic controller state machine for continuous motion devices.

#### State Machine

The following state diagram describes the NGAO Controller state machine for a continuous motion control device. Simplified versions of this state machine will be developed for non-motion control devices.

#### INIT

When a device controller is launched it will be in the START state. Before the device can be used it must be initialized. When a controller is first initialized it will load configuration information, connect to dependencies, and home the device. This is known as a full-init, and can be performed from the START and HALT states. In contrast a fast-init -- which is performed during all other Initializations -- will reload configuration, refresh middleware connections, and configure the device: although the device may move the encoder position is not zeroed.

The homing process is responsible for calibrating the motor and physical location of the device in its range of motion. During the home process the device will typically move to one of its limits and then back-off the limit until it hits the home sensor. Once the home sensor has been triggered the encoder position is zeroed and the device calibration is complete. For devices that do not have limits, such as a filter wheel, the calibration process may simply move the device in a particular direction until a dedicated home sensor is crossed. Once the homing and initialization process is complete the controller will enter the INPOS state, signaling that it is ready for use.

#### SLEW

Assuming the device does not encounter any problems while operating, the motion controller will typically iterate between the SLEW and INPOS state as the device is repositioned during an observing program. The SLEW state is a transient state: once the device position read back matches the demanded position the controller will automatically transition to INPOS. Depending on the properties of the device and the distance it has to travel the controller may stay in the SLEW state for an extended period. If the slew process would cause the device to reach or pass one of its defined limits the control will automatically enter the LIMIT state, aborting the slew command.

#### LIMIT

The LIMIT state is a software safety-state that prevents clients from driving the device into a limit-fault state.  The controller and device can still be commanded while in this state, but must be given a new command that does not drive the device into a limit. The advantage to using the LIMIT state is that rather than relying on the motion controller to detect a limit, with its only recourse being to fault and reinitialize, we can keep the device in an operational state and warn the user the commanded position is unachievable.  From the LIMIT state the controller can be halted, slewed or told to track.

#### TRACKING

The Controller design also allows developers to implement and utilize device tracking functionality. Tracking is the process of synchronizing a device’s state to the state of other devices in the system. An example of a tracking device is the image rotator, which is positioned based on the changing orientation of the telescope and field. Devices with tracking capabilities implement the TRACKING and transient TRACK state to perform coordinated motion control with other components in the system (note: not all continuous devices will have tracking requirements.)

When a Track is issued from one of the open-loop states (INPOS or HALT) the controller will enter the TRACK state. This state is responsible for calculating the starting position for the device (based on the external sources, e.g. telescope position), and slewing the device to this position. TRACK is very similar in functionality to the SLEW state, except on completion of the move the servo loop is not opened (the device remains powered). Instead, new positions are continually processed and feed to the device, causing it to follow an ever changing demanded position. The controller will enter the TRACKING state once it achieves its starting position, and will stay in this state until it is commanded to halt or slew to a new position.

As with the SLEW state, the tracking states can enter LIMIT if the calculated position will cause the device to go beyond its permitted range of motion. On the LIMIT transition the servo loop is opened and power is cut from the device.

#### HALT

At any time post initialization a client can interrupt the controller and device by issuing a Halt. The HALT state is designed to immediately disable power to the device by opening the servo control loop. If the device had been tracking it will also open the software control loop causing it to stop responding to system events and calculating demand position. Since a halt can be issued at any time the designer must make sure that there is no point in a device’s standard operation where immediately severing power can result in unsafe conditions. If such a risk exists, the component / motion controller developer must take this into account when developing the HALT functionality. A user can resume operations from the HALT state by issuing a new slew or track operation, or if desired, a full-initialization can be performed (configure and home). Although HALT will immediately disable power from the device there is still the possibility that the device could continue to drift before coming to a stop. If the device requires that all motion stop immediately a breaking system will also have to be implemented and added to the functionality of the halt task.

#### FAULT

For tracking and non-tracking continuous devices, errors can occur during operation that need to be caught and corrected before the device can be used. FAULT represents a catch-all state for the device, and any transition could potentially cause the controller to enter this state. When a fault occurs the operator will need to know the cause of the fault, and possibly take action to resolve it. For most errors an initialization or recover may suffice.

#### RECOVER

The RECOVER state offers a convenient alternative to initialization for fault recover. When a device is initialized to clear a fault it will most likely move away from its current position, and will loose all sense of state prior to the fault. RECOVER allows the client to clear a fault and restore the same device and controller state that existed just prior to the fault event; all of which will be stored internally by the controller. Depending on what caused the device to fault the controller may issue a TRACK, SLEW, or enter INPOS. As the recover will attempt to return a device back to its previous state this could cause the device and controller to enter FAULT again, depending on the original error. The client should be aware of the cause of the fault beforehand to ensure that this recover-fault cycle does not occur.

### Composites

Composites are similar to controllers in basic functionality: both provide a state driven software component with set, get, and execute methods. However, they are different from controllers in the role they fill in the controls hierarchy. As the name implies, a composite is a collection of components. Specifically, composites manage the life-cycle and coordinate the execution of one or more low-level controllers and sub-composites. Composites themselves do not directly interact with hardware. Instead they build the software hierarchy by dividing the system into discrete units. Each composite and its subordinate components are intended to be fully self-contained, allowing clients to treat a particular portion of the system as a black-box.

The NGAO Controls system will be designed as a hierarchy of composites, where each layer abstracts away the underlying detailed functionality of the sub-components. At each layer, the client is provided with a high-level interface to command the execution of common sequences required by the controls subsystem. As the composite move farther away from the low-level devices the interfaces become more general and identify less of the underlying infrastructure. At the highest levels a set of global commands are offered to initialize, acquire, track, and halt the system.

### Device Control Interface

The following sections detail the common input and output interfaces available to most NGAO components. In the AO and Laser system portions of this document the individual subsystem interfaces are detailed. Each subsystem will expand on these interfaces and include any device specific functionality involved in processing these actions and events. Subsystem may implement additional interface capabilities or omit specific functionality entirely. See each subsystem for more information.

#### Inputs

Inputs are provided through the standard KCSF component interface using direct peer-to-peer connections through proxies created by the Connection Service. These actions and related parameters are wrapped in an attribute list and set to the client for handling through the Execute method. The following actions are available to most components.

Initialize

The Initialize action is responsible for configuring software and hardware, and homing motion control devices. A component initialization can be categorized as a fast or full initialization. A fast-initialization typically involves refreshing soft connections and configuring hardware. A full-initialization primarily relates to hardware and involves device calibration and homing. The type of initialization executed during a call to Initialize will depend on the current state of the device; full-initializations are typically performed on system startup, and fast-initializations are performed for most executions thereafter.

All NGAO controllers and composites implement the Initialization action, which must be performed prior to use. For software composites and subsystems the Initialize action is primarily responsible for loading configuration, setting default properties, connecting to dependencies, and initiating telemetry. Controllers perform a similar initialization process but also include device preparation and calibration which involves: slewing the device to a home switch; setting the encoder count (absolute position) to zero; and moving to a starting position (e.g. middle of range).

When the initialization process completes the component will be ready for use.

SetMode

The SetMode action is responsible for configuring the AO or Laser system for a particular observing configuration. The highest level components in the NGAO Control System (AO and Laser Sequencer) implement the SetMode action, and are aware of the required configurations of their individual subsystems for the available observing mode. The configuration process can include slewing devices, configuring device properties, making connections to remote components, disabling device controller, and powering down hardware. On completion of a SetMode action the subsystem and its devices should be ready for observing in the desired mode.

Offset

The Offset action is available to all continuous motion control devices and allows clients to apply a specific offset to the device position. This offset can be defined relative to the current position or an absolute position, and is specified by physical units (e.g. mm) or an enumerated representation of common offsets.

Slew

The Slew action is common to all motion control devices are is responsible for commanding the device to a specified position. The specific format of the data is dependent on the type of device (tip-tilt, linear stage, etc.) and the number of degrees of freedom. Most controllers also permit the use of named positions allowing clients to specify an enumerated value in place of device counts or physical positions. On completion of the Slew command the controllers associated devices should reach their commanded position.

Track

The Track action is available to all motion control tracking devices, and some higher level composites. A tracking device is any device that runs in an event driven (or periodic) mode based on system events -- also known as closed loop motion control. The Track action commands the component and associated hardware to enter closed loop control to respond to one or more stimuli. For example, a focus stage that responds to focusing errors measured by a sensor. A component in track mode will stay in track until it receives a command to Halt or a problem is encountered.

Acquire

The Acquire action is a high-level operation available to composites and subsystems. This action is responsible for implementing the component and device sequence to acquire an observing target(s) (e.g. a tip-tilt NGS). Typically this action accepts target coordinates and properties that are used to calculate device positions, loop rates, and other system configurations. Depending on the responsibilities of the subsystem an acquire may cause devices to slew, sensors to reconfigure and control loops to close. On completion of this action the subsystem should be properly configured for observing.

Halt

The Halt action is available to all NGAO components and is designed to immediately open all tracking / offloading loops and cut power to all devices managed by the component. This action is executed to terminate a sequence -- either in response to a system failure or the end of an observation. The Halt action should leave all devices and components in a useable state: they should not require an initialization or special recovery sequence; although the acquisition sequence may not to be restarted before observing can continue.

Recover

The Recover action is available to all NGAO components, and is used in the context of selective initialization of one or more devices during a system fault. During observing, devices may encounter problems that cause the hardware motion controller and NGAO component to enter their FAULT state. Rather than being forced to terminate the observing sequence to resolve the problem, it may be desirable to continue observing, although in a potentially degraded capacity. In such an event the operator will be able to issue the recover action at the subsystem level.

When a recover is issued to a subsystem it is propagated to all of its child components, even if they report that they are functioning properly. Essentially the recover action is sent blindly to the entire control subsystem hierarchy. As each component receives the recover action it will check to see if it is currently in a faulted state. If it is, an initialization is performed to attempt to clear the fault, and if successful, the device is repositioned and loops are closed to bring the device(s) back into the proper operating state. (Note: device acquisition is only performed if the device had faulted while in the TRACK state.)

If the device can not be recovered it will remain in its faulted state, and direct operator intervention may be required at the device level. If a device fault is cleared at the device level (e.g. through an engineering GUI), the control system will still need to be told to recover from the fault by issuing the recover action: a component will not clear its fault just because the device reports “in-position”. This is to ensure that all control system dependencies have been properly informed of the recovery and have responded as needed.

#### Outputs

Component outputs are provided through telemetry topics. Topics are identified by a unique name typically appended to the full-qualified component name. Data published by a server is formatted based on the type of data and requirements of the component, and wrapped in an attribute list. The following telemetry topics are available to most components.

State

Every NGAO component will publish its current state to allow the parent composite and interested clients to monitor its progression. The state of a component is the primary means by which composites and subsystems ensure that an action has been properly processed, and the component is operating properly. The state is typically represented by an enumerated value.

Health

Every NGAO component will publish its current health status as a good, ill, or bad enumeration. The health is used to tell operators the overall performance or condition of a component, and will also reflect the state of the motion control device (if applicable). The health status of a component is used in conjunction with the state value to ensure that the component is in the proper state and is functioning at optimal performance.

Error

An error message is available to all components, and contains a readable string message related to a fault or error that may occur for the device. Typically this telemetry is used by the operator control GUIs and should only need to be read when a device enters a faulted state.

Position

All motion control KCSF components are able to report their current device position. Depending on the controller design this may have one, two, or three degrees of freedom (e.g. x, y, z). Linear devices (such as a focus stage) will typically report their position in millimeters of motion. Rotational and tip-tilt device position is usually represented as angle or radians of displacement. Discrete devices can be in one of two possible states, which is reported as the activation of one limit switch, usually as an enumerated value (e.g. ‘open’).

Offset

Most motion control components (excluding discrete devices), allow clients to set a specific offset in base units (e.g. millimeters). These controllers will report the current cumulative offset in the appropriate unit and for each of the degrees of freedom of the device.

Framerate

Specific to camera components, this telemetry topic identifies the current operation frame rate of the camera.

Gain

Gain is common to most power devices including stimulus sources, power controllers, and cameras. Gain defines the sensitivity or voltage applied to a device.

Binning

Binning is specific to a camera device, and identifies the pixel averaging / grouping used to take images.

Wavelength

Wavelength telemetry is available for cameras and wavefront sensors, and represents the wavelength (or range) of light for which the device is sensitive. This value does not change but it required when performing the correct atmospheric refraction calculations.

SensorReading

Sensor telemetry is available for all sensors, and represents the current value as reported by the device. This can be a temperature, count, intensity, or percentage. The specific format of the data will depend on the type of sensor.

## Low Level Motion Device Control

### Introduction

The devices to be controlled by AO Controls are many, and are detailed in the configuration controlled KAON 682: NGAO Master Device List spreadsheet.  As of this writing there are a total of 83 motion devices, 29 of which are laser devices.  There are also a substantial number of non-motion devices that is covered under the AO Control System.  One thing to note is that AO Controls does refer to real-time control of devices, as in realtime control of deformable or tip-tilt mirrors, but to the basic control of the devices themselves.  This includes power control, initialization, basic parameter control, etc…

Motion control refers to the control of any device that has to be positioned before it can be used.  This control depends on the type of motion device being controlled.  The particular devices are being designed in parallel and are described in KAON 715: NGAO Preliminary Motion Control Design.

Although we haven’t converged on a particular hardware control system, we are very familiar with Delta Tau PMACs as they are used extensively in our current AO and Interferometer systems.  Our current thought is to continue to use PMACs, especially given its programmability and scalability. We are however considering moving away from the VME design to a smaller PXI/CPCI form factor.  Also, for low precision intermittent motion devices we may use smart motors for simplification of cabling as they can be chained together and would be controlled over a serial interface.

The low level motion control functions consist of one or more device drivers that implement an API for interacting with the motion control system to provide the most basic level of control. This device driver will most likely be supplied by the manufacturer of the motion control electronics. An abstraction layer will be added above the low level driver as part of AO Controls.   This layer would hide all the hardware specifications and implementation of the actual motion control hardware.  At the top of this control block is a generic state machine that implements the basic coordinated functions required for controlling a motion control device with an arbitrary number of degrees of freedom. This concept is used in the existing Keck AO and Interferometer systems.

KAON 715 identifies all of the current device types / categories required by NGAO. In order from least complicated to most demanding, they are:

- Shutters

- Low precision  non-tracking devices

- Medium precision non-tracking devices

- High precision non-tracking devices

- Tracking devices

- Extremely high precision tracking  and non-tracking devices

- Coordinated high precision tracking and non-tracking devices

From a software perspective, these can be grouped into the following generic categories:

- discrete commanded devices

  - shutters

- non-tracking devices

  - low precision  non-tracking devices

  - medium precision non-tracking devices

  - high precision non-tracking devices

  - extremely high precision non-tracking devices

  - coordinated high precision non-tracking devices

- tracking devices

  - tracking devices

  - extremely high precision tracking devices

  - coordinated high precision tracking devices

The intent would be to develop each type once and instantiate it for all the necessary devices.   This will simplify implementation and allow for quick and easy changes or addition of new devices. Each of these types is discussed in the following sections.

### Discrete Commanded Devices

#### Description

This is the simplest of device types and consists of in/out devices with loose positional requirements; actuators other than motors (e.g., solenoid, pneumatic, etc); and devices with switches or hard stops used to define positions.

Software controls for this device type will involve sending a discrete output signal to a digital I/O controller board and status information will be returned in the same manner.  More than likely the status will consist of limit switch readings where only one limit can be active at any given time.  When no limit switches are active, this indicates that the device between positions and thus assumed to be “travelling”.  When both limit switches are active, this indicates an error.

#### Inputs

Commands to this device type will be simple boolean 0 or 1 commands where 0 means one state and 1 means the opposite state.  For instance open/close or in/out.

#### Outputs

Since these are simple discrete devices, the outputs will consist of two boolean state readback values which correspond to opposite positions, i.e. open/close, in/out, or cw/ccw, etc…  A valid status will be when only one state is active and the other is open.  When both states are open, this indicates that the device is travelling between states.  When both states are active, this will indicate a fault condition.

### Non-tracking Motion Devices

#### Description

Non-tracking motion devices are devices that can be positioned anywhere within a fixed position range.  Examples of non-tracking devices are:

- dichroics

- fold mirrors

- focus stages

- lenslet stages

These devices are positioned during configuration or acquisition and then left in place with its servo loops open, i.e. no power.  They are not moved during an observation. There will be varying degrees of precision: low, medium, and high, all based on encoder resolution with matching DAC output control.  However from a software perspective these are all treated the same.  They must all be initialized, which means that their encoders must be homed.  This involves driving to a home switch, which could be a limit switch, and loading a particular position once the switch is sensed then applying a home offset to position the device to a known position.  Depending on the device, this could be a specific position or some where in the middle of travel, which we call encoder position 0.  This homing process is built into the PMAC and can be configured on a motor by motor basis as to which direction to home, what to use for a home switch, and  the specific home offset.  In most cases this is a straight forward operation meaning that the device axis can be freely homed without any limitations.  However in the case of the LOWFS OSM (object selection mechanism), care will need to be take to prevent them from colliding into one another during the initialization process.  The actual homing of the OSM pickoff arms will also be tricky and require a customized homing process.

Once the device is homed, it will then need to be positioned to specific locations such as in beam or out of beam.  This will be done by using named positions instead of actual position values or encoder counts.  The named position will be mapped to a specific encoder position which can change as needed but the named position used will always be the same, i.e. in-beam, out-beam, etc.

For some devices, i.e. focus stage or a lenslet stage, positioning will need to be based on phsical positions then adjusted throughout the night.  The initial position, in the case of a focus device, will be pre-calibrated and can have a named position, such as initial-focus, but subsequent moves / adjustments will need to be commanded in some form of engineering unit that is logical to the device.  So instead of adjusting focus by some number of encoder counts, you would specify a move in terms of millimeter of focus, for instance.

#### Inputs

- Home / Initialize command

- Abort command

- Move to named position command

- Move to logical unit position command

- Move to raw encoder count command

#### Outputs

- Current position in

  - raw encoder counts

  - equivalent logical unit position

  - applicable named position

- In position status

- Home status

- Servo loop status

- DAC output

- Following / position error

- Limit conditions

### Tracking Motion Devices

#### Description

Tracking motion devices are devices that are constantly commanded to a new position during an observation.  The actual position is based on and synchronized with an external source, i.e. telescope az or el, and is updated at a periodic rate.  Like non-tracking devices, tracking devices require various levels of precision as well as coordination of moves with other DOFs, i.e. steering mirrors, and in some cases spatial position constraints to avoid collisions and occlusion, i.e. pickoff arms.

From a software perspective, tracking devices are basically non-tracking devices with the added capability of being able to handle continuous position demands by keeping its servo control loop closed until explicitly opened by an explicit command.  Ideally the controller will be able to handle velocity and position loops to help keep motion smooth and constant.

#### Inputs

- Home / Initialize command

- Abort command

- Move to named position command

- Move to logical unit position command

- Move to raw encoder count command

- Continuous tracking demanded position commands

#### Outputs

- Current position in

  - raw encoder counts

  - equivalent logical unit position

  - applicable named position

- In position status

- Home status

- Servo loop status

- DAC output

- Following / position error

- Limit conditions

## Low Level Non-Motion Device Control

### Introduction

As defined by KAON 668: NGAO Device Control Architecture, non-motion device control refers to the control of any device or subsystem that is not creating motion and is not one of the devices requiring real-time control for wavefront sensing and correction. This includes power control, environmental control, camera control, mirror control, general sub-system control (e.g., the laser units, the RTC) and the controls infrastructure.

Non-motion device control is, in general, low speed without tight timing constraints. The commands are frequently single events (e.g., operator initiated) and are expected to be executed on the timeframe of seconds.  Ethernet will be the primary means of communicating with devices for control. If a device does not have an Ethernet interface there may be simple Ethernet solutions which exist (e.g., terminal servers to communicate with RS-232 based devices, or digital I/O modules with an Ethernet interface to generate digital I/O signals).  To reduce software integration efforts, Ethernet devices should support an industrial protocol (Modbus TCP/IP, Fieldbus, etc) rather than a web (HTTP) based interface whenever possible.

The following sections describe the various types of device control which would be supported by this device layer.   Currently the following device interface control methods have been identified in KAON 668:

- Remote power control

- Camera / detector control

- Digital input/output

- Analog input

- Analog output

- Remote reset control

- Video

- RS-232 and USB

As part of the synergy with the TCS Upgrade project we hope to use the same generic IDaqDevice interface currently under development by TCSU.   below, illustrates the concept of a generic IDaqDevice interface taken from a TCS Upgrade example. The DAQ specific controller would be created to match the IDaqDevice interface. In this case we show the use of the NIDAQ Device Controller. This device controller uses the NI-DAQmx API from Nation Instruments and supports the ICE IDaqDevice interface.  The NI-DAQMmx drivers supports a number of device types:

- Multifunction DAQ

- Analog output

- Digital input/output

- Counter/timer

- High speed digitizer

- Dynamic signal acquisition

A DAQ Client will run as part of the application and will be responsible for setting up the IO, alarming and reading/writing to the devices via the ICE interface

Figure 4: Generic KCSF Device Interface Layer

Serial and Ethernet interface device types, which are not supported by NI-DAQmx, as well as special camera or detector controls, will be handled as part of the existing KCSF device/controller infrastructure.  For more details refer to KAON 679: NGAO Control Software Architecture.

Non-motion device control not covered by the above IDaqDevice interface will  be described case by case in the following sections.  However, they should all be able to fall under the KCSF device infrastructure and be designed with the generic IDevice interface. The IDevice interface allows a device to access the get, set, execute and xxxMonitor methods which should be able to support the full command set for each device.

### Remote Power Control

Remote power control is used to power on / off devices remotely through COTS power controller device.  The observatory currently uses two different supplied power controllers (Pulizzi and APC) both of which are serial device types.  Serial device types would be custom implemented using KCSF’s device/controller concepts.

Should a more customized solution based on solid state relays and digital controls be necessary, it would easily fit into KCSF’s generic iDaqDevice interface.

#### Inputs

Execute command attributes:

- power on <port|# | all>

- power off <port# | all>

#### Outputs

Get command attributes:

- port status

### Camera / Detector Control

There will be a number of cameras used in the NGAO system.  Most will be controlled by AO Controls but there will be some that have control interfaces as part of the data readout interface, as is the case with Camera Link. For these cameras, the control functions will be implemented directly by the RTC, but the RTC is required to make specific control commands available to the NGAO control system for initiation of these functions.

For the other cameras and detectors, we will use the KCSF device infrastructure to implement a generic IDevice interface. The IDevice interface allows a device to access the get, set, execute and xxxMonitor methods which should support the full command set.

#### Inputs

Set command attributes:

Execute command attributes:

- power on

- power off

- set frame rate

- expose

#### Outputs

Get / monitor command attributes:

- status

- frame rate

- frame count

### Video

Video servers will be used to monitor beam positions for use in alignments as well as a remote access tool to verify beam position.  The Interferometer project currently uses AXIS video servers which does not need much control once it’s configured via a web interface.

#### Inputs

Set command attributes:

Execute command attributes:

#### Outputs

Get / monitor command attributes:

- centroid position

- fwhm

- intensity

- image

### RS-232 and USB

A goal is communicate with every controllable device via Ethernet. To accomplish this, some equipment will require a protocol converter. In the case of RS-232 and RS-422/485 devices, this “protocol converter” can be implemented by  the use of a terminal server hooked up to the Eternet.  The observatory currently uses a number of Lantronix terminal servers.  Lantronix also produces a USB server if our only option for specific devices are USB interfaces.  As of this writing, it is unclear what type of devices these would be and is thus difficult to design to.

#### Inputs

Set command asttributes:

Execute command attributes:

#### Outputs

Get / monitor command attributes:

# AO System

## Introduction

For the preliminary design we’ve chosen to diverge from the original concept presented in KAON 569: NGAO System Design Report, Non-Real-Time Controls. We are proceeding with a more subsystem view approach to help us better visualize and map the control system to the way the hardware is actually laid out.  The original concept had us treating each device as a standalone entity sitting off by itself.  Software-wise the same functionality will still be developed but by visualizing things from a system perspective it helps us to better understand how the pieces fit together and ultimately how the system will function as a whole.  This new view is shown in Error: Reference source not found below.

Figure 5: AO Control Block Diagram

You can consider each block as a high level control state machine  that implements the complex coordinated control required by the devices in that subsystem. They accept high level commands from the AO sequencer and issue the appropriate commands to the low-level motion control blocks. The sequences will be implemented using a state machine compiler (SMC). The high level motion control sequences identified above are described in detail in the following sections.

## Wide Field Relay

### Description

The Wide Field Relay covers the low-order atmospheric correction hardware, and miscellaneous first-relay and low-order devices. The Wide Field Relay subsystem is a logical grouping of these hardware devices, but is not actually implemented as a standard KCSF component due to the very different management and control requirements of the individual devices. Although other devices may also be considered part of the wide field relay, they are grouped and managed under other subsystems (see the LGS and LOWFS subsystems).

Figure 6: Wide Field Relay Collaboration Diagram

The low-order DM and TT stage are under the control and management of the real-time control system. The NGAO Controls system provides KCSF component interface classes to initialize and configure the RTC devices. During regular operations the operation of the DM and TT will be handled entirely by the RTC control loops.

The IF fold dichroic, hatch cover, and image rotator are managed by the NGAO Controls system. The dichroic and cover are simple discrete devices. The fold mirror feeds the interferometer DSM with light and is only used on IF nights. The hatch cover protects the AO system input and will typically be left open for the duration of the observing period. The image rotator is the only tracking device in the low-order relay under the NGAO control system management, and is responsible for maintaining a fixed-field or fixed pupil orientation with respect to the science instrument.

### Control Loops

The following sections detail the NGAO controls loops implemented by the Wide Field Relay devices.

#### Image Rotation

Due to the design of the Keck telescopes, as a target is tracked across the sky it will appear to rotate on the science instrument. Similarly, as the telescope elevation changes, additional rotation is added to the field and pupil due to the changing optical axis of the telescope with respect to the nasmyth platform. The image rotator is designed to correct for both of these rotations to keep the field or pupil stable relative to the instrument.

Figure 7: Image Rotator Control Loop

In fixed-field observing we want the field to remain at a fixed orientation relative to the detector. In this mode the rotator should track the position angle of the target (the apparent angle of rotation on sky) which will be constantly changing as the telescope tracks the target. While in fixed-field the pupil is allowed to rotate freely.

In fixed-pupil we want the pupil to remain at a fixed orientation relative to the detector. To maintain a stable pupil the rotator will track the telescope vertical angle. The vertical angle changes as the telescope slews in elevation. While in fixed-pupil mode the field is allowed to rotate freely.

#### LOWFS Rotation Correction

During LGS observing the LOWFS sensors will be tracking the tip and tilt of the natural guide stars. From this information the system can determine the appropriate tip-tilt correction to apply to TTM1, and determine if there is any residual field rotation not corrected by the Image Rotation control loop.

Figure 8: Field Rotation Offset Control Loop

The NGAO RTC Interface Controller will calculate the residual rotation seen by the three LOWFS and make this information available as telemetry. The Image Rotator Controller will subscribe to this telemetry and correct for any residual field rotation.

### Devices

The following devices are managed under the Wide Field Relay.

#### Hatch Cover

The hatch cover is a simple discrete device positioned at the input of the AO system. When the AO system is not use the hatch cover is closed. During observing the hatch must be open.

#### Image Rotator

The image rotator is a 1-DOF rotational tracking device responsible for maintaining a fixed field / fixed pupil orientation as the telescope follows a target during observing. The rotator controller will be responsible for calculating the required position of the device based on telescope azimuth and elevation. The rotator will also receive rotation corrections based on outputs from the LOWFS RTC.

#### IF Fold Dichroic

The IF fold dichroic is a beamsplitter fold mirror mounted on a discrete motion stage that is designed to feed the near-IR AO corrected beam to the interferometer. The visible light is propagated to the NGS WFS for processing.

#### Low-Order Relay DM and TT

The low-order relay DM and TT are managed and controlled directly by the RTC system. The NGAO Controls system will be responsible for powering on the devices; loading reconstructors; setting properties; and initiating RTC servo loop commands.

## Acquisition System

During the NGAO acquisition phase of the observing sequence, the operators, science tools, and controls system will need to be able to view the sky to determine the current pointing and alignment of the telescope. This is possible by imaging the field with an on-axis camera. With this device clients can determine the precise pointing of the telescope and the orientation and position of the laser asterism. The functionality for this capability is provided by the Acquisition Subsystem.

### Description

The Acquisition Subsystem, implemented by the AcquisitionComposite, is responsible for the management and configuration of the acquisition camera and related devices. The composite implements a very simple set of actions and has no specific tracking or control loop functionality.

Figure 9: Acquisition Subsystem Collaboration Diagram

### Interfaces

The following sections detail the input and output interface to the Acquisition Subsystem controller.

#### Inputs

The primary inputs to the Acquisition Composite are implemented as KCSF actions. Actions are defined through an attribute list and may contain parameters. The following define all of the recognized actions available to the WFS composite.

Initialize

The Initialize action is responsible for configuring the subsystem composite, controllers, and devices, and must be executed prior to acquisition and observing. The initialize action will connect the Acquisition Composite to each of the subsystem dependencies and create the required telemetry streams. This includes:

- Acquisition Fold Controller

- Filter Wheel Controller

- Acquisition Focus Stage Controller

- Acquisition Camera Controller

The Acquisition Camera Controller will connect to the low-level camera driver and initialize the device with the default operating parameters. Each of the motion control devices will be initialized and homed as required.

Once the initialization of the subsystem and devices are complete the Acquisition Composite will be ready for observing.

Acquire

The Acquire action is responsible for initiating acquisition imaging. This involves inserting the acquisition fold mirror; configuring the camera based on the seeing and target characteristics (frame-rates, filters, etc.); and adjusting the camera for optimal focus position.

Once the Acquire action is complete the camera will be publishing images to the operator tools, which will also be available to the Controls System through telemetry.

Halt

The Halt action is responsible for halting the acquisition camera and pulling the fold mirror out of the beam-path. Motion controller servo loops will also be opened to halt the devices. The client can issue an Acquire to continue imaging.

Recover

The Recover action is a state dependent process that is responsible for recovering the Acquisition Composite and controllers in the event of a fault. Since the acquisition system is not used during observing, and none of the motion control devices poses tracking or closed-loop functionality, this action can be implemented with functionality identical to Initialize. Recover is defined to conform to the standard NGAO component interface.

#### Outputs

Subsystem output is provided through named telemetry topics. The following describe the telemetry provided by the subsystem, in addition to the standard outputs discussed in 4.1.2.

Image

While the acquisition camera is operating, images captured by the device will be formatted and published.

### Control Loops

The Acquisition Subsystem does not implement any control loops.

### Devices

The Acquisition Subsystem manages the following hardware devices.

#### Acquisition Fold

The acquisition fold mirror is a simple discrete device that is used to pickoff the light from the AO, and feed it to the acquisition camera. This device is controlled as part of the Acquire process and is implemented by a Discrete Controller.

#### Filter Wheel

The filter wheel is a 1-DOF rotational device that is intended to hold an array of camera filters that may be required to properly attenuate the light for the acquisition camera. The filter wheel will use enumerated values to position the desired filter in-beam, and is implemented by a Rotation 1-DOF Controller.

#### Focus Stage

Focus for the acquisition camera is provided through the linear focus stage. This 1-DOF device is positioned based on the seeing conditions and target. Although the focus can be adjusted if conditions change it should not be required frequently and as such, does not require the implementation of a focus control loop. The focus stage is implemented by a Linear 1-DOF Controller.

#### Acquisition Camera

The acquisition camera performs the imaging for the acquisition system. The camera can be configured for the desired observing and target conditions, and will publish images through telemetry and a direct Ethernet connection to a video server. The acquisition camera is implemented by the KCSF Camera Controller class.

## LGS Wavefront Sensors

The Laser Guide Star (LGS) Wavefront Sensors (WFS) are used by the RTC to determine the appropriate corrections to apply to all deformable mirrors (LODM, HODM, and LOWFS) to correct for atmospheric turbulence.

There are a total of seven LGS WFS: four are used to perform tomographic correction; the remaining three are used to perform localized corrections for each of the three tip-tilt stars. The four tomographic WFS are stationary within the LGS assembly, and are used by the RTC to calculate the three dimensional atmospheric profile. The real-time corrections are applied directly to low and high order DMs. The patrolling WFS information is used to correct for atmospheric distortion at each of the tip-tilt NGS stars (3), as seen by the LOWFS. The DMs for each tip-tilt star are located within the LOWFS assembly.

The patrolling WFS each have their own 2-DoF pickoff arm that allows the LGS subsystem to acquire the patrolling LGS beacons from anywhere within the field. As the LGS light path is also separate from the science light path the only positional restrictions for the pickoffs are that they must not obstruct one another or the tomography WFS.  A caveat to this is that the BGS patrolling LGS steering arms are restricted from entering the central 30 arcsecond diameter of the field due to the optics that form the central asterism. As a result, the LGS WFS pickoffs will not normally be positioned within the narrow field area of regard.

### Description

The LGS Wavefront Sensor Subsystem is responsible for the management and control of the LGS wavefront sensors and related hardware. This includes the WFS pickoffs and LGS focus stage.

Figure 10: LGS WFS Collaboration Diagram

The LGS WFS Subsystem provides the initialization and acquisition functionality to position the patrolling pickoffs; align the patrolling and tomography LGS beacons; configure the WFS RTC cameras; and manage the focus for the entire LGS assembly. The LGS WFS Subsystem also provides capabilities to perform background calibrations and respond to alignment errors by offloading to the laser facility global tip-tilt and asterism rotator.

### Interfaces

The following sections detail the input and output interface to the LGS WFS subsystem controller.

#### Inputs

The primary inputs to the LGS WFS composite are implemented as KCSF actions. Actions are defined through an attribute list and may contain parameters. The following define all of the recognized actions available to the WFS composite.

Initialize

The Initialize action is responsible for configuring the subsystem composite, controllers, and devices, and must be executed prior to acquisition and observing. The initialize action will connect the LGS WFS Subsystem to each of the subsystem dependencies, and system inputs. This includes:

- Laser System Composite

- Tomography WFS Controller

- Patrolling WFS Controllers

- LGS WFS Focus Stage

- AO Rotator Controller

- DCS (Channel Access Service)

Each of the WFS controllers will in turn connect to the NGAO RTC Interface and their offload dependencies; power on the WFS devices; and load the default camera properties. The Patrolling WFS Controllers will also home their associated pickoff arms. The WFS Focus stage controller will connect to its input sources and home the focus stage.

Once the initialization of the subsystem and devices are complete the LGS WFS Composite will be ready for observing.

Slew

The Slew action is responsible for positioning the LGS WFS pickoff arms in preparation for acquisition. This action will calculate the required position of the pickoff arms, ensuring that all targets are within the permitted field of regard and do not obstruct each other or the tomography sensors.

The WFS Composite will be provided with the telescope pointing target (optical axis) and the positions of up to three laser beacons. These targets will be converted to image plane coordinates taking into account telescope and rotator orientation to determine the N and E vectors after the rotator. This information will be used to determine the orientation of the targets in the field and the corresponding positions of the pickoffs. The composite will then command each patrolling WFS to acquire its associated target. The mapping between the targets and WFS will be based on the relationship between the LGS and LOWFS, specifically the TTFA target pair (beacon and NGS).

If the target position can not be achieved the action will return a failure status. If only a subset of the patrolling targets are used (including zero), the unused pickoff arms will be safely stowed outside of the field.

Background

The Background action is responsible for initiating a WFS background calibration. The WFS Composite will open the RTC control loops (if currently tracking); tune the lasers off of the 589nm wavelength; initiate WFS background calibrations and tune the lasers back on.

This action can be performed at any time post-initialization, and can be performed as a system wide performance optimization during observing (this will interrupt the observing sequence). If performed during an observation the WFS RTC control loops will have to be closed through a call to Acquire when complete.

Acquire

The Acquire action is responsible for aligning the LGS asterism with the WFS, and closing the AO and offload control loops. There are a total of seven WFS: four tomography sensors managed under the TomographyWFSController; and three patrolling sensors, each managed by a PatrollingWFSController. Alignment of the LGS beacons will be performed with the WFS as the reference, by commanding the BGS tip-tilt, rotator and steering arms. The WFS were selected as the preferred reference because the pickoff devices should be capable of being positioned with higher precision than the patrolling beacons, and will not be affected by telescope flexure – unlike the BGS which is mounted on the telescope secondary socket.

After the devices have been positioned through the execution of the Slew action, the device controllers will be commanded to acquire their respective laser beacons. This will involve reading back flux information from the WFS, and adjusting the position of the BGS devices to align the LGS beacon. For the tomography sensors this will involve a combination of asterism tip-tilt and rotator adjustments. For the patrolling WFS the acquisition will be performed by adjusting their associated steering arm. As the asterism tip-tilt and rotator impact the entire asterism acquisition and alignment will start with the tomography sensors, followed by each patrolling WFS.

If the beacons are not visible during the acquisition process a spiral search will be performed to find the missing beacons using the WFS respective alignment devices.  The spiral search for the tomography sensors will be performed by moving the asterism tip-tilt stage, and will have to balance the offsets across all four WFS. The patrolling WFS will utilize the NGS steering arms, and will be restricted from crossing over or passing behind the other hardware in the asterism generator.

If the acquisition is successful, the composite will enter the track state and close the RTC and offloading control loops.

Halt

The Halt action is responsible for opening the control loops and halting the WFS. The servo power loop for motion devices such as the pickoff arms and focus stage will opened; and RTC will be commanded to open the AO control loop and idle the sensors. The targets can be reacquired and observing can continue with a call to Acquire.

Recover

The Recover action is a state dependent process that is responsible for recovering the WFS Composite and controllers in the event of a fault. During observing if the pickoff arms, focus stage, or WFS sensors enter a fault state observing may be interrupted. Rather than requiring a full subsystem initialization to clear the fault – consequently ending the science sequence – a recover can be issued, which will selectively reinitialize the faulted devices only. If the fault can be cleared the device(s) will be put back into its previous state prior to the fault, and observing should be able to continue, uninterrupted.

#### Outputs

Subsystem output is provided through named telemetry topics. The following describe the telemetry provided by the subsystem, in addition to the standard outputs discussed in 4.1.2.

Targets

The Targets telemetry item contains the sky information used to calculate the position of the patrolling WFS.

TomographyOffload

The TomographyOffload telemetry item is updated during offloads to the BGS Asterism Tip-Tilt and Rotator for the central asterism tip-tilt. This item contains the latest offload applied to these devices and can be used to track total offloads for the control loop.

PatrollingOffload

The PatrollingOffload telemetry item is updated during offloads to the BGS patrolling steering arms. Each patrolling offload task will have its own offload telemetry item.

Focus

The Focus telemetry item defines the calculated focus position for the WFS Focus Stage. This value is updated whenever the focus is recalculated.

### Control Loops

The following sections detail the offload control loops of the LGS WFS subsystem. Each control loop is implemented in part of in full by one or more of the NGAO controllers.

#### WFS Focus

As a tracking device, the WFS Focus Stage Controller will be monitoring the state of the system and responding to events. Focus information is derived from the LOWFS calculated focus and the telescope zenith angle.

Figure 11: WFS Focus Control Loop

WFS focus tracking is an event driven control loop triggered by telemetry events from the RTC, the TWFS camera, and DCS telescope zenith angle keywords. The RTC provides time-averaged data from the Tip-Tilt Focus and Astigmatism (TTFA) WFS to determine the error in the LGS WFS focus position. More accurate focus information is provided at a slower rate (~1Hz) directly by the Truth Wavefront Sensor (TWFS). The TWFS images will be processed by the NGAO controls system to determine focus and centroid offsets and published as telemetry. When the controller receives input from these items it will recalculate the sodium layer altitude and apply the appropriate compensation to the stage.

As the focus position is driven by measurements calculated from LGS beacon information, disturbances in seeing, hardware obstructions, or device failures can cause the control loop to receive erroneous data from the WFS. To ensure the focus is tracking a real target and not noise or reflections, the controls system will monitor the sensor data quality. This may involves periodic checks of flux intensity or the rate of focus degradation. Extreme changes in the monitored properties can be used to automatically stop focus correction until the seeing improves.

#### LGS Offloading

During observing the LGS WFS will periodically have to offload tip-tilt errors to the Beam Generation System to maintain alignment of the LGS beacons. Two types of offload control loops are implemented: one for the tomography sensors, and another for the patrolling WFS. The offload control loop has been divided in this way based on the specific offloading hardware. As with the LGS WFS focus control loop, data quality monitoring will be performed to ensure the offloads are genuine and not a product of tracking noise or reflections.

Tomography WFS

The Tomography WFS Controller uses a single offload task to manage the tip-tilt offload for all four WFS. Offloads are performed when the average of all the tomography tip-tilt stages exceeds a specified threshold. This task is responsible for maintaining a balanced tip-tilt average across the asterism.

Figure 12: Tomography LGS WFS Offloading

Tip-tilt information from each of the WFS TT mirrors is provided through the RTC telemetry, and is used to calculate average tip-tilt across the tomography sensors. The offload task will determine the offset to apply to the BGS up-link Tip-Tilt and Rotator for asterism shift and rotation, respectively. The BGS offsets will be seen by the WFS as rapid motion away from the increasing tip-tilt. This will result in the RTC compensating by moving the WFS TT mirrors towards their zero point.

To ensure that offloads are valid the control loop will regularly monitor the flux as seen by the tomography WFS. If any of the tomography sensors loose their corresponding beacon the offloading task will automatically compensate by removing tip-tilt readings for the errant beacon. If the flux drops for all WFS the offload control loop will be suspended.

Patrolling WFS

Each of the patrolling WFS Controllers implements their own tip-tilt offload control loop. This loop is responsible for periodically offloading the accumulated tip-tilt of the WFS TT mirror (the mirror itself is controlled by RTC) to the BGS patrolling steering arm.

Figure 13: Patrolling LGS WFS Offloading

When the RTC WFS loop is closed the real-time system will be using the WFS tip-tilt mirror to keep the LGS spot stable on the sensor. Over time the tip-tilt will build up and if not offloaded can cause the TT mirror to reach its limit, and eventually loose the LGS beacon all together. When the tip-tilt exceeds a defined threshold, the LGS WFS Controller will calculate the offset required by the steering arm to bring the beacon closer to true center (zero tip-tilt). The WFS will see this as the beacon moving rapidly out of center in the opposite direction of drift, and will compensate by moving the TT mirror to bring the spot back to center. This should consequently return the TT mirror to near its center of range of motion.

As with the tomography WFS offload control loop, each patrolling offload loop will monitor the flux of its associated WFS. If the flux drops below a specified threshold the offload loop will be suspended.

### Devices

The following devices are managed by the LGS WFS subsystem. Specific details, interfaces, and discussions will be available in related design documents.

#### Focus Stage

All seven LGS WFS are mounted on a rigid assembly. Focus for the wavefront sensors are achieved by moving the entire assembly along the z-axis with a linear 1-DOF focus stage. The required position of the stage is based on the telescope zenith angle, and focus measurements determined by the TTFA LOWFS and, at a slower rate, by the TWFS. The TTFA measurements are produced by RTC, while the TWFS focus control loop is implemented directly by the NGAO Controls system.

#### Pickoff Arms

The three patrolling LGS wavefront sensors utilize a 2-DOF rotational pickoff arm to acquire their LGS beacons from the field. The pickoff arms are based on a theta-phi motion control design, allowing each WFS to pickoff a target anywhere in the field. Although the three pickoff arms are stacked along the optical axis to prevent physical collision with each other, motion control and positioning requirements exist to prevent vignetting of the neighboring sensors and ensure the safety of the system.

#### LGS WFS

Although the seven LGS WFS perform different functions in the NGAO system, they share identical hardware and similar control interfaces. The wavefront sensors are under the direct control of the RTC subsystem. The NGAO Controls system is responsible for initializing and configuring the WFS cameras, and commanding RTC to open or close the control loops.

All four tomography sensors are managed by TomographyWFSController. As there are no operating modes where only a portion of the tomography WFS would be active, the Controller exposes a single abstract interface to command all four sensors simultaneously.

The three patrolling wavefront sensors are each represented by an instance of the PatrollingWFSController. The Patrolling Controller interface is similar to the tomography controller, but adds additional functionality for slewing the pickoff arms and acquiring the patrolling LGS beacons.

## Low Order Wavefront Sensors

To correct for atmospheric turbulence you need to know how the light is distorted as it travels through the atmosphere. The LGS WFS do not measure atmospheric tip-tilt and cannot distinguish between focus changes and changes in the altitude of the sodium layer. For this, natural guide stars are needed to calculate the tip-tilt caused by the atmosphere, and the required focus of the AO system. This information and functionality is provided by the low-order wavefront sensor (LOWFS) system.

### Description

The Low Order Wavefront Sensor subsystem is responsible for the management and configuration of the low order wavefront sensors (LOWFS) and related hardware. The LOWFS are responsible for providing tip-tilt and focus information to the RTC and AO Controls system. RTC will apply corrections to the low-order relay TTM1 to correct for full-field tip and tilt. Focus and rotation offloading will be managed by the AO Controls and will be applied to the LGS focus stage and AO rotator.

Figure 14: Low-Order WFS Collaboration Diagram

There are a total of three LOWFS assemblies, each with a dedicated deformable mirror, pickoff arm, focus stage, tip-tilt mirror and sensor. In addition to tip and tilt, one of the LOWFS sensors also provides focus and astigmatism information (TTFA) to the RTC. Co-mounted with the TTFA is a truth wavefront sensor (TWFS). This sensor provides more accurate focus information, but at a slower rate than the TTFA. The TWFS sensor is under the direct management and operation of the AO Control system.

All three LOWFS and the TWFS are managed by the LOWFSComposite. This NGAO component provides a simple interface to initialize, acquire, and manage the devices during observing. The following sections detail the LOWFS Composite.

### Interfaces

The following sections detail the input and output interface to the LOWFS subsystem controller.

#### Inputs

The primary inputs to the LOWFS composite are implemented as KCSF actions. Actions are defined through an attribute list and may contain parameters. The following define all of the recognized actions available to the LOWFS composite.

Initialize

The Initialize action is responsible for configuring the subsystem composite, controllers, and devices, and must be executed prior to acquisition and observing. The initialize action will connect the LOWFS Subsystem to each of the subsystem dependencies, and system inputs. This includes the LOWFS and TWFS controllers.

Each of the LOWFS controllers will in turn connect to the NGAO RTC Interface and dedicated focus stage and pickoff arm controllers; power on the WFS devices; and load the default camera properties. The TWFS Controller will connect to the RTC Interface to provide it with the slow-rate focus information and create the telemetry streams to be used by the LGS focus controller.

Once the initialization of the subsystem and devices are complete the LOWFS Composite will be ready for observing.

Slew

The Slew action is responsible for positioning the LOWFS pickoff arms in preparation for acquisition. This action will calculate the required position of the pickoff arms, ensuring that all targets are within the permitted field of regard, and do not obstruct each other, or the science field. Although the pickoff arms can be placed anywhere within the field, they will typically not be positioned much closer than the wide-field cutoff (~30 arcsecond diameter) due to the restrictions of the LGS patrolling system (each NGS target needs to be LGS corrected) and obstruction of the science field.

The LOWFS pickoff position will be calculated based on the N and E vectors after the rotator and the affects of DAR on the NGS field position. If the targets are not reachable the action will return a failure status. If only a subset of the patrolling targets are used (including zero), the unused pickoff arms will be safely stowed outside of the field.

The LOWFS Composite will be provided with the telescope pointing target (optical axis) and the positions of up to three NGS targets. These targets will be converted to image plane coordinates taking into account telescope and rotator orientation to determine the N and E vectors after the rotator. This information will be used to determine the orientation of the targets in the field and the corresponding positions of the pickoffs. The composite will then command each LOWFS pickoff to acquire its associated target. The primary requirement for this process is that the NGS TTFA target must be obtained by the TTFA WFS.

Acquire

The Acquire action is responsible for aligning the NGS tip-tilt targets with the LOWFS, and closing the AO control loops. There are a total of three LOWFS: two tip-tilts and one TTFA. All three LOWFS controllers share the same acquisition sequence and are treated identically from the perspective of the LOWFS Composite.

After the initial positioning of the devices through the Slew action, the controllers will be commanded to acquire their respective stars. The LOWFS controllers will read back the centroid position from the sensors and reposition the pickoff to center the NGS. If the target is not visible a spiral search will be initiated to find the target. When the targets are acquired a focus calibration will be performed followed by a background. The controllers will finally be put into track to close the RTC control loops.

Background

The Background action is responsible for initiating a sequential background calibration of all three LOWFS, and can only be performed after the LOWFS have been acquired. Essentially, this action exists to perform servo loop optimizations while observing without interrupting the science sequence. The following actions are performed for each LOWFS controller:

- Issue a Background command. This will open the control loops, offset the pickoff arm, take a background over the region of interest, and reacquire the NGS.

- Issue a Track command. This will close the LOWFS control loops.

This action can be performed at any time post-acquisition, and is typically performed as a system wide performance optimization during observing. If new frame rates are required they should be specified at the LOWFS Controller level before initiating a background. If only one or two LOWFS need to be recalibrated the background should be done individually, rather than being performed from the LOWFS Composite level. Acquiring a background during observing should be done sequentially to avoid having to interrupt the observing sequence.

Focus

The Focus action is responsible for initiating the LOWFS calculations to determine and apply the required focus compensation for each of the sensors, conjugate to the science instrument focal plane. The actual focus position is based on the position of the TT target in the field; optics in the science path; and the properties of the science instrument. For each new acquisition or hardware / system configuration change (such as inserting or removing optics, filters, etc.) the LOWFS focus should be recalculated.

Halt

The Halt action is responsible for opening the control loops and halting the LOWFS. Motion devices such as the pickoff arms and focus stages will be powered down and RTC will be told to open the AO control loop and idle the sensors. The devices and control loops can be reacquired with a call to Acquire.

Recover

The Recover action is a state dependent process that is responsible for recovering the LOWFS Composite and controllers in the event of a fault. During observing if the pickoff arms, focus stages, or LOWFS sensors enter a fault state observing may be interrupted. Rather than requiring a full subsystem initialization to clear the fault – consequently ending the science sequence – a recover can be issued, which will selectively reinitialize the faulted devices only. If the fault can be cleared the device(s) will be put back into its previous state prior to the fault, and observing should be able to continue, uninterrupted.

#### Outputs

Subsystem output is provided through named telemetry topics. The following describe the telemetry provided by the subsystem, in addition to the standard outputs discussed in 4.1.2.

Targets

The Targets telemetry item contains the sky information used to calculate the position of each LOWFS.

Focus

The Focus telemetry item defines the focus position for the LOWFS stage. Each LOWFS will have its own focus item.

TruthFocus

The TruthFocus telemetry item defines the focus error calculated by the TWFS. This value is used by the LGS WFS Focus Stage to correct for apparent laser focus.

### Control Loops

The following detail the control loops managed by the LOWFS subsystem.

#### TWFS Focus and Centroiding

The TWFS Controller will be responsible for processing images to determine high accuracy focus and centroiding errors. A Shack-Hartmann technique is used to estimate the wavefront and calculate the overall focus of the light and its average centroid. The focus information will be published to the LGS WFS focus stage, and centroid information will be sent to RTC to improve AO corrections.

If the controller detects that the NGS target has been lost or the flux is too low to track against, the focus and centroid calculations will be suspended until seeing improves.

### Devices

The following details all of the devices managed by the LOWFS subsystem.

#### LOWFS Pickoffs

Each LOWFS sensor utilizes a 2-DOF rotational pickoff arm to acquire the NGS targets from the field. The pickoff arms are based on a theta-phi motion control design, allowing each LOWFS to pickoff a target anywhere in the field. Although the three pickoff arms are stacked along the optical axis because they are on independent focus stages they could physically collide during a focus readjustment or target acquisition. These three degrees of freedom will need to be considered when moving the LOWFS. Additional motion control and positioning requirements exist to prevent vignetting of the science field and ensure the safety of the system.

#### Focus Stages

Each LOWFS unit is mounted onto a 1-DOF linear focus stage. Focus for each LOWFS must be conjugate to the science instrument and is affected by the position of the TT target in the field; optics in the science path; and the configuration of the instrument itself. As the focus for the science instrument shouldn’t change during observing the LOWFS focus position should only need to be set at the beginning of a sequence. If any hardware or optics are added, removed, or reconfigured in the beam train that may affect the focus the LOWFS focus position will need to be recalculated and applied.

#### LOWFS

The LOWFS sensors are under the direct control of the RTC subsystem. The NGAO Controls system is responsible for initializing and configuring the LOWFS cameras, and commanding RTC to open or close the control loops. Each of the LOWFS sensors are managed by instances of the LOWFSController

#### TWFS Camera

Unlike the other LOWFS sensors, the TWFS is under the direct control and management of the NGAO Control system. The TWFS Controller implements the functional control logic for the sensor, and exposes a simple interface to initialize, configure, and control TWFS image processing. The controller is responsible for providing the slow-rate focus and centroid calculations to RTC to optimize system performance.

## Narrow Field Relay

The Narrow Field Relay subsystem covers the devices responsible for applying the high order corrections to the science field. These include the MEMS DM and high-order TT stage. The DM is mounted to the TT stage, and together they control the propagation of the field to the NGS WFS and science instrument. The TT stage is used to perform slow-rate TT / field alignment corrections, and can also be used to perform dithering of the science field.

### Description

There are two devices that comprise the Narrow Field Relay: the high-order DM and TT mirror. The DM is under the direct control and management of the RTC, and is used to correct for the remaining atmospheric distortion that could not be removed by the Wide-Field relay DM. The high-order tip-tilt device (TTM2) is under the direct control of the NGAO controls system, and acts as the mount for the MEMS DM.

Figure 15: Narrow Field Relay Collaboration Diagram

As a low-bandwidth device, TTM2 is used primarily for tracking of differential atmospheric refraction (DAR); correcting ADC tip-tilt errors; and offloading for the Keck Angle Tracker (KAT). In addition, TTM2 will be able to execute science dithering routines, either as scripts or built in sequences, to move the field in a coordinated fashion.

### Control Loops

The following identifies the control loops defined for the Narrow Field Subsystem.

#### ADC Tip-Tilt Errors

It is expected that the Atmospheric Dispersion Compensator (ADC) will cause position translation errors in the science image. Fortunately these errors can be compensated for by the High-Order Tip-Tilt device.

Figure 16: ADC Tip-Tilt Error Correction Control Loop

The High-Order Tip-Tilt Controller will listen to position events from the ADC (as orientation of the corrector stages), and using a table of premeasured tip-tilt error values find the appropriate compensation to apply to the stage.

#### KAT Offloads

During IF observing the Keck Angle Tracker (KAT) will need to periodically offload target centroids while tracking. The TTM2 mirror will correct for KAT offloads.

Figure 17: ADC Tip-Tilt Error Correction Control Loop

The High-Order Tip-Tilt Controller will listen to offload events from KAT as an X / Y field offset. The controller will convert these offsets into the appropriate tip-tilt offset which will be corrected up-stream by the NGS WFS.

### Devices

The following devices are managed under the Narrow Field Subsystem.

#### High-Order Tip-Tilt stage

The High-Order Tip-Tilt stage is a 2-DOF tip-tilt device and is under the direct control and management of the NGAO Controls System. The stage is implemented by the HighOrderTipTiltController.

#### High-Order Relay DM

The high-order relay DM (HODM) is managed and controlled directly by the RTC system. The NGAO Controls system will be responsible for powering on the device; loading reconstructors; setting properties; and initiating RTC servo loop commands. The NGAO Controls System represents the device through the HighOrderDMController: effectively a wrapper for RTC HODM control.

## NGS Wavefront Sensor

Just after the narrow field relay is the natural guide star WFS system. This system is responsible for calculating the wavefront, tip-tilt, and focus of a guide star in NGS mode, and applying the corrections to the low and high-order relays (depending on instrument). The NGS subsystem can also provide truth and focus information to the Real-Time Control system during laser guide star IF and fixed-pupil instrument observing.

### Description

The NGS WFS Subsystem is implemented by the NGSWFSComposite, and is responsible for the management and operation of the natural guide star (NGS) devices. This includes the NGS WFS, focus stage, lenslet array, and field steering mirrors. The Composite presents a simple high level interface to the client to initialize, configure, and acquire the guide star from the field. The NGS WFS Composite is managed directly by the AO Sequencer subsystem.

Figure 18: NGS WFS Subsystem Collaboration Diagram

The NGS WFS is under the direct control of the real-time system. The NGAO controls system is responsible for the initialization and configuration of the sensor, and commanding RTC to open and close the WFS control loops. The other NGS devices are all under the direct control of the Controls System. Outside of the RTC NGS control loops, the NGS Subsystem does not implement any tracking or offload tasks. The majority of device control is performed during configuration and acquisition.

### Interfaces

The following sections detail the input and output interface to the NGS WFS Subsystem.

#### Inputs

The primary inputs to the NGS WFS Composite are implemented as KCSF actions. Actions are defined through an attribute list and may contain parameters. The following define all of the recognized actions available to the NGS WFS Composite.

Initialize

The Initialize action is responsible for configuring the subsystem composite, controllers, and devices, and must be executed prior to acquisition and observing. The initialize action will connect the NGS WFS Subsystem to each of the subsystem dependencies, and system inputs. This includes:

- Field Steering Mirror Controller

- NGS Dichroic

- WFS Lenslet Array

- Camera Focus Stage

- Assembly Focus Stage

- NGS WFS Controller

The NGS WFS Controller will in turn connect to the NGAO RTC Interface; power on the WFS devices; and load the default camera properties. The FSM, focus and lenslet stages will be home and positioned appropriately. Additionally, the FSM Controller will connect to TTM2 to receive dither offset inputs (see Tip-Tilt Compensation control loop.)

Once the initialization of the subsystem and devices are complete the NGS WFS Composite will be ready for observing.

Focus

The Focus action is responsible for determining and applying the required focus compensation for the NGS WFS sensor conjugate to the science instrument focal plane. The actual focus position is based on the optics in the science path, and the properties of the science instrument. Any hardware / system configuration change (such as inserting or removing optics, filters, etc.) will require the focus for the NGS WFS focus to be recalculated and applied.

Acquire

The Acquire action is responsible for placing the fold dichroic in-beam and positioning the field steering mirrors to acquire the specified NGS target. The Acquire action accepts a set of coordinates for a target in the field, and will compute the relative position of the target with respect to the optical axis, the current offset applied to the High-Order Tip-Tilt mirror, and the affects of DAR. This information will be sent to the FSM to position the target on the WFS. After peak-up, the RTC control loops will be closed on the target.

Halt

The Halt action is responsible for halting the motion control devices and idling the camera. All NGS WFS RTC servo control loops will be opened. Halt can be issued from any of the post-initialization states.

Recover

The Recover action is a state dependent process that is responsible for recovering the NGS WFS Controller in the event of a fault. During observing if the sensor enters a fault state observing may be interrupted. Rather than requiring a full subsystem initialization to clear the fault – consequently ending the science sequence – a recover can be issued, which will selectively reinitialize the faulted device. If the fault can be cleared the device will be put back into its previous state prior to the fault, and observing should be able to continue, uninterrupted.

#### Outputs

Subsystem output is provided through named telemetry topics. The following describe the telemetry provided by the subsystem, in addition to the standard outputs discussed in 4.1.2.

Targets

The Targets telemetry item contains the sky information used to calculate the position of the field steering mirrors.

Focus

The Focus telemetry item defines the combined focus position of the NGS assembly and camera stages.

### Control Loops

The NGS WFS subsystem does not currently implement any servo control loops.

### Devices

The following details all of the devices managed by the NGS WFS subsystem.

#### Dichroic

The NGS dichroic is a simple linear 1-DOF device that is responsible for splitting the NGS WFS visible light from the science field to the sensor. During NGS and LGS inteferometry the dichroic will be on-beam. For LGS with an instrument the WFS are not used, and the dichroic will be positioned off-beam.

#### Camera Focus Stage

The NGS detector focus stage is a linear 1-DOF device that is responsible for maintaining the focus of the NGS WFS. The focus of the WFS is based on the properties of the system and instrument. Once the system is configured the focus should not have to be updated. If any new optics are inserted or removed from the beam path, or the instrument properties are changed, the focus may have to be recalculated.

#### Assembly Focus Stage

The NGS assembly focus stage is a linear 1-DOF device that is responsible for maintaining the focus of the image to the lenslet mount and optics. The focus is based on the properties of the system and instrument. Once the system is configured the pre-detector focus should not have to be updated. If any new optics are inserted or removed from the beam path, or the instrument properties are changed, the focus may have to be recalculated.

#### Field Steering Mirrors

The field steering mirrors are a pair of coordinate 2-DOF tip-tilt fold mirrors designed to shift the field to change the apparent view of the WFS NGS. This is the reverse concept of using a pickoff arm to select a point in the field: instead you move the field to acquire the target on the WFS. The FSMs will be able to shift the field to acquire any target in the 40” x 60” science field. Offsets may be applied to the FSMs during a science dither (open-loop) from the MSCS or Observing Tools.

#### Lenslet Array

The lenslet array is a linear 2-DOF device used in producing Shack-Hartmann wavefront calculations. This WFS system allows the RTC to determine focus, tip-tilt, and astigmatism for the natural guide stars. The lenslet array will be positioned based on pre-observing calibration tests determined to produce an optimal flat wavefront, with minimal focus and tip-tilt errors.

#### NGS WFS

The NGS WFS is implemented by the NGSWFSController and provides an interface to the RTC NGS wavefront sensor. The controller provides all of the functionality necessary to initialize, configure, and acquire the WFS. The NGS WFS Composite will be responsible for managing the NGS WFS controller during regular observing.

## Atmospheric Dispersion Corrector

A well known phenomenon of land based observing is that atmosphere can act as a prism and will cause light to diffract and separate into its component wavelengths. As the incidence angle between the light and atmosphere decreases the effects of the diffraction increases. This will have the effect of changing the apparent position and shape of the targets. To reverse the effects of diffraction a set of prisms are used to converge the wavelengths and correct for shape distortion. This prism assembly is known as the Atmospheric Dispersion Compensator.

### Description

The ADCComposite is responsible for the management and operation of the ADC subsystem. The subsystem comprises the ADC device and in-out stage.

Figure 19: ADC Collaboration Diagram

The design of the ADC hardware has not yet been settled, but it will be based on the control of a pair of rotating optics or a linear prism.

### Interfaces

The following sections detail the input and output interface to the ADC subsystem controller.

#### Inputs

The primary inputs to the ADC composite are implemented as KCSF actions. Actions are defined through an attribute list and may contain parameters. The following define all of the recognized actions available to the ADC composite.

Initialize

The Initialize action is responsible for configuring the subsystem composite, controllers, and devices, and must be executed prior to acquisition and observing. The initialize action will connect the ADC Subsystem to each of the subsystem dependencies, specifically the ADC Controller and ADC Stage. The ADC Controller will in turn connect with its motion control dependencies and control loop inputs.

Once the initialization of the subsystem and devices are complete the ADC Composite will be ready for observing.

Track

The Track action will place the ADC in-beam and put the Controller in the track state (closed loop operation). While in the track state the ADC will be actively responding to system events from DCS and the AO Rotator to correct for the expected atmospheric dispersion. The dispersion correction values will be based on theoretical and empirical tests, and maintained in a look-up table. Additionally, pressure and humidity data may also be used to optimize the positioning of the ADC. The ADC tracking loop will index this table to compute the appropriate correction to apply.

Halt

The Halt action is responsible for halting the ADC and stage controllers, and taking the composite out of the TRACK state. Halt can be issued from any of the post-initialization states.

Recover

The Recover action is a state dependent process that is responsible for recovering the ADC Composite and controllers in the event of a fault. During observing if the ADC device enters the fault state, dispersion affects will begin to increase. Although this will not require that the observing sequence be interrupted, it will continue to degrade performance and data quality as the sequence continues.

In the event of a fault a recover can be issued to selectively re-initialize the device and attempt to clear the error. If the error can be cleared the device will return to its original state before the fault occurred.  If the fault can not be cleared the device will remain in the fault state, but the composite will still be able to receive commands. If the client attempts to command the faulted device through the composite the composite will in turn enter its fault state, requiring a response from the operator.

#### Outputs

This subsystem does not implement any additional outputs outside of the standard telemetry items defined in section 4.1.2.

### Control Loops

The following detail all of the tracking control loops implemented by the ADC subsystem.

#### ADC Tracking

As a tracking device the ADC Controller will be listening to system event and responding automatically to changes. The ADC dispersion corrector control loop takes as input the telescope zenith angle and rotator orientation. (The science instrument wavelength is also needed, but this is set before the start of a sequence.)

Figure 20: ADC Tracking Control Loop

The input values are used to index a table of premeasured dispersion values and combined to determine the appropriate correction to apply with the ADC. The ADC correction loop is event driven and can potentially run at the rate of the fastest telemetry item. Decimation can be used to reduce the correction rate if appropriate.

### Devices

The following lists all of the devices managed by the ADC Subsystem.

#### ADC Stage

The ADC stage is a simple discrete device that will insert or remove the ADC device from the beam path. The position of the stage is implicitly assumed by the current observing mode selection. If the user wishes to manually insert or remove the stage the device controller can be addressed and commanded directly.

#### Atmospheric Dispersion Corrector

The Atmospheric Dispersion Corrector is implemented by the ADCController, and may be implemented as a hybrid controller or 2-DOF rotational controller. The device is designed to control the amount and direction of the dispersion. Although the controller allows clients to explicitly position the device, for the most part this device will simply be run in closed loop tracking mode, automatically applying correction as required.

## RTC Interface

The Real-Time Controller (RTC) is responsible for calculating the atmospheric distortion and applying corrections to the NGAO system by controlling the operation of the deformable mirrors and various tip-tilt stages through feedback from the wavefront sensors. For the most part the RTC system can be seen as a black box by the AO controls system, requiring an initial configuration but only minimal input during observing. The RTC performs the tomographic and atmospheric reconstruction calculations using an array or processors, with throughput of multiple kilohertz rates. The NGAO Controls System interfaces with the real-time system through the RTC Interface Controller.

### Description

The RTC Interface is implemented by the RTCController, and provides all of the controls functionality exposed by RTC Control Processor (CP). The CP module is the server interface to the RTC system and is responsible for receiving and dispatching commands to the RTC control units. The CP receives and send information through an Ethernet based message protocol defined in the RTC Supervisory Controller Interface Document.

Figure 21: RTC Interface Block Diagram

Through the CP, the RTC Interface is able to interact with most of the RTC hardware and is able to execute many low-level device functions.  The CP can also provide diagnostic and telemetry information about the current state and position of devices; WFS and centroiding data; and the results of turbulence and correction calculations. All of the information and functionality provided by the CP will be implemented by the NGAO RTC Interface Controller, and is detailed in the following sections.

### Interfaces

The following details the input and output functionality provided by the RTC Controller.

#### Inputs

The primary inputs to the RTC Interface Controller are implemented as KCSF actions. Actions are defined through an attribute list and may contain parameters. The following define all of the recognized actions available to the RTC Controller.

Initialize

The Initialize action is responsible for connecting the RTC Interface Controller to the real-time control processer; powering on the AO and data storage system; loading the default configuration and properties; and creating the output telemetry streams. Once the initialize action completes the real-time system will be ready for use.

PowerOn

The PowerOn action allows the client to power on a specific device in the RTC control system. This includes:

- LGS WFS

- LOWFS

- NGS WFS

- Low-order DM

- High-Order DM

- Tip-Tilts

The client would specify the device ID which the RTC CP will power on through a networked power controller (e.g. Pulizzi).

PowerOff

The PowerOff action performs the reverse of PowerOn and accepts the same set of devices.

ConfigureDiagnostics

The ConfigureDiagnostics action allows clients to create and modify telemetry streams to be published by the RTC. The client is able to configure the rates of telemetry and apply select filtering mechanisms to the data. The following telemetry streams are available:

- Raw camera data from all WFS,

- Centroids from all WFS,

- Reconstructed wavefront from all HOWFS,

- High-order science wavefront,

- Low-order science wavefront,

- Science on-axis high-order commands,

- Science on-axis low-order commands,

- Device states, configuration and loop status,

- Control Processor and RTC Unit states,

- WFS and camera settings,

- Debug levels

ConfigureSensor

The ConfigureSensor action allows clients to configure a WFS or camera. The following parameters can be set (not all parameters may be valid for a specific camera or sensor):

- Frame rate,

- Frame transfer rate,

- Exposure Time,

- Clock speed,

- Pixel rate,

- Gain,

- Dark pattern,

- Threshold level,

- Centroiding algorithm,

- Centroid position,

- Image pixel offsets,

- Area of interest

ConfigureTipTilt

The ConfigureTipTilt action allows clients to configure the RTC tip-tilt devices. The client would specify the ID of the tip-tilt device and the values for any or the configurable parameters. The following parameters can be set:

- Position,

- Apply a cyclical dither rate to any tip-tilt actuator,

- Zero position

ConfigureDM

The ConfigureDM action allows clients to configure any of the RTC deformable mirrors. The client would specify the ID of the DM and the values for any or the configurable parameters. The following parameters can be set:

- Voltage of any of the DM actuators,

- Bias for any voltage or pattern,

- Apply alignment patterns,

- Cycle one or more actuators as a desired rate,

- Cycle all actuators in a round-robin fashion,

- Apply a control matrix,

- Load a flat pattern,

- Create calibration system matrix,

- Clear any manually set values

CloseLoop

The CloseLoop action allows clients to close any of the available RTC servo loops. When a loop is closed sensor information will be processed and applied to the associated DM and tip-tilt stages. The following control loops have been identified:

- LOWFS DM servo loops,

- Low-Order Woofer DM servo loop,

- LOWFS Tip-Tilt servo loops,

- Low-Order Tip-Tilt servo loop,

- TTFA focus servo loop,

- High-Order Tweeter DM servo loop

OpenLoop

The OpenLoop action allows clients to open and currently closed RTC servo loop. This action accepts the same set of control loop defined by CloseLoop.

Halt

The Halt action is responsible for halting all atmospheric corrections and opening the RTC servo control loops.

Recover

The Recover action is responsible for recovering the RTC hardware in the event of a fault.

Shutdown

The Shutdown action is responsible for shutting down the RTC system and devices at the end of observing. The real-time system can be restored by issuing an Initialize.

#### Outputs

In addition to the standard outputs discussed in 4.1.2, the RTC Interface defines an output telemetry item for every property and value that can be read from RTC. Nearly every item that can be set in an action has an associated output telemetry item. For brevity they will not be listed here. See the previous section for a summary of the properties available as telemetry, or the RTC Control Processor Interface Design Document.

### Control Loops

The RTC Interface Controller does not implement any control loops.

### Devices

The RTC Interface Controller does not directly control the operation of the real-time devices, but through the real-time Control Processor allows clients to configure a number of these devices arbitrarily. The following devices are controlled through the RTC Interface Controller

#### Deformable Mirrors

There are a total of five deformable mirrors within the NGAO system: woofer, tweeter and three mems DMs found in the LOWFS. Each of the DMs is built with a set of actuators to control the shape of the mirror surface and correct for disturbances in the wavefront. Actuator control is issued in the 2-kilohertz range by the RTC.

#### Tip-Tilt Stage

There are a total of 11 tip-tilt stages under the control of the RTC: seven LGS WFS tip-tilts; three LOWFS tip-tilts; and the on-axis low-order tip-tilt (TTM1). The WFS tip-tilts are responsible for keeping the LGS and NGS targets centered on the respective sensors. Tip-tilt information from the LOWFS is sent to TTM1 by RTC. The LGS WFS tip-tilt is eventually offloaded by the Controls System to the Beam Generation System (non-real time).

#### Cameras and Sensors

There are a total of 11 WFS under the control of the RTC: seven LGS WFS, three LOWFS, and one NGS WFS. Each of the sensors provides the real-time system with data required to calculate the atmospheric distortion including centroiding, focus, and wavefront information. The data from these sensors is sent to the tomography and turbulence correction engines which will calculate the appropriate actuation to apply to the DMs and tip-tilt stages.

#### Power Controllers

Each of the RTC devices is attached to a dedicated power controller, and must be explicitly powered by the AO Controls System prior to use. Once powered, the devices can be used by the RTC in regular observing operation.

## Offloading

During observing a number of tip-tilt and offload events must occur to track targets as the telescope and field moves. Most of the closed-loop tasks are implemented and handled directly by controllers; however some system offloads must be implemented as stand-alone tasks. The Offloading Subsystem is responsible for configuring and starting these offload tasks.

### Description

The Offloading Subsystem manages the software controllers responsible for performing the telescope and Differential Atmospheric Refraction (DAR) offloading. This subsystem is implemented by the OffloadingComposite.

Figure 22: External Offloading Collaboration Diagram

The Offloading Composite manages two stand-alone offload tasks: TelescopeOffloadController and DAROffsetController. These controllers are responsible for calculating and applying offsets to the associated devices.

The Telescope Offload task offloads the focus and tip-tilt build-up (as measured by DM1 and TTM1) to the telescope secondary and pointing, respectively. In open loop mode (during target acquisition) focus and tip-tilt information will be provided by the LOWFS and NGS WFS. The DAR Offset task is responsible for calculating the effects of DAR between the science instrument and WFS (either LOWFS or NGS), while taking into account the other optics and filters in the beamtrain. These offsets are applied to the High-Order Tip-Tilt stage, or as centroid offsets to RTC. Both tasks are event driven; where the telescope offload is based on focus and tip-tilt thresholds; and the DAR task is based on changes in telescope elevation.

### Interfaces

The following sections detail the input and output interface to the External Offloading subsystem controller.

#### Inputs

The primary inputs to the External Offloading subsystem are implemented as KCSF actions. Actions are defined through an attribute list and may contain parameters. The following define all of the recognized actions available to the Offloading subsystem.

Initialize

The Initialize action is responsible for configuring the subsystem composite and offload controllers, and must be executed prior to acquisition and observing. The initialize action will connect the Offloading Subsystem to the Telescope and DAR offload controllers. The controllers will in turn establish peer-to-peer and telemetry connections with their associated dependencies.

Once the initialization of the subsystem is complete the offloading tasks will be ready for observing.

SetMode

The AO Control System will be receiving offsets and offloads from a number of sources and is responsible for calculating the appropriate offsets to apply to specific devices, depending on the observing mode. The SetMode action is responsible for configuring the required input and output streams for the DAR Offset and Telescope Offload Controller. While in track mode, these controllers will read from the input streams and send offloads to the defined output devices.

Track

The Track action is responsible for starting the offloading tasks, and represents the bulk of the subsystems functionality. While tracking, the controller will be listening to events from DAR and RTC. Offsets will be calculated based on these input values, and output to the relevant device and telescope.

The composite and controllers will stay in track until the loop is explicitly opened through a call to Halt.

Halt

The Halt action is responsible for opening the AO Offset and Telescope Offload control loops and taking the subsystem out of the TRACK state. Halt can be issued from any of the post-initialization states.

Recover

The Recover action is designed to selectively reinitialize a component in the event of a system fault during observing. As the External Offload Subsystem is comprised of simple software controllers, and does not directly manage any hardware, a recover has essentially the same functionality as performing an initialization.

#### Outputs

Subsystem output is provided through named telemetry topics. The following describe the telemetry provided by the subsystem, in addition to the standard outputs discussed in 4.1.2.

DAR

The DAR telemetry item defines the direction and magnitude of the calculated atmospheric refraction, as well as the tip-tilt offset or centroid value that is used to compensate.

TelescopeOffload

The TelescopeOffload telemetry item defines the low-order tip-tilt offload sent to DCS to track the target.

FocusOffload

The FocusOffload telemetry item defines the low-order focus offload sent to DCS to control the secondary mirror.

InstrumentWavelength

The InstrumentWavelength telemetry item defines the name and wavelength set for the science instrument. This value is used as the reference which the DAR is calculated against.

### Control Loops

The following control loops are managed by the Offloading Subsystem.

#### Telescope Offload

The TelescopeOffloadController is responsible for offloading tip-tilt, focus, and coma from the low-order relay to the telescope. During observing, focus and field tip-tilt will be corrected by the low-order DM and TTM1 stage, respectively. Periodically these quantities will need to be offloaded to the telescope to ensure the AO devices have sufficient range to correct for the atmosphere. Tip-tilt is offloaded to the telescope AZ / EL position of the telescope, and focus / coma is offloaded to the secondary piston and tip-tilt.

Figure 23: Telescope Offload Control Loop

The Track action is responsible for starting the telescope offload task, and represents the bulk of the controller’s functionality. While tracking, the controller will be listening to events from the low-order relay DM and Tip-Tilt Mirror (TTM). If the controller determines that the focus applied to the DM or the position of the TTM exceed configurable thresholds, the controller will take action to offload the buildup to the telescope.

#### DAR Offsets

The DAR Offset controller is responsible for calculating and compensating for the affects of DAR on the NGAO system. DAR is a measure of the refraction of light between two wavelengths caused by atmosphere at different elevations. From the perspective of the WFS it will appear as if the target is slightly offset from its expected position due to the wavelength of light the sensor can see. If the diffraction effects are not corrected the WFS alignment will suffer and the science target will not be on-axis for the science instrument. (See KAON 221 for more information on Differential Atmospheric Refraction.)

Figure 24: DAR Offset Control Loop

DAR is a relative calculation based on the wavelength of the science instrument, tip-tilt WFS (either LOWFS or NGS WFS), and the current telescope elevation and vertical angle. In NGS mode or LGS with the Interferometer the DAR calculation is performed relative to the NGS WFS, and is applied to the RTC as centroid offsets. During LGS with an instrument the DAR is calculated relative to the LOWFS wavelength and is applied directly to TTM2. The DAR calculation is triggered whenever the telescope elevation / vertical angle changes.

The DAR Offset Controller is not responsible for initializing, configuring or acquiring the device dependencies. The Offset Controller assumes the devices are in the proper operating state and capable of accepting offsets when tracking.

### Devices

Although the offload control subsystems interacts with a number of devices they do not actively manage their life-cycle or state. See the relevant subsystems for more information about the offload devices.

- High-Order Relay

- LOWFS

- RTC Interface

## Calibration Simulation

Periodically the NGAO system will need to be calibrated to obtain the baseline measurements for device and system performance. Calibrations are usually performed once at the beginning of every run, and involve low-level tests of the AO hardware using a simulated star field. The star field and other stimulus sources are generated by the calibration subsystem.

### Description

The Calibration Subsystem is comprised of a number of devices responsible for the generation and propagation of simulated sources including: stimulus devices, tip-tilt fold mirrors, and focus stages. The Calibration System is implemented by the NGAO CalibrationComposite class.

Figure 25: Calibration Subsystem Collaboration Diagram

The calibration subsystem is a loose collection of device controllers. The subsystem is primarily responsible for initializing the devices and catching run-time errors. Control and execution of calibration sequences will be performed by the MSCS or other high-level operator tools, which will interface directly with the devices (see KAON 739 for a summary of the calibration sequences). The following devices and stimulus sources are available:

Motion Control Devices (non-tracking):

- NGS Source Focus Stage

- LGS Source Focus Stage

- NGS Target Selector

- AO Input Fold Stage

- AO Input Fold Mirror Tip-Tilt

- Output Fold Mirror Tip-Tilt

- Pupil Mask Selector

- Astrometric Grid

- Phase Plates (I&T only): Rotation1DOFController

Sources:

- NGS Source (2)

- LGS Source (N)

- Flat-field Source

- Spectral Source

### Interfaces

The primary inputs to the Calibration Composite are implemented as KCSF actions. Actions are defined through an attribute list and may contain parameters. The following define all of the recognized actions available to the Calibration Composite.

#### Inputs

The primary inputs to the Calibration subsystem are implemented as KCSF actions. Actions are defined through an attribute list and may contain parameters. The following define all of the recognized actions available to the Calibration subsystem.

Initialize

The Initialize action is responsible for configuring the subsystem composite and devices, and must be executed prior to acquisition and observing. The initialize action will connect the Calibration Subsystem to each of the subsystem device dependencies, and issue its own initialization for each device. The devices will in turn be configured, and for motion control devices, homed.

Once the initialization of the subsystem and devices are complete the Calibration Subsystem will be ready for use.

Halt

The Halt action is responsible for halting the Calibration Subsystem controllers and devices. Halt can be issued from any of the post-initialization states.

Recover

The Recover action is responsible for recovering the Calibration Composite and controllers in the event of a fault. Recover can be issued to selectively re-initialize the faulted device(s) and attempt to clear the error. If the error can be cleared the device will return to its original state before the fault occurred.  If the fault can not be cleared the device will remain in the fault state, but the composite will still be able to receive commands. The operator will need to take direction action to clear the fault from the affected devices.

#### Outputs

This subsystem does not implement any additional outputs outside of the standard telemetry items defined in section 4.1.2.

### Devices

The following device types are managed by the Calibration Subsystem.

#### Stimulus

A stimulus device is a simple ON / OFF light source, with basic capabilities to change the intensity of the output. The calibration system may utilize up to six sources, although it is possible this number will be reduced based on the need for spares and the ability to share sources for different calibrations.

#### Linear

A number of 1-DOF linear devices will be required by the subsystem to position fold mirrors, select input configurations, and control stimulus focus. None of these devices have any tracking requirements.

#### Rotational

Up to three rotational 1-DOF devices are planned for the Calibration System. The astrometric grid will have a permanent place in the system, but the two phase plates will likely only be used during I&T and annual system calibrations, and will be inserted when needed.

#### Tip-Tilt

Two tip-tilt devices will be used to control the propagation of the source light. These are mounted to the calibration system output fold mirror and the AO input fold mirror. The tip-tilt devices will be positioned to align the source light with the AO optical axis, after which they will remain in a stationary state and will not be used in a motion control capacity for calibrations.

## AO Enclosure

The AO enclosure, which includes the cold-bench, AO room, and electronics vault, contain a number of diagnostic and environment sensors. These sensors will be distributed throughout the AO enclosure to obtain global and localized measurements of environmental conditions. The readings from these sensors will be used primarily to monitor these conditions and ensure the safe operating environment for the AO hardware.

### Description

The AO Enclosure sensors are managed under the AO Enclosure Subsystem.

Figure 26: AO Enclosure Collaboration Diagram

Currently this subsystem does not manage any motion control devices, and as such has a very simple software design and limited functional requirements. Although the sensor information will be freely available to the NGAO Controls System, it is not actively used during regular observing.

Unlike other NGAO control subsystems, the AO Enclosure Subsystem will typically be left running at all times, even during shutdown at the end of an observing night. Information published by the enclosures should always be available so operators and technicians can monitor the well being of the system, even during down time.

### Interfaces

The following sections detail the input and output interface to the AO Environment subsystem controller.

#### Inputs

The primary inputs to the AO Environment Enclosure subsystem are implemented as KCSF actions. Actions are defined through an attribute list and may contain parameters. The following define all of the recognized actions available to the Environment Enclosure subsystem.

Initialize

The Initialize action is responsible for configuring the subsystem composite, controllers, and devices. The initialize action will connect the AO Enclosure Subsystem to each of the subsystem dependencies, specifically the AO Cold Bench, AO Room, and AO Electronics Vault composites. These composites will in turn connect to their device drivers, initialize the sensors, and create the telemetry topics to publish environmental data.

#### Outputs

The AO Enclosure subsystem does not implement any additional outputs outside of the standard telemetry items defined in section 4.1.2

### Control Loops

As AO Environment Subsystem data is not required by the NGAO controls system during AO operation, and does not manage or interact with any motion control devices, no control loops are implemented.

### Devices

The following sensor devices are managed by the AO Environment Subsystem. Each sensor instance will be monitored by the Alarm Handler to provide operators with immediate feedback and alarm notification in case any of the sensors exceed a defined threshold. Thresholds are based on the LOLO to HIHI alarm categorization, and can be defined in the configuration database.

#### Temperature Sensor

The AO Environment Subsystem manages two temperature sensors: one on the AO bench, the other in the AO room enclosure. The temperature sensors will continually report the current temperature of the local environment. Large fluctuations in temperature can indicate failure of one or more cryo units, and run the highest risk in causing damage to the system. High temperatures can also significantly degrading system performance.

#### Humidity Sensor

High humidity in the AO system can cause problems for the electronics, and even more so for the optics and AO performance: this could be especially disastrous for the MEMS if powered up. As there is a variance between the ambient and cryo temperatures, condensation can also form and collect if the humidity in the environment is not controlled.

#### Particulate Sensor

As with the temperature and humidity sensors the AO facility also employs two particulate sensors to monitor the purity of the air in the AO enclosure. A high percentage of dust and debris in the air will reduce the throughput and quality of light in the AO system, and will require more frequent cleanings to bring the system back up to peak performance. A surge in particulates can also indicate smoke from electronic failures or friction between surfaces that may not register a high enough heat to be noticed by the temperature sensors.

## AO Control Sequencer

NGAO is a complex system that will require a controls architecture capable of coordinating and managing the execution of dozens of devices. Each device must be configured and operated in a coordinated and synchronized fashion to ensure the optimal performance and safety of the system. A key to defining this device coordination is through the development of control sequences.

A sequence is a step-by-step breakdown of events and operations that a control system must perform to accomplish some task. Each step in a sequence represents a distinct system command. In many cases a step can itself be divided into sub-steps allowing the developer to design a hierarchical command sequence. These hierarchies are critical for discovering the collaboration and communication between components within the system.

Steps within a sequence can be executed serially, parallel or a combination of the two. A sequence or an operation within a sequence can be defined synchronously or asynchronously, where the execution of the sequence may depend on a specific system event.

The NGAO system can be divided into two main sequences: up-link and down-link. The up-link sequence pertains to the tasks and operations that must be performed by the Laser facility to generate, align, and project the laser asterism on the sky. The down-link sequence defines all of the steps that must be performed by the AO facility to acquire, calibrate, and perform atmospheric correction. The terms down-link and up-link indicate the direction of travel the light takes in and out of the system, respectively.

The following sections detail the implementation of the down-link control sequence for NGAO.

### Description

The AO control sequence is implemented by the AOFacilityComposite, which is the main interface between the various AO controls functions and the rest of the NGAO system.

Figure 27: AO Control Sequencer Collaboration Diagram

The AO Sequencer is responsible for the management and coordinated execution of the entire AO controls system. The Facility Composite implements the high-level device sequences that will perform the system configuration, target acquisition, and closed-loop observing. The Composite executes these sequences by directing commands to each of the dependent subsystems depicted in Figure 27, and discussed in the previous sections.

#### AO Sequence

The AO control sequence covers everything from bringing up the software system to acquisition, and finally shutting down for the night. The sequence itself can be divided into phases each covering a specific state of the system during a typical night’s observing. These phases are (listed in order of execution):

- Startup

- Initialize

- Configure

- Acquire

- Observe

- Halt

- Power-down

For the most part these phases can be executed straight through to perform NGAO corrections for a science target. However there is flexibility allowing the sequence to return to a previous phase or jump forward to a latter phase. The following state diagram describes the relationship between each phase in the AO sequence:

Figure 28: AO Sequencer State Diagram

The main branch of the sequence proceeds through each of the phases as outlined. At specific points in the sequence the control can be sent back to a previous step, such as going from Halt to Acquisition. This flexibility allows the system to be reconfigured multiple times throughout a night for greater flexibility. (Note: this diagram is a simplified version of the actual AO control sequence. The complete sequence employs a more complex state machine, with some details left out here for brevity. The full AO Sequencer state machine is documented in the NGAO AO Facility Component Design Document.)

A detailed review of the AO sequences is described in KAON 735: NGAO AO Facility Sequence.

### Interfaces

The AO Control Sequencer implements a set of high-level tasks that are used to control the entire AO system during standard observing. Additional device control and configuration can be performed by connecting directly to the AO subsystems or devices. The following sections detail the input and output parameters of the AO Facility Composite.

#### Inputs

The primary inputs to the AO Control Sequencer are implemented as KCSF actions. Actions are defined through an attribute list and may contain parameters. The following define all of the recognized actions available to the AO Sequencer.

Initialize

The Initialize action is responsible for configuring the AO sequencer, subsystems, and devices, and must be executed prior to acquisition and observing. The initialize action will connect the AO sequencer to each of the composite’s dependencies, specifically:

- LGS WFS Subsystem

- LOWFS Subsystem

- NGS WFS Subsystem

- ADC Subsystem

- Acquisition Subsystem

- Offloading Subsystem

- RTC Interface Controller

- High-Order DM Controller

- High-Order Tip-Tilt Controller

- IF Fold Dichroic Controller

- Hatch Cover Controller

- Low-Order DM Controller

- Low-Order Tip-Tilt Controller

- Image Rotator Controller

Each of the subsystem dependencies will in turn initialize their subordinate controllers and devices as discussed in the corresponding sections. The individual device controllers will be commanded to connect to their low-level hardware and initialize / home the devices. Any additional communication dependencies (such as tracking loop inputs) will also be established.

In addition to the primary dependencies listed above, at different stages in the observing sequence the AO Control Sequencer will connect to other system devices to perform specific mode-dependent configurations and commands. These operations will be performed outside of the knowledge and control of the parent composites and subsystems. This allows us to separate the mode and instrument dependent functionality from the low-level system and keep it all within the AO Control Sequencer.

Once the initialization of the AO facility is complete the system can be configured for observing.

SetMode

Before observing can begin the operator must configure the system for the desired observing mode. This process commands the control system to configure the hardware and software for observing and may include loading hardware parameters, positioning stages, and changing the operating mode of software controllers.

Configuring the system for a particular observing mode will typically be performed from the top-level MSCS. An operator will select one of the available modes and command the system to configure. This command is passed down to AO Control Sequencer, which will be responsible for determining how the subsystem and low level devices are affected.

The following sections detail the general system and component configuration of each observing mode. System configuration information was derived from KAON 550: NGAO System Configurations

Fixed-Field (Imager or IFS)

Fixed Field (FF) mode will be the most common observing mode of the NGAO system, as required by the majority of science use cases. In this mode the laser and AO system are configured such that the star field remains fixed relative to the science camera, while the pupil can rotate freely. Up to three patrolling LGS beacons may be used.

Fixed field observing is specified by the configuration names FixedField_LGS and FixedField_NGS.

Fixed-Field NGS

In Fixed-Field NGS with an instrument, the laser system and dependent AO hardware will not be used. The primary beam train will consist of the rotator, low-order relay, high-order relay, NGS subsystem and ADC. The following notable configurations will be obtained for this mode:

- Rotator tracking position angle.

- LGS WFS Subsystem stowed:

  - Pickoffs positioned evenly in field.

  - Cameras idle.

  - Focus and offloading loops open.

- LOWFS Subsystem stowed:

  - Pickoffs stowed outside of science field.

  - Cameras idle.

  - Focus loops open.

- NGS WFS Subsystem on beam.

- ADC Subsystem on-beam, tracking dispersion correction.

Fixed-Field LGS

In Fixed-Field LGS the majority of the AO hardware and control software will be in use. The following notable software and hardware configuration will be obtained:

- Rotator tracking position angle.

- LGS WFS subsystem in use.

- LOWFS Subsystem enabled in use.

- NGS WFS Subsystem stowed: off-beam.

- ADC Subsystem on-beam, tracking dispersion correction.

Fixed-Pupil (Imager or IFS)

Fixed Pupil (FP) mode is designed to maintain the structure of the Point Spread Function (PSF) in a fixed location on the science camera, while allowing the star field to rotate freely. This observing mode is ideal for detecting faint sources next to bright companions. Since there is no plan for tracking wide-field tip-tilt in fixed-pupil mode, the patrolling LGS beacons will not be available.

Fixed pupil observing is specified by the configuration names FixedPupil_LGS and FixedPupil_NGS.

Fixed-Pupil NGS

Fixed-pupil NGS observing with an instrument follows an identical configuration as fixed-field, except the image rotator will maintain a fixed pupil orientation in place of a fixed field (e.g. track the vertical angle.)

Fixed-Pupil LGS

Fixed-pupil LGS observing will not use the wide-field beacons, and will therefore have to get the truth information from a different source. The NGS WFS can be put into TTFA / TWFS mode allowing the AO to obtain truth information from the on-axis science target. The following notable software and hardware configuration will be obtained:

- Rotator tracking vertical angle.

- LGS Tomographic WFS in use. Patrolling WFS stowed.

- LOWFS Subsystem stowed

- NGS WFS Subsystem on-beam. Sensor in TTFA / TWFS mode.

- ADC Subsystem on-beam, tracking dispersion correction.

Interferometer

Interferometer (IF) observing can be performed in single target (V2 and NULLING), and dual star (ASTRA) modes. Single target inteferometry can be performed in either fixed field or fixed pupil modes, while dual-star must be performed in fixed field. The configuration for Interferometer observing is similar to instrument observing, but with a slightly reduced hardware set. Since the IF fold dichroic picks off the near-infrared light, the LOWFS sensors do not have enough signal to correct for full-field TT. Additionally, since the fold mirror is between the first and second relays the tweeter DM does not contribute to the corrected light seen by the interferometer, and is put in its optimal flat shape. Also the science instrument ADC will not be used during IF observing so will be idle and stowed off-beam.

The following interferometer NGAO observing configurations will be available:

- FixedField_LGS_IF: Fixed field LGS inteferometry.

- FixedPupil_LGS_IF: Fixed pupil LGS inteferometry.

- FixedField_NGS_IF: Fixed field NGS inteferometry.

- FixedPupil_NGS_IF: Fixed pupil NGS inteferometry.

Fixed-Field NGS

The following configuration will be obtained when observing in Fixed-Field NGS IF:

- Rotator enabled, tracking position angle.

- IF Fold Dichroic on beam.

- LGS WFS Subsystem stowed:

  - Pickoffs positioned evenly in field.

  - Cameras idle.

  - Focus and offloading loops open.

- LOWFS Subsystem stowed:

  - Pickoffs stowed outside of science field.

  - Cameras idle.

  - Focus loops open.

- High-Order DM flat.

- NGS WFS Subsystem on-beam.

Fixed-Field LGS

In Fixed-Field LGS with an instrument, the laser system and dependent AO hardware will be enabled. The majority of the AO hardware and control software will be active in this mode. The following software and hardware configuration will be obtained:

- Rotator enabled, tracking position angle.

- IF Fold Dichroic on beam.

- LGS Tomographic WFS in use. Patrolling WFS stowed.

- LOWFS Subsystem stowed.

- High-Order DM flat.

- NGS WFS Subsystem on-beam. Sensor in TTFA / TWFS mode.

Fixed-Pupil NGS

Fixed-pupil NGS observing with the interferometer follows an identical configuration as fixed-field NGS, except the image rotator will maintain a fixed pupil orientation in place of a fixed field (e.g. track the vertical angle.)

Fixed-Pupil LGS

Fixed-pupil LGS observing with the interferometer follows an identical configuration as fixed-field LGS, except the image rotator will maintain a fixed pupil orientation in place of a fixed field (e.g. track the vertical angle.) Fixed-pupil mode can not be used during dual-star observing (ASTRA).

Acquire

The Acquire action is responsible for initial positioning and acquisition of devices in preparation for observing. This action is typically executed concurrently with the telescope slew. The following steps outline the Acquire sequence, and will vary based on each observing mode (i.e. LGS, NGS, Fixed-Field, and Fixed-Pupil) and instrument.

- Command Image Rotator Controller to obtain desired orientation for field or pupil, and enter track mode.

- Command ADC Composite to obtain required position and enter track mode (instrument observing only).

- Command the High-Order TT controller to center the second relay Tip-Tilt (TTM2). (Alternatively this mirror can be positioned at an observer demanded orientation.)

- Command the RTC Interface Controller to center tip-tilt mirrors and zero devices.

- Command Acquisition Composite to prepare the camera and stages for target acquisition.

- Command LGS WFS Composite to slew to the calculated LGS target positions.

- Command LOWFS Composite to slew to the calculated Tip-Tilt target positions (instrument observing mode only.)

- Command the NGS WFS Composite to slew the FSM to the calculated NGS target position (IF and fixed-pupil modes only.) [?]:

- Configure RTC system (frame-rates, gains, reconstructors, etc.) for observing.

These operations will be executed in parallel by the sequencer and it is expected that they will complete before the telescope pointing corrections are completed. (Acquisition Composite preparation must be complete before the telescope slew ends). Once the Acquire action is complete the wavefront sensors can be commanded to obtain their target and enter closed loop operation.

Track

The Track action is responsible for initiating the WFS target acquisition and closed loop operation. Once the slew to the science target has completed the AO sequencer can be commanded to acquire the WFS targets and close the RTC correction and AO offloading loops. The track sequence will acquire the LGS WFS first followed by the LOWFS, and finally the NGS WFS.  Each WFS acquisition phase in this sequence must be explicitly initiated by an operator before continuing. The following outlines the track sequence and will vary based on each observing mode (i.e. LGS, NGS, Fixed-Field, and Fixed-Pupil) and instrument.

LGS WFS Acquisition

The following steps are performed serially:

- Command the LGS WFS Composite to perform a background calibration for each WFS.

- Command the Laser Unit Composite (laser facility) to tune lasers on 589nm.

- Command the LGS WFS Composite to acquire the LGS beacons and close the tracking loops. This involves motion control of the Asterism Tip-Tilt and Rotator and the BGS Patrolling Steering Arms to align the beacons with the WFS.

LOWFS Acquisition

The following steps are performed serially:

- Command the LOWFS Composite to acquire the TT NGS targets and close the TT and mems DM control loops. This involves imaging the field and moving the pickoff arms to acquire the NGS target.

NGS WFS Acquisition

The following steps are performed serially:

- Command the NGS WFS Composite to acquire the NGS target. This will cause the composite to steer the FSMs to center the target on the WFS.

- Command the NGS WFS Composite to perform a background calibration. This involves dithering the High-Order Tip-Tilt and taking a measurement of the background signal.

Once the subsystem acquisition is complete the RTC DM and Tip-Tilt loops can be closed, and the remaining AO Controls System offloading and offsetting tasks can be started. At this point the system is tracking the science target, and instrument observing can begin.

Halt

The Halt action is responsible for opening the servo and correction control loops for the AO and RTC control system, and is executed to terminate an observing sequence. The following command sequence is executed to terminate an observation.

- Command the ADC Composite to halt.

- Command the Offloading Composite to halt.

- Command NGS WFS Composite to halt. This will open the RTC servo loops and halt any device motion.

- Halt the High-Order Tip-Tilt Controller and open the tweeter-DN RTC servo loop.

- Command the LOWFS Composite to halt. This will open the RTC servo loops and TWFS focus control loop.

- Command the LGS WFS Composite to halt. This will open the LOWFS and woofer DM RTC servo loops, LGS focus tracking loops, and LGS WFS tip-tilt offloading process.

- Command the Image Rotator Controller to halt.

After the halt sequence is executed the telescope can be re-pointed and the AO system can reacquire a new set of targets. Halt should also be performed prior to shutting down for the night.

Recover

The Recover action is a state dependent process that is responsible for recovering the AO system in the event of a fault. Faults events for the AO Sequencer typically occur because on or more sub-devices have themselves entered a fault state. The recover command will be issued to each of the faulted composites / controllers. The components will attempt to clear the error and return the device to its original state before the failure occurred. If the fault can not be cleared the device will remain in the fault state, but the Sequencer and subsystems will still be able to receive commands. If the client attempts to command the faulted device(s) the fault state will be passed up to the AO Sequencer, and will require a response from the operator.

#### Outputs

Subsystem output is provided through named telemetry topics. The following describe the telemetry provided by the subsystem, in addition to the standard outputs discussed in 4.1.2.

ObservingMode

The ObservingMode telemetry item defines the mode for which the AO Control Sequencer and subsystems have been configured.

RotatorTarget

The RotatorTarget telemetry item defines the target and value the rotator is tracking. This can be PA (fixed-field), VA (fixed-pupil), or stationary.

AdcTipTilt

The AdcTipTilt telemetry item defines the tip-tilt offset correction the science Tip-Tilt stage (TTM2) has applied to correct for ADC aberrations. This value is combined with the overall TTM2 Offset value.

### Control Loops

The AO Control Sequencer does not directly implement any control loops.

### Devices

The AO Control Sequencer is the highest level component in the AO hierarchy, and as such, has the responsibility of managing the overall state of the entire AO facility. Most devices, however, are managed by at least one intermediate subsystem composite, which in turn may define its own component hierarchy: these subsystems have been detailed in the previous sections of this document. Although most of the AO facility device control is delegated to subsystems, a few devices are under the direct control and management of the AO Control Sequencer. These are defined under the Wide Field Relay and Narrow Field Relay subsystems.

# Laser System

## Introduction

Figure 29 Laser Control Block Diagram

## Laser Switchyard

The NGAO laser system is designed for center mounted propagation, with the laser asterism launched from the back of the telescope secondary. As there are size and weight restrictions on the secondary socket, the actual generation of the lasers will have to be done elsewhere. This will require a system of optics to transport the laser beams from the laser facility to the launch telescope. The transport optics and related hardware are all part of the laser switchyard (the laser units themselves are considered a separate subsystem), and are found in the laser enclosure on the nasmyth platform.

The transport optics consists of tip-tilt mirrors and an associated laser beam positioning sensor. Together the steering mirrors and sensors allow the system to control the path of the lasers as they travel to the secondary. The switchyard also includes a set of polarizers for each laser; a low-power fold mirror that can be used for alignment and acquisition; and the fast safety shutter which is controlled by the Laser Safety Subsystem.

The laser switchyard optics are controlled and managed by the Laser Switchyard Subsystem.

### Description

The Laser Switchyard control system is implemented by the LaserSwitchyardComposite. This component is responsible for controlling the individual devices and distributed software controllers that implement the switchyard functionality. This includes the low-power fold mirror, steering mirrors, beam polarizers, and beam positioning sensors.

Figure 30: Laser Switchyard Collaboration Diagram

The Switchyard Composite is managed directly by the Laser Facility Composite, and presents a simple high-level interface to initialize, acquire, and halt the switchyard devices. The Composite’s primary responsibility is to prepare the switchyard devices for closed loop observing. The majority of the Switchyard Composite functionality exists as part of the configuration and acquisition sequence. During observing the tracking components are independently controlled or run closed-loop with the rest of the system.

The Laser Switchyard Subsystem implements a total of nine tracking control loops: one for each of the quarter and half polarizers, and for each laser steering mirror. The half and quarter polarizer controllers track the required polarization of the laser as a function of the telescope elevation. The Laser Steering Mirror Controllers are responsible for maintaining beam alignment with the BGS, and track with a combination of the telescope elevation and centroid information from the positioning sensors.

### Interfaces

The following sections detail the input and output interface to the Switchyard subsystem controller.

#### Inputs

The primary inputs to the Laser Switchyard composite are implemented as KCSF actions. Actions are defined through an attribute list and may contain parameters. The following define all of the recognized actions available to the Switchyard composite.

Initialize

The Initialize action is responsible for configuring the subsystem composite, controllers, and devices, and must be executed prior to acquisition and observing. The initialize action will connect the Laser Switchyard Subsystem to each of the subsystem dependencies, and system inputs. This includes:

- Low-Power Fold Controller

- ¼ and ½ Polarizer Controllers

- Laser Positioning Sensor Controllers

- Laser Steering Mirror Controllers

The Controllers will in turn connect to their related telemetry and component dependencies; load configuration information; and initialize, home and position the motion control hardware

Once the initialization of the subsystem and devices are complete the NGS WFS Composite will be ready for observing.

Acquire

The Acquire action is responsible for commanding the laser switchyard components to prepare the switchyard for laser propagation and transport.

This action has the following effect:

- Track the laser steering mirrors.

- Track the ¼ polarizers.

- Track the ½ polarizers.

The composite will wait for each device to complete the acquisition process. The collective result of the acquisition process will be returned by the action.

LowPower

The LowPower action is responsible for configuring the switchyard system for low power laser propagation. This is performed by inserting the Low-Power fold mirror into the laser beam path, effectively redirecting 99.9% of the laser light to a dump and allowing the leakage to continue through the system. As a requirement, the Laser shutters must be closed before the fold can be inserted. If the lasers are not shuttered the action will fail. This validation step is performed by the Low Power Fold Controller.

HighPower

The HighPower action is responsible for configuring the switchyard system for high power laser propagation. This is performed by removing the Low-Power fold mirror from the laser beam path allowing the laser to propagate at full strength. As a requirement, the Laser shutters must be closed before the fold can be removed. If the lasers are not shuttered the action will fail. This validation step is performed by the Low Power Fold Controller.

Halt

The Halt action is responsible for halting device motion and opening the tracking loops. This action will cause the laser steering mirrors, polarizers and beam positioning sensors to all enter their HALT state. As these devices are no longer tracking, flexure compensation will not be performed. If the switchyard is not tracking the laser units or safety shutter should be closed.

Recover

The Recover action is responsible for selectively initializing faulted sensors devices by Laser Switchyard Subsystem. When the subsystem receives the recover action it will send a recover command to each of the switchyard controllers. If any of the controllers are faulted they will attempt an initialization, and clear the fault if the device reports the initialization was successful. . If the recover is successful the switchyard composite will clear its own fault state. The recover action can be issued at any time, and will not affect healthy components.

#### Outputs

Subsystem output is provided through named telemetry topics. The following describe the telemetry provided by the subsystem, in addition to the standard outputs discussed in 4.1.2.

LaserPositionCentroid

The LaserPositionCentroid telemetry item defines the centroid position of the laser beams as seen by the beam positioning sensors. This information is used by the steering mirrors to compensate for telescope flexure and maintain beam alignment. Centroid telemetry is defined for each positioning sensor.

### Control Loops

The following sections detail the tracking and offload control loops of the Switchyard subsystem. Each control loop is implemented in part or in full by one or more of the switchyard subsystem controllers.

#### Laser Transport

The laser transport control loop is responsible for maintaining the proper alignment of the lasers from the laser enclosure to the BGS. This control loop is needed to compensate for flexure as the telescope tracks a target; and uses a combination of the telescope elevation and feedback from the beam positioning sensors to calculate the appropriate position of the laser steering mirrors.

Figure 31: Laser Steering Control Loop

Although the laser transport tracking loop (implemented by the Laser Steering Mirror controller) uses inputs from DCS and the positioning sensors, it does not use both of them at the same time. The laser facility can propagate the lasers in high-power and low-power modes, depending on the position of the low-power fold mirror. When the fold mirror is in the beam path, all but a small percentage of the laser light is diverted to a beam dump. The leakage is allowed to propagate through the system and can be used for alignment and tests with little risk to hardware or personnel. When the fold mirror is retracted the full beam is allowed to propagate.

The laser transport control loop will use the flux value from the beam positioning sensors to determine if the system is configured for low or high-power propagation. If the flux is below a certain threshold the control loop will use the current telescope elevation from DCS. The elevation will be used to index a simple table to determine the appropriate flexure compensation to apply to the steering mirrors. These values will have been derived from tests empirical tests and will be interpolated to find the appropriate tip-tilt to apply to the stage.

If the positioning sensor threshold reports sufficient laser light, centroid information will be obtained from the detector and will be used to calculate the precise compensation required to keep the laser aligned with the BGS.

In either low or high-power track mode the operator can apply an offset to the steering mirrors to redefine the ideal centroid position. The Laser Steering Mirror tracking loop should be closed whenever the lasers are propagating.

#### Laser Polarization

During observing the laser asterism needs to have the proper polarization as seen by the AO wavefront sensors.  Fortunately this polarization is directly related to the elevation of the telescope and incidence angle of the laser with the sodium layer. The control loop to maintain the desired polarization is implemented by each of the polarizer controllers.

Figure 32: Laser Polarization Control Loop

The polarizer controller tracking loop is event driven and responds to telescope elevation events from DCS. When an event is received the controller will lookup the required orientation of the polarizer in a table based on the current elevation. The table will be defined through empirical tests at regular telescope elevations. The actual polarizer orientation will be interpolated based on these values.

The Half and Quarter Polarizer tracking loop can be remain closed throughout a night’s observing.

### Devices

The following devices are managed by the Laser Switchyard subsystem, and are listed in the order the laser light propagates through the switchyard. Specific details, interfaces, and discussions are available in their related design documents.

#### Half & Quarter Polarizers

Each laser unit has a dedicated half and quarter polarizing lens. The polarizers are responsible for maintaining the proper polarization of the lasers for observing. The required polarization is based on the angle of incidence of the lasers and the sodium layer, which can derived from the current telescope elevation.

#### Low-Power Fold Mirror

The low-power fold mirror is designed to redirect all but a small fraction of the laser light into a cold dump. The leakage from the mirror is allowed to propagate through the system to the BGS. At this power level the laser light can not cause any damage to the system during alignment, calibration or testing. The low-power fold mirror is common for all three laser beams.

#### Laser Steering Mirrors

There are a total of three tip-tilt laser steering mirrors – one for each laser – responsible for propagating the lasers from the switchyard to the BGS. The steering mirrors will compensate for flexure as the telescope moves by using pre-measured flexure values, and feedback from beam positioning sensors on the BGS.

#### Beam Position Sensors

The three beam positioning sensors work closed loop with the laser steering mirrors to correct for telescope flexure. As the telescope changes in elevation the gravity vectors will exert forces on the telescope structure causing flexure. The sensors will see this as a change in the beam position for which the laser steering mirrors will then compensate.

## Laser Beam Generation System

NGAO is a multi-laser adaptive optics system designed to correct for atmospheric turbulence over multiple narrow fields. As part of the science requirements of the system, a constellation of lasers will need to be produced and positioned on sky to perform tomography on the seeing column. In addition, up to three natural guide stars may also be used to provide full-field tip-tilt and focus correction.

To satisfy these requirements the laser system will make use of four fixed laser guide star (LGS) beacons for tomography, and three patrolling LGS beacons to individually sharpen the wide-field tip-tilt stars. The beacons must be flexible enough to be positioned anywhere in the wide-field, and the entire asterism must be able to be orientated arbitrarily.

All of the hardware needed to produce this laser asterism are part of the laser beam generation system (BGS), and is located behind telescope secondary mirror. The BGS is controlled and managed by the Laser Beam Generation Subsystem.

### Description

The BeamGenerationSystemComposite is a high-level multi-device manager that is responsible for acquiring and configuring the hardware on the BGS. These devices include the patrolling steering arms, beam splitters, asterism rotator, tip-tilt mirror, and beam expander focus.

Figure 33: Beam Generation System Collaboration Diagram

The BGS Composite is managed directly by the Laser Facility Composite, and presents a simple high-level interface to initialize, acquire, and halt the BGS devices. The majority of the BGS Composite functionality exists as part of the configuration and acquisition sequence. During observing the majority of the subcomponents are independently controlled or run closed-loop with the rest of the system.

The BGS subsystem is responsible for configuring and positioning the hardware to obtain the desired laser constellation. The constellation is based on the number and position of the wide-field targets, the position of the telescope, and the desired image orientation as seen from the science detector. The BGS subsystem will take this information and calculate the required positions for each of the associated devices.

The BGS subsystem also manages the focus and propagation of the lasers and implements various safety processes to ensure the safety of the laser generation system.

The following sections detail the BGS subsystem and devices.

### Interfaces

The following sections detail the input and output interface to the BGS subsystem controller.

#### Inputs

The primary inputs to the BGS composite are implemented as KCSF actions. Actions are defined through an attribute list and may contain parameters. The following define all of the recognized actions available to the BGS composite.

Initialize

The Initialize action is responsible for configuring the subsystem composite, controllers, and devices, and must be executed prior to acquisition and observing. The initialize action will connect the Beam Generation Subsystem to each of the subsystem dependencies, and system inputs. This includes:

- Asterism Tip-Tilt Controller

- Asterism Imager

- Asterism Rotator

- Patrolling Steering Arm Composite

- Beam Expander Focus Controller

- Final Shutter Controller

- Laser Beam Splitter Controllers

The Asterism Imager Controller will connect to and configure the asterism camera by setting the default sensor parameters and creating the image output telemetry stream. The Patrolling Steering Arm Composite will connect to and initialize each of the patrolling steering arm controllers, and position them evenly in the field. The Asterism Tip-Tilt, Rotator and Beam Expander Controllers will connect to their related tracking input streams. All motion control devices will be initialized and homed.

Once the initialization of the subsystem and devices are complete the NGS WFS Composite will be ready for observing.

Acquire

The Acquire action performs the target acquisition of the BGS system. This action requires the pointing target for the telescope (including any telescope offsets) and up to three LGS targets, all in sky coordinates. The BGS Composite will calculate the offsets of the LGS targets relative to the optical axis in arcseconds, and command the Patrolling Steering Arm Composite to acquire the field positions. The number of LGS targets will also be used to determine the appropriate configuration for the patrolling laser Beam Splitters.

The Acquire action is also responsible for putting the BGS devices into the proper state for observing:

- Asterism Tip-Tilt Controller is put into TRACK.

- Asterism Rotator is put into TRACK.

- Asterism Imager Controller is put into free-run mode.

- Beam Expander Focus Controller is put into TRACK.

If any of the devices fail to enter their required states, or there is a problem acquiring the patrolling target positions, the Acquire action will fail.

Close

The Close action commands the BGS Composite to close the final shutter – blocking laser propagation. This action will typically be executed during a configuration mode change and when an observing sequence has been halted.

Open

The Open action commands the BGS Composite to open the final shutter, and is performed as the last step of laser acquisition and propagation sequence.

Halt

The Halt action is responsible for halting all devices on the BGS and is typically issued to interrupt or terminate a sequence. Any devices currently in motion will be immediately stopped, and all loops will be opened: this includes stopping automated asterism validation (Validation Controller) and taking the asterism imager out of free-run.

Halt can be issued from any of the post-initialization states, and from the HALT state the controller can be reinitialized or perform a new acquisition.

Recover

The Recover action is responsible for selectively initializing faulted devices managed by the Beam Generation System.  When the subsystem receives the recover action it will send a recover command to each of the BGS controllers. If any of the controllers are faulted they will attempt an initialization, and clear the fault if the device reports the initialization was successful. If the recover is successful the BGS will clear its own fault state. The recover action can be issued at any time, and will not affect healthy components.

#### Outputs

Subsystem output is provided through named telemetry topics. The following describe the telemetry provided by the subsystem, in addition to the standard outputs discussed in 4.1.2.

Targets

The Targets telemetry item contains the sky information used to calculate the position of each patrolling steering arm.

Focus

The Focus telemetry item defines the focus position of the beam expander stage.

RotatorTarget

The RotatorTarget telemetry item defines the target and value the asterism rotator is tracking. This can be PA (fixed-field), VA (fixed-pupil), or stationary.

### Control Loops

The following sections detail the tracking and offload control loops of the BGS subsystem. Each control loop is implemented in part of in full by one or more of the NGAO controllers.

#### Focus Tracking

During observing the BGS subsystem will have to track the changing focus of the LGS asterism. Focus is primarily based on the current telescope elevation and launch telescope temperature. As the telescope elevation changes, the beam length of the laser from the telescope to the sodium laser will fluctuate. As the apparent distance to the sodium layer changes the beam expander will have to compensate to maintain LGS focus. The temperature of the launch telescope also affects the focus of the laser. This information is derived from the environmental sensors within the LLT.

Figure 34: Laser Focus Tracking Control Loop

Beam expander tracking is an event driven control loop triggered by telemetry events from the LLT environment sensors and DCS telescope elevation keywords. When the controller receives input from these items it will calculate the LGS asterism focus compensation and apply it to the stage. A lookup table should be sufficient to define focus compensation based on telescope elevation and temperature.

#### Rotator Tracking

During observing the laser asterism will need to track the LGS targets as the telescope rotates. This is to ensure that the asterism maintains alignment with the LGS WFS detectors. In fixed field mode this is based on the position angle (PA). In fixed pupil mode the rotator will track the vertical angle (VA).

Figure 35: Asterism Rotator Tracking Control Loop

Rotator tracking and offloading is event drive. The Asterism Rotator Controller will receive the position and vertical angle from the AO rotator through a telemetry stream. The asterism rotator will process this information and track with telescope and field rotation.

#### Flexure Compensation

Changes in telescope elevation will induce flexure on the BGS and launch optics. As the telescope tracks a target the asterism tip-tilt will need to compensate for flexure to keep the asterism stable on the sky.

Figure 36: Flexure Compensation Control Loop

The Asterism Tip-Tilt Controller implements an event driven tracking loop. Primary input to the control loop is the telescope elevation and is made available from DCS through the Channel Access service. A simple look-up table can be used to interpolate the required flexure compensation for the current elevation.

### Devices

The following devices are managed by the BGS subsystem, and are listed in the order the laser light propagates through the BGS. Specific details, interfaces, and discussions are available in their related design documents.

#### Beam Splitter

The BGS subsystem manages two linear 1-DOF beam splitters. These devices are used to split the patrolling LGS laser into the required number of wide-field LGS beacons. The beam splitters can be configured to produce one, two, or three beacons: off-beam, one on-beam or both on-beam, respectively. The specific configuration will depend on the number of tip-tilt guide stars used. The beam splitters are controlled through KCSF Linear 1-DOF Controllers.

#### Patrolling Laser Beacons

The patrolling steering arms allow the BGS to position the LGS beacons anywhere in the permitted range of the asterism. There are a total of three steering arms, each implemented as a linear 2-DOF device. Each arm can be positioned arbitrarily in the field; with software restrictions preventing them from being within the central asterism and from occluding one another. The three steering arms are managed by the Patrolling Steering Arm Composite, which is responsible for approving target acquisition, preventing laser obstruction, and managing the device interaction. Additionally, this composite will receive offloads from the AO Controls System to align the LGS beacons with the WFS (see section 6.4.3.2 for more information). The steering arm devices are controlled through the KCSF Linear 2-DOF Controller.

#### Asterism Rotator

The asterism rotator is a rotational 1-DOF K-mirror device responsible for the orientation of the entire laser asterism. The asterism rotator is designed to keep the laser asterism aligned to a specific orientation with respect to the sky (fixed field) or detector (fixed pupil). This tracking device receives inputs from the AO facility as offsets and must maintain a relative orientation to the AO image rotator. The asterism rotator is controlled through the NGAO Asterism Rotator Controller.

#### Asterism Tip-Tilt

The asterism tip-tilt mirror is a 2-DOF device that allows the AO controls system to shift the entire laser asterism on the sky. This common mirror is used during the acquisition and observing sequences for LGS WFS alignment and flexure compensation. The tip-tilt mirror will compensate for flexure by tracking the telescope elevation, and any additional offsets issued by the control system. The tip-tilt mirror is controlled through the NGAO Asterism Tip-Tilt Controller.

#### Beam Expander

The beam expander stage is a linear 1-DOF device that is responsible for maintaining the required LGS focus.  This tracking device responds to changes in the telescope elevation and the environmental conditions in the LLT. The beam expander focus stage runs in a nearly continuous automated fashion, and is controlled through the NGAO Beam Expander Focus Controller.

#### Asterism Imager

The asterism imager is a slow-rate diagnostic camera that provides a view of the LGS asterism prior to the launch telescope. This camera picks up the leakage from the LLT fold mirror on the BGS, and shows the real-time structure of the asterism after the rotator, tip-tilt and steering arms. The asterism imager camera is controlled through a KCSF Camera Controller and is primarily used by the Asterism Validation task.

## Laser Safety System Interface

NGAO uses three high-powered lasers to create the LGS asterism. These lasers have the potential to cause significant damage to the laser facility and personnel if not used properly. To reduce the chances of damage or injury when using the lasers, an independent safety system will be developed to implement the laser safety protocols. This system, known as the Laser Safety System (LSS), will monitor the state and health of the lasers and control their propagation.

The LSS is responsible for monitoring the state of the laser system and ensuring safe operating conditions. A number of devices are under the control of the safety system, and can not be controlled directly by the software system. In order to offer a level of control to these devices (such as the beam dump and BGS fast shutter) functionality will exist in the safety system to accept control requests. If the requests do not risk the safety of the system or personnel they will be performed.

For the most part the LSS is a fully independent system that will run continually while the lasers are powered. Direct hardware and sensor connections are made to the LSS to ensure it has access to all of the critical system state information. The safety system itself manages a number of devices in the NGAO system that can not be controlled directly by the software system; for example the BGS fast shutter. In order to offer a level of control to these devices an interface will be provided to accept device control commands.  Additional software based permissives are also used to notify the LSS of state changes and propagation requests outside of its physical connections. These LSS permissives and commands are made available through the NGAO Laser Safety System Interface.

### Description

The LSS Interface is implemented by the LaserSafetySystemController, and provides the operational interface to the laser control system. The LSS Controller provides clients with actions and attributes to set permissives, execute device commands, and read-back state information about the lasers.

Figure 37: Laser Safety System Collaboration Diagram

A number of facility components interact with the Safety System (depicted by the yellow boxes). The LSS utilizes these connections for commanding and feedback. The NGAO Controls system does not need to be directly aware of the state of these components, and instead can rely on the LSS to provide telemetry information indicating the overall safety / health of the laser system. Other safety permissives such as spotter and OA acknowledgments must be provided by the NGAO LSS Controller.

### Interfaces

The following sections detail the input and output interface to the Laser Safety Subsystem.

#### Inputs

The primary inputs to the LSS composite are implemented as KCSF actions. Actions are defined through an attribute list and may contain parameters. The following define all of the recognized actions available to the LSS composite.

Initialize

The Initialize action is responsible for connecting the Laser Safety System Controller to the LSS device driver and creating the telemetry output streams. Once the initialization has completed clients will be able to interact with the LSS through the NGAO controls system.

SetPermissive

The SetPermissive action is responsible for setting the state of one or more LSS permissives. This action can only be executed from within the post-initialization states.

A permissive is defined by a string and a boolean value, and can be any one of the recognized names available to the LSS. The permissives include, but are not limited to, the following:

- Observing Assistant Acknowledgement: The final permissive from the OA to permit laser propagation on-sky.

- Laser Traffic Control System: Confirms there is no laser or observing collisions between the other Mauna Kea observatories.

- Spotter Acknowledgment: Confirmation from the spotters that it is safe to propagate.

- Telescope Control Acknowledgement: Confirmation that the telescope is tracking and in the safe telescope elevation window for laser propagation.

- Satellite Safety Acknowledgment: Confirmation from space traffic control that there are no imminent satellite crossings over the laser path.

- Laser Transport Acknowledgement: Software confirmation that the lasers are being transported to the BGS properly. (This can double for a valid asterism confirmation.)

- Reduced Performance Acknowledgment: Permits (or prevents) the system from running with a reduced set of laser units. This may be needed in the event that one of the lasers is malfunctioning, but observing should continue.

These permissives cover the majority of system state properties that are not directly visible by the Laser Safety System. The full set of software supplied permissives will be finalized after the laser vendor has been chosen.

Close

Closes the specified shutter. This can be any of the subordinate shutters that comprise the laser safety system, including the beam dump, fast shutter, and laser unit shutters.

Open

Opens the specified shutter. This can be any of the subordinate shutters that comprise the laser safety system, including the beam dump, fast shutter, and laser unit shutters.

FullStop

This action tells the Laser safety System that it should immediately put the system into an emergency safe state. This includes closing shutters, powering down lasers, and taking any additional full-stop safety precautions.

Acknowledge

The Acknowledge action is used to force the LSS Controller out of an alarm state by acknowledging the existence of an alarm condition. Ideally alarm conditions will be satisfied by either correcting the problem identified by the LSS or setting the appropriate permissives. This should then cause the LSS Controller to return to the READY state. In the case where the problem can not be immediately handled and observing is to continue the acknowledge action can confirm receipt of problem, and put the controller back in the READY state allowing it to respond to other alarm conditions that may occur.

Note: the acknowledge action does not change the state of the LSS, it simply silences the current alarm state from within the NGAO control system. If the LSS issues the same error condition again it will result in reentrance of the ALARM state.

Halt

The Halt action is responsible for putting the LSS Controller into the HALT state. As the LSS Controller does not have any motion control requirements this action is implemented as a no-op routine, and exists primarily for consistency with the NGAO Controller design.

#### Outputs

Subsystem output is provided through named telemetry topics. In addition to the standard outputs discussed in 4.1.2, the LSS Interface Subsystem will provide the following telemetry status information:

- Shutter state for each laser unit and the fast / safety shutters.

- LTCS and Satellite / Aircraft warnings.

- Laser unit health and operating status.

- Values for all of the software provided permissives.

### Control Loops

The Laser Safety System Controller is a passive interface to the safety system, and expects clients to provide it will all of the required permissives. The LSS Controller does not implement any control loops related to monitoring or setting safety permissives.

### Devices

The following devices are managed by the Laser Safety Subsystem. Specific details, interfaces, and discussions are available in their related design documents.

#### Laser Safety System

The Laser Safety System is a safety and operational control system that will be built in-house for the NGAO project based on the current laser safety system developed for K2 LGS. The safety system design is covered in detail on the Laser System Design Wiki.

## Laser Service Enclosure

The laser facility enclosure contains a number of diagnostic and environment sensors. These sensors will be distributed throughout the laser enclosure to obtain global and localized measurements of environmental conditions. The readings from these sensors will be used primarily to monitor and ensure the optimal operating conditions for the laser hardware.

### Description

The LaserEnclosureComposite manages the environmental sensors found in the laser unit enclosure. These sensors are responsible for providing the current temperature, humidity, and particulate level.

Figure 38: Laser Service Enclosure Collaboration Diagram

As sensors are simple on / off devices without motion control requirements, the composite itself is very simple, and provides the basic capabilities to initialize the system for observing. The composite will collect data from each of its sensors and will provide a summarized health status as telemetry and alarming.

### Interfaces

#### Inputs

The primary inputs to the Laser Service Enclosure subsystem are implemented as KCSF actions. Actions are defined through an attribute list and may contain parameters. The following define all of the recognized actions available to the Laser Service Enclosure subsystem.

Initialize

The Initialize action is responsible for configuring the subsystem composite, controllers, and devices. The initialize action will connect the Laser Enclosure Subsystem to each of the subsystem sensors, specifically the temperature, particulate and humidity sensors. Multiple instances of each sensor may be found within the service enclosure to obtain localized environmental conditions.

#### Outputs

The Laser Enclosure subsystem does not implement any additional outputs outside of the standard telemetry items defined in section 4.1.2

### Control Loops

As the Laser Enclosure subsystem does not manage any motion control devices, and the environmental information of the laser enclosure is not used in a controls context, this subsystem does not implement any control loops.

### Devices

The following devices are managed by the Laser Enclosure subsystem. Specific details, interfaces, and discussions are available in their related design documents.

#### Environmental Sensors

The Laser Enclosure subsystem manages the temperature, humidity, and particulate sensors found in the laser enclosure. These sensors run continually, even outside of the regular NGAO control system, and provide environmental information to the operators through the observing tools. Although not used in a direct controls context, the environmental sensors will report excessive environmental conditions to the NGAO alarm system, allowing the operator to take action if needed.

## Laser Interface

The NGAO laser system is designed to produce up to seven laser beacons using three laser sources. Two of the laser units are responsible for producing the central asterism of four beacons (each split 50 / 50). The third laser unit is responsible for producing the three patrolling beacons (split 33 / 33 / 33). Depending on the observing mode all three lasers may be used, just the central asterism lasers, or no lasers (NGS observing). The laser units are all managed under the Laser Interface Subsystem.

### Description

The Laser Interface Subsystem is implemented by the LaserSystemComposite, and manages the three individual laser units and provides propagation control through the Laser Safety System.

Figure 39: Laser Interface Collaboration Diagram

This Laser System Interface provides a single unified command and status interface to all three lasers units. Clients can treat the laser sources as a single instance, allowing them to issue a single command in place of three. Additionally, health and status information for all three laser units will be consolidated and made available to the control system for easy reference.

The Laser System Composite is managed directly by the Laser Facility Composite.

### Interfaces

#### Inputs

Initialize

The Initialize action is responsible for configuring the Laser Interface Subsystem composite, controllers, and devices, and must be executed prior to acquisition and observing. The initialize action will connect the Laser Interface Subsystem to each of the subsystem dependencies, which includes the three laser controllers and the Laser Safety System. The Laser Controllers will in turn connect to their associated laser controller and perform a laser calibration.

Once the initialization of the subsystem and devices are complete the NGS WFS Composite will be ready for observing.

SetMode

The SetMode action configures the Laser Interface Subsystem for the current observing mode. Specifically, it tells the composite whether it should interact with the patrolling laser unit when it receives global laser control commands or status requests (e.g. Open). Since this laser unit is not used during certain observing modes (e.g. fixed-pupil), we want the control system to exclude it from regular operation – the laser unit itself is still functional and can be commanded directly through the safety system or associated Laser Controller.

The composite will maintain this configuration until it is explicitly changed in a call to SetMode.

Close

The Close action is responsible for closing the active laser units’ main shutter and is performed through the Laser Safety System.

Open

The Open action is responsible for opening the active laser units’ main shutter and is performed through the Laser Safety System. While open the laser light will propagate through the beam train.

TuneOn

The TuneOn action is responsible for tuning the active lasers on to the 589nm wavelength. This must be performed for the AO wavefront sensors to operate properly.

TuneOff

The TuneOff action is responsible for tuning the active lasers off of the 589nm wavelength. The actual wavelength will be defined in configuration for each unit. This action is performed when taking measurements of the background signal.

Halt

The Halt action is responsible for putting the laser units in a halted / safe state. Specifically this action will shutter the laser units to prevent laser propagation. The halt action is required prior to disabling and shutting down the composite and controllers.

Recover

The Recover action is responsible for selectively initializing faulted laser devices and controllers managed by the Laser Interface Subsystem.  When the subsystem receives the recover action it will send a recover command to each of the laser controllers. If any of the controllers are faulted they will attempt to clear the error and return the device to an operational state. The recover action can be issued at any time, and will not affect healthy components.

#### Outputs

Subsystem output is provided through named telemetry topics. The following describe the telemetry provided by the subsystem, in addition to the standard outputs discussed in 4.1.2.

ShutterState

This attribute returns the current global shutter state of the active laser units. If the active laser shutters are in different states Open will be returned.

TuneState

This attribute returns the current global tuned state of the active lasers. If the active lasers are in different states Off will be returned.

### Control Loops

The Laser System Interface does not implement any control loops.

### Devices

The following devices are managed by the Laser Interface subsystem. Specific details, interfaces, and discussions are available in their related design documents.

#### Lasers

The Laser Interface Subsystem manages the controllers for the three laser units. These controllers provide the capabilities to initialize, calibrate, and propagate the lasers. The controllers themselves do not actually interface directly with the lasers, but acts as a wrapper around the LSS to provide the set of laser commands available to clients.

## Laser Control Sequencer

NGAO can be divided into an up-link (laser) and down-link (AO) system. The up-link system is responsible for producing and propagating the laser asterism on sky. The down-link system is responsible for correcting distortion in the field image caused by atmospheric turbulence. As there is such a clear and distinct functional separation between the two systems each can be modeled with their own control system hierarchy.

The up-link system is implemented by the Laser Facility System. This hierarchical control system implements a number of tracking and non-tracking devices with varying motion control requirements. At the highest level of the Laser Facility System is the Laser Control Sequencer. This KCSF Composite provides a high-level interface to the entire laser facility, and is responsible for executing the main acquisition and observing sequences; and managing the operation and state of the laser subsystems. The following sections detail the Laser Control Sequencer.

### Description

The Laser Control Sequencer is implemented by the LaserFacilityComposite, and is responsible for managing and configuring the entire laser facility, and executing the various sequences. The Composite sits above all of the other laser system components, and coordinates their state and task execution during observing. The Laser Sequencer is under the direct control of the MSCS, and is the primary interface to the laser system.

Figure 40: Laser Control Sequencer Diagram

The Laser Sequencer exposes a simple set of high level actions to the MSCS. These actions provide the Command Sequencer with the ability to initialize and configure the laser facility; acquire LGS targets; and shutdown the system. The composite in-turn takes the action data and distributes it among the relevant subcomponents; invoking their corresponding initialization, configuration, acquisition, and shutdown actions.

As with other composites, the Laser Sequencer will reflect the overall state of its subcomponents: effectively the health and state of the entire laser system. The MSCS will monitor the state of the composite to ensure its continued safe operation, and to respond to any errors that may occur.

#### Laser Sequence

The laser control sequence covers everything from bringing up the software system to acquisition and finally shutting down for the night. The sequence itself can be divided into phases each covering a specific state of the system during a typical night’s observing. These phases are (listed in order of execution):

- Startup

- Initialize

- Configure

- Acquire

- Observe

- Halt

- Power-down

For the most part these phases can be executed straight through to perform NGAO corrections for a science target. However there is flexibility allowing the sequence to return to a previous phase or jump forward to a latter phase. The following state diagram describes the relationship between each phase in the laser sequence:

Figure 41: Laser Sequencer State Diagram

The main branch of the sequence proceeds through each of the phases as outlined. At specific points in the sequence the control can be sent back to a previous step, such as going from Halt to Acquisition. This flexibility allows the system to be reconfigured multiple times throughout a night for greater flexibility. (Note: this diagram is a simplified version of the actual Laser Facility control sequence. The complete sequence employs a more complex state machine, with some details left out here for brevity. The full Laser Facility Sequencer state machine is documented in the NGAO Laser Facility Component Design Document.)

A detailed review of the Laser sequences is described in KAON 736: NGAO Laser Facility Sequence.

### Interfaces

The following details the Laser Control Sequencer interface.

#### Inputs

The primary inputs to the Laser Sequencer are implemented as KCSF actions. Actions are defined through an attribute list and may contain parameters. The following define all of the recognized actions available to the Laser Sequencer.

Initialize

The Initialize action is responsible for configuring the Laser sequencer, subsystems, and devices, and must be executed prior to acquisition and observing. The initialize action will connect the Laser Sequencer to each of the composite’s dependencies, specifically:

- Sky Alignment Camera Controller

- Asterism Validation Task

- Laser Safety System Controller

- Laser Enclosure Subsystem

- Laser Interface Subsystem

- Laser Switchyard Subsystem

- Laser Launch Telescope Subsystem

- Beam Generation Subsystem

Each of the subsystem dependencies will in turn initialize their subordinate controllers and devices as discussed in the previous sections. The individual device controllers will be commanded to connect to their low-level hardware and initialize / home the devices. Any additional communication dependencies (such as tracking loop inputs) will also be established.

In addition to the primary dependencies listed above, at different stages in the observing sequence the Laser Control Sequencer will connect to other laser system devices to perform specific mode-dependent configurations and commands. These operations will be performed outside of the knowledge and control of the parent composites and subsystems. This allows us to separate the mode and instrument dependent functionality from the low-level system and keep it all within the Laser Control Sequencer.

Once the initialization of the Laser facility is complete the system can be configured for observing.

SetMode

The SetMode action is responsible for configuring the laser facility composite and subcomponents for a particular observing mode. Mode configuration will position devices as required by the observing mode.

In fixed-field observing the entire laser facility is available for use. In this mode up to three patrolling LGS beacons are used to correct for atmospheric effects for wide-field tip-tilt as measured by the corresponding NGS stars. As the telescope rotates to track the science target the LGS asterism must also rotate on sky to compensate for the changing position angle of the field. This effectively fixes the laser constellation with the star field as seen from down stream.

In fixed-pupil mode the field is allowed to rotate freely, but the pupil must remain fixed. In this mode the patrolling LGS beacons are not used. As a result, the majority of the wide-field laser beamtrain can be idled and shuttered. To stay aligned will the pupil the laser asterism will only need to match the orientation of the telescope’s vertical angle.

The system must be configured each time the observing mode changes, and must be performed prior to target acquisition and observing.

Acquire

The Acquire action is responsible for commanding the laser facility components to acquire a set of LGS targets. Ultimately the laser asterism will be configured to match a set of sky coordinates based on the orientation of the field and pointing of the telescope. The acquisition process is typically initiated during the slew to the next pointing target.

This action requires the current pointing information for the telescope and up to three LGS targets, all in sky coordinates. The action arguments will be used to configure and command devices within the laser facility; specifically:

- Prepare for low-power laser propagation

- Select the appropriate laser Beam Splitter configuration,

- Position the patrolling steering arms,

- Track the switchyard devices (steering mirrors and polarizers),

- Track the BGS devices (focus stage, asterism rotator and asterism tip-tilt).

The composite will wait for each device to complete the acquisition process. The collective result of the acquisition process will be returned by the action.

Propagate

The Propagate action is responsible for validating the asterism and projecting the high-power laser beams on sky. This action can only be executed when the telescope is tracking, and is typically performed immediately after the slew to the science target.

When initiated, this action will perform a set of steps to validate the asterism before and after switching to high-power laser propagation. If the asterism structure appears correct the final shutter will be opened and the lasers will be allowed to propagate on sky. Once this action completes the entire laser facility will be properly configured for observing and all tracking loops will be closed.

The combined result of the propagate action will be returned.

Halt

The Halt action is responsible for halting all devices in the laser facility, and is typically issued to interrupt or terminate a sequence. Any devices currently in motion will be immediately stopped, and all loops will be opened. The laser facility shutters will be put into a safe state, and propagation will be stopped at the switchyard. Note: as the AO system relies on the laser light to perform atmospheric correction the AO Facility controller must be halted as well. If performed through the observing tools, the MSCS will automatically take care of halting both systems.

Halt can be issued from any of the post-initialization states. From the HALT state the laser facility can be reinitialized or perform a new acquisition.

Recover

The Recover action allows the client (i.e. MSCS) to initiate a selective system re-initialization if a fault or error occurs. The recover command will be passed down the laser control system hierarchy to every NGAO component. If one of the components is in the fault state an initialization and recovery sequence will be executed to attempt to return the device to an operation state. If the recovery succeeds the device will be returned to its previous run-time state (e.g. TRACK). If the recovery fails the client may still be able to continue the sequence; however observing performance may be degraded. If the sequence has to be halted a full initialization should be performed in place of a recover.

#### Outputs

Subsystem output is provided through named telemetry topics. The following describe the telemetry provided by the subsystem, in addition to the standard outputs discussed in 4.1.2.

AsterismHealth

The AsterismHealth telemetry item defines the current status of the LGS asterism as determined by the Validation Task. This can be good, ill, or bad, where:

- Good – The asterism beacons are all accounted for and in their expected positions.

- Ill – The asterism beacons are all accounted for, but are not positioned where expected. Performance may suffer but observing can continue.

- Bad – One or more beacons are missing or occluded. Immediate action must be taken to safety the system.

### Control Loops

The following sections detail the control loops of the Laser Control Sequencer. Each control loop is implemented in part of in full by one or more of the NGAO components.

#### Asterism Validation

NGAO makes use of high-intensity lasers to produce the LGS beacons. If the lasers are not properly aligned they can cause significant damage to the laser facility hardware and injury to personnel. As such, the laser propagation must be periodically validated to ensure the proper alignment and performance of the system. An asynchronous task will be created to monitor the health of the laser system. If the task identifies any major changes to the asterism structure, such as the loss of an LGS beacon, an automated response will be taken to ensure the safety of the laser facility.

Figure 42: Asterism Validation Control Loop

The AsterismValidationController is designed to perform asterism validation and error handling during the acquisition and observing sequence. The control loop implements an image processing routine that takes images from the Asterism Imager Camera and compares the laser beacons against their expected positions in the field. These positions are based on the coordinates of the tip-tilt targets, position of the asterism tip-tilt mirror, and the orientation of the rotator. If there are any major changes to the asterism structure, such as the loss of an LGS beacon, the validation controller can take an automated response to ensure the safety of facility.

Three failure modes have been identified for the laser asterism and will be monitored by the validation controller. The failure modes are:

- Loss of a laser or beacon – If the asterism validation identifies that there are fewer than the expected LGS beacons in the image the sequence will be halted and the laser system will be put into a shuttered safe state.

- Additional beacons – Potentially the laser system could be configured incorrectly to produce additional beacons when a reduced subset is desired (e.g. fixed-pupil or a single TT target). If the validation task identifies more than the expected number of LGS beacons the sequence will be halted and the laser system will be put into a shuttered safe state.

- Incorrect beacon position – Potentially caused by a number of factors (offloading, flexure compensation, etc.) the laser beacons may drift over time. As all of the beacons are accounted for, only displaced, there is no imminent safety concern for the system. The validation task will assume the system is performing correctly, but will warn the operator that the asterism has changed and that performance could suffer.

The validation task is event driven and will be executed for each new frame produced by the asterism imager camera. The initial positions of the beacons will be calculated when the validation task is started based on the sky position of the tip-tilt targets. The positions of the beacons will be recalculated whenever the BGS rotator or tip-tilt mirror is moved, and should only require a translation and rotation matrix transformation. The validation task is started at the beginning of the science sequence and halted at the end of integrations.

### Devices

As a high-level subsystem, the Composite relies on the low-level subsystems for the majority of hardware control. Although the Laser Control Sequencer is capable of directly interacting with any hardware in the laser facility, there are only a few hardware devices it manages directly. Specific details, interfaces, and discussions are available in their related design documents.

#### Laser Safety System

The Laser Safety System (LSS) acts as a constant supervisory safety controller for the laser system, and must be running at all times the laser units and hardware is active. The LSS monitors a number of system and device states to ensure the lasers are operating within their defined safe limits. All real-time and safety critical device information is connected directly to the LSS allowing the system to immediately respond to protect the system and personnel from hazards. Non-critical and external system information is provided in the form of software permissives, and includes events from LTCS and operator acknowledgements.

See the Laser Safety Subsystem for more information on interfacing with the LSS.

#### Sky Alignment Camera

The Sky Alignment Camera exists primarily for system testing and laser alignment: this camera is not used during regular observing operations. The camera is physically located in the BGS, just after the beam expander, and can be used in either an up-stream (viewing asterism) or down-stream mode (viewing sky).  To use the camera, a fold mirror is inserted into the beam path configured to pickoff the light from the lasers or to view the current pointing alignment of the telescope.

The Laser Sequencer Subsystem is responsible for initializing the sky alignment camera. Control and configuration of the device is performed through the observing tools.

# Data Server

Although it is part of the AO Controls System, the Data Server can be and is being worked independently.  As such, the details of the Data Server design can be found in the separate document: KAON 732: NGAO Data Server Preliminary Design.

# Risk Analysis

The largest risk to the non-real-time controls system design is still the lack of specificity regarding the many devices to be controlled, both motion control devices and otherwise. In most cases we understand the type and number of devices to be controlled, but we still have not decided on and specific hardware device controllers or platforms. As a result, all the work estimates must have a high level of risk assigned until more information is available and the design can be specified.  However, since the System Design phase, we have tried to simplify the design by consolidating and specifying the minimum number of controller types for the AO controls and Laser controls, thus minimizing the work estimates as much as possible.

Also affecting the work estimates is the lack of details for calibration procedures for most aspects of the system.  Without enough details, good work estimates cannot be made.

Similarly, a choice still needs to be made on the machine architecture and operating system that we will use to host the NGAO control system. Currently we are considering Solaris and Linux machines, but there may still be requirements for using VxWorks or other platforms. Although the KCSF architecture should be able to operate on any OS and platform compatible with Java there are still questions about the compatibility of the underlying communication middleware (ICE or DDS) with certain architectures. This decision will impact which middleware architecture we will want to use, and which version (community or commercial) we will need to purchase.

Another fairly large risk is posed by the fact Erik Johannson (AO Controls Lead) has left WMKO and with him went a wealth of AO knowledge.  This has left a huge void for the AO Control Team and must be filled by the remaining team members coming up to speed on AO, which is difficult while continuing to work on the software design.  The NGAO project as a whole will also need to provide much more details when specifying tasks and algorithms for the AO Controls Team.  Also, Erik was allocated at the 80 - 90% level, while his replacement, Kevin Tsubota is supposed to be working up to the 50% level by the end of PD.  WMKO is currently in the process of hiring a Control Systems Engineer which will help with man power but none of the existing candidates have any AO experience.

# Issues Requiring Resolution During DD Phase

The following items require attention during the DD design phase:

- Device controllers, workstations, operating systems, and architecture need to be defined before an overall device control philosophy and architecture can be better specified.

- The middleware (DDS or ICE) must be selected.

- Subsystem ICDs need to be defined and documented.

- AO and Laser Controls telemetry streams needed by other systems need to be defined and documented.

- Calibration control and sensor calibration measurement algorithms need to be specified and documented for the AO Controls software team.

- The laser software interface must be defined (this is the software interface to the laser device delivered by the vendor).

- The RTC interface needs to be detailed.

- Fixed-pupil observing mode control loops must be defined.

- Configuration system must be designed.

- Develop KCSF prototypes to prove NGAO controls system design.

- Interface to Science Operation Tools.

- Functional Requirements

The following table is a snapshot of the Controls Infrastructure Functional Requirements included as a convenience for the reader.  The see the most current and complete set of requirements refer to the Contour Database.

| ID | Short Name | Description |
| --- | --- | --- |
| FR-456 | Definition | The overall controls system infrastructure is defined as all the infrastructure control functions that will be used throughout the AO Controls and LGS Controls system. Throughout the remainder of these requirements, the overall controls system infrastructure control functions will be referred to simply as "Controls Infrastructure". Controls Infrastructure shall provide the following basic functions for the AO and LGS Controls system: 1. Device control for all controllable devices. 2. Motion control for all controllable opto-mechanical devices not directly concerned with the high speed measurement and correction of atmospherically distorted wavefronts. 3. Configuration and control of the RTC. 4. Support for system configuration, calibration and operations. 5. Monitoring the health of all subsystems and providing notification to other systems as required. 6. Control of the AO and Laser enclosure environment. |
| FR-1839 | Health Monitoring | Controls Infrastructure shall provide a health monitoring function that periodically checks all the AO and LGS subsystems for status. If an error is detected, Controls Infrastructure will provide notification to the other NGAO subsystems and users. |
| FR-1840 | Acquisition Control | AO Controls will provide an automated acquisition capability to acquire all guide stars and science targets with limited operator oversight. |
| FR-1841 | Configuration Control | The Controls Infrastructure will provide an automated capability to configure the AO and LGS system for all modes of the system: startup, operational, calibration, and test. This includes configuring all the required cameras with the appropriate default parameters, configuring the appropriate devices on the optics bench, configuring the RTC, and configuring the AO related functions of the IFS (if necessary). The specifics of these configurations will be documented as use cases in the operations concept document. |
| FR-474 | Acquisition and Offload Control | AO Controls will provide functionality to interact with the telescope control system (currently DCS) to implement control for: Acquisition (possibly an interface to MAGIQ or son-of-MAGIQ) Offload (tip-tilt, focus and coma offloading) |
| FR-1828 | General Device Control | Controls Infrastructure shall provide a control function for all non-motorized devices in the system capable of computer control. The devices to be controlled include: detectors, mirrors, the RTC, and the environmental system for the both AO and LGS enclosures. The control function shall provide, at a minimum, power control, device configuration and basic device control, unless these functions are not available. The motorized devices are addressed in a separate requirement on motion control. |
| FR-1834 | Motion Control | Controls Infrastructure shall provide a motion control function for all opto-mechanical devices requiring remote computer control. The control function shall include basic device control, configuration control, position control, tracking control for those devices that require it, and support coordinated moves of multiple devices. |
| FR-1832 | Environmental Control | Control Infrastructure shall provide a configuration control function for all the environmental control devices in the systems that require remote computer control. |
| FR-1812 | Acquisition System data products | One hundred per night (TBC) acquisition images shall be stored on the NGAO data server. Images shall be in FITS file format and have standard FITS header information. The specific contents of the header are TBD. The data products from the acquisition system and the information from available catalogs or literature shall be recorded in the same photometry system and with comparable spatial resolution. |
| FR-1913 | Image Size and Profile Measurements | The acquisition system shall include tools to measure the FWHM and ellipticity of images, and to provide a profile of these images. |
| FR-3360 | Logging | The Controls Infrastructure shall provide logging of important system events, parameter changes, etc., to one or more log files. |
| FR-3361 | Alarms | Controls Infrastructure shall provide an alarm system which allows both the reporting of and response to abnormal system conditions. It shall provide both automated and manual response options, multiple levels of alarm severity, a centralized alarm manager, and flexible alarm system GUIs which implement sorting capabilities. |
| FR-3362 | Data Recording | The Controls Infrastructure systems shall support recording of parameter data, telemetry and diagnostics to the Data Server. |
| FR-3363 | Motion Control Tracking Rates | Controls Infrastructure shall support minimum sustained input rates of 40 Hz for tracking motion control applications. |
| FR-3364 | Data Types | The Controls Infrastructure shall handle transmission of arbitrary data types, both scalars and arrays, with data sizes up to at least 1 Mb. The control system must support arbitrary data types, including structures. Many RTC parameters are arrays, which are also large in size. The maximum array size used by the RTC is estimated to be 1 Mb. |
| FR-3121 | Data Server | The NGAO Data Server is defined as the system that provides all the data recorder and server capabilities required by the NGAO system. This includes all database support, logging, archiving, general storage and retrieval. The data server must support all the data storage and retrieval requirements of the NGAO system. |
| FR-3098 | Data Server - Definition | The NGAO Data Server is defined as a computer in the network, with associated software, that is dedicated to the data storage and retrieval needs required by the AO Controls System. It holds the database management system, the databases and is part file server capable of housing non-database files such as FITS. In addition to providing data persistence, retrieval and data management there will be a telemetry data recorder, an ad-hoc data recorder, an alarm recorder, general logger, and configuration. |
| FR-3132 | Data Storage Capacity | At a minimum the Data Server shall store at least three month's worth of collected data [TBC] with a buffer of 1 month’s worth. Once it reaches this limit, the oldest data of a configurable timeframe, i.e. 1 day, 1 week, 10 days, etc., will be automatically deleted. |
| FR-3367 | Data Recording | All data recording will be performed through supplied APIs. It is expected that the underlying control system framework will provide services to support data persistence to the various data stores. |
| FR-3368 | Time Stamping | All data stored in the data server should be time-stamped in a consistent manner. The recommendation is that UT is used for all time stamps. The timestamp should have at least a millisecond resolution or better. It is expected that the data server be clock synchronized with the rest of the AO Controls system. |
| FR-3099 | Support systems and services | The data server shall provide the database support for the following systems and services: a. The configurator b. The general logger c. The telemetry recorder d. The ad-hoc recorder e. The alarm recorder |
| FR-3369 | Telemetry Recorder | The Data Server shall provide a telemetry data recorder capable of recording time-stamped structured data. Grouped information will be stored and retrieved in an atomic and consistent manner. Ideally the telemetry recorder should not require the manual use or creation of meta data in order to function. The maximum burst and sustained rates for the telemetry recorder are TBD. |
| FR-3370 | Ad-hoc Recorder | The Data Server shall provide the capability to record any requested set of published AO Controls System data at the fastest rate that the data is published for any specified amount of time. It is expected that data can be archived at a particular scan rate or based on some configurable event. |
| FR-3371 | General Logger | The Data Server shall provide general AO Controls System logging. It will record all messages in a persistent store. The recording will include automatic timestamp and source identification There should be separate categories for any message type. Logging should be capable of being enable/disabled programmatically per component and should be capable of changing its logging level programmatically (log level can be used to mediate density). |
| FR-3372 | Alarm Recorder | The Data Server shall have the capability to record all published alarm data. An alarm recorder will be available to record all generated alarms and alarm transitions. The recorder can be enabled and disabled as appropriate. |
| FR-3373 | Configurator | The Data Server shall provide for the storage and retrieval of system configuration information. Calibration data should be made available for an indefinite period |
| FR-3374 | Auto Recovery | The Data Server shall auto recover from disk failures so an observing nights data is not lost. This could be through the use of a RAID system, disk duplication, etc… |
| FR-3375 | Scalable Data Storage | The data store shall be capable of scaling to increase the amount of data storage as needed. |
| FR-3376 | Permanent Archival Storage of Data | The data server shall support permanent archival storage of both raw and processed data. A permanent archive shall include removable media. The archival mechanism must be able to write to a removable media (tapes, hot-swap disks, or a future transportable media) and possibly HQ through its network interface. |
| FR-3377 | Extract and Save | The data server shall have the ability to save off portions of a database to external files for later processing. |
| FR-3103 | Simultaneous queries | The data server shall support simultaneous high-speed storage and recall (queries). |
| FR-3107 | Error handling | The Data Server shall provide error logging, alerts, and messages. It should also make its alarm visible to an operator during operations. |
| FR-3109 | Reliability | The Data Server shall have a MTBF of 10hrs (TBC) or a minimum of a full observing night. An unfortunate loss of the Data Server should not prevent the AO Controls System from continuing to run. |
| FR-1842 | Calibration Control | AO Controls will provide a semi-automated capability to calibrate the AO system, to include computation of DM actuator gains, DM to WFS registrations, centroid gains, interaction matrices, WFS calibrations, etc. The term semi-automated is used here to indicate that the calibration capability shall minimize the amount of user interaction required. |
| FR-470 | Sensor Calibration Measurements | AO controls will provide an automated capability to compute the dark frames, sky frames, and flat fields for all the various sensors in the AO system as required. |
| FR-1843 | Telemetry | AO Controls shall provide appropriate telemetry streams for the other AO subsystems, instruments, telescope control system, and top-level MCS, as required. The list of telemetry streams required by these systems is currently unknown; as they become defined, individual requirements will be written for each stream to be generated. |
| FR-479 | Atmospheric Profiler | AO Controls shall publish atmospheric turbulence profile (Cn2) information from the atmospheric profiler. |
| FR-1829 | Detector Control | AO Controls shall provide a configuration control function for all the detectors in the system. In addition to the basic device control functions, the detector control shall include the setting of exposure times and frame rates, bias fields and flat fields, filter and gain settings, and shutter control, if applicable. |
| FR-1831 | RTC Control | AO Controls shall provide a configuration control function for the RTC. In addition to the basic control functions, the RTC control function shall include the setting/readback of all RTC parameters, such as reconstructor arrays, centroid offsets, centroid gains, loop gains, etc., and loop control for all the AO correction loops in the system, includes both go-to and feedback loops. |
| FR-467 | Altitude tracking | AO controls must provide altitude tracking for the LGS WFS (in LGS mode). |
| FR-461 | Truth Wavefront Sensor | AO Controls is responsible for control of the TWFS and calculating low-order WFS focus correction for the RTC. |
| FR-3188 | Field Rotation Offload | The AO Controls shall compute the field rotation error from the average tip-tilt signal of the NGS low order wavefront sensors as output from the RTC telemetry. The field rotation error should be computed at a rate of 10 Hz (TBC). The AO Controls system will use this information along with Telescope position information to update the position of the AO field derotator. Offload data from the RTC includes the separate average tip/tilts from each NGS LOWFS channel. |
| FR-3365 | Loss of Communications | In the event AO Controls fails to establish communication or losses communication with any controller or service, other than a safety system, it shall generate an appropriate alarm and attempt to continue operations to the best of its ability or enter a safe state depending on the condition. The operator, upon receiving the alarm, will be able to take action as needed. |
| FR-2483 | Overall | This folder contains overall requirements that span multiple engineering disciplines |
| FR-2130 | Definition | The LGS control system is defined as all the LGS control functions that are not directly concerned with the high speed measurement and correction of atmospherically distorted wavefronts. Throughout the remainder of these requirements, the LGS control functions will be referred to simply as "LGS Controls". |
| FR-2131 | Basic Functionality | LGS Controls shall provide the following basic functions for the LGS facility: 1. A control system, supporting: Simple and complex commands Synchronous and asynchronous commands Command completion monitoring Error detection, handling, notification and recovery User interfaces supporting graphical, command line, and script modes 2. Device control for all controllable devices in the LGS facility. 3. Motion control for all controllable opto-mechanical devices in the LGS facility. 4. Configuration and control of the laser system. 5. Support for LGS facility configuration, calibration and operations. 6. Logging of all telemetry data to the data server. 7. Monitoring the health of all LGS facility subsystems and providing notification to other systems as required. 8. Control of the Laser enclosure environment. 9. Interfaces to the other NGAO and observatory subsystems: AO System LTCS Telescope Control System Data Server Multi-System Command Sequencer |
| FR-2143 | General Device Control | LGS Controls shall provide a control function for all non-motorized devices in the LGS facility capable of computer control. The devices to be controlled include: detectors, mirrors, and the environmental system for the Laser enclosure. The control function shall provide, at a minimum, power control, device configuration and basic device control (initialize, standby, start, stop, power up/down, read status, etc.), unless these functions are not available. The motorized devices are addressed in a separate requirement on motion control. |
| FR-2144 | Detector Control | LGS Controls shall provide a configuration control function for all the LGS facility detectors. In addition to the basic device control functions, the detector control shall include the setting of exposure times and frame rates, bias fields and flat fields, filter and gain settings, and shutter control, if applicable. |
| FR-2146 | Laser System Control | LGS Controls shall provide a configuration control function for the laser system. In addition to the basic control functions (initialize, standby, start, stop, etc.), the laser control function shall include the setting/readback of all laser parameters. |
| FR-2147 | Environmental Control | LGS Controls shall provide a configuration control function for all the environmental control devices in the laser enclosure that require remote computer control. |
| FR-2149 | Motion Control | LGS Controls shall provide a motion control function for all opto-mechanical devices in the LGS facility requiring remote computer control. The control function shall include basic device control (initialize, standby, start, stop, etc.), configuration control, position control, tracking control for those devices that require it, and support coordinated moves of multiple devices. As of this writing, the opto-mechanical devices requiring motion control are as follows (number of degrees of freedom shown in parentheses): Laser Shutter (1 DOF) Polarization waveplates (3 DOF) BTO bottom mirrors (4 DOF) BTO mid mirrors (4 DOF) BTO top mirrors (4 DOF) Shutter (1 DOF) Point and shoot beam splitter (2 DOF) Star imager pickoff (1 DOF) Point and shoot steering (6 DOF) Asterism rotator (1 DOF) Asterism tip/tilt (2 DOF) Fast shutter (1 DOF) Beam expander focus (1 DOF) LTO cover (1 DOF) LTO polarization sensor (1 DOF) LTO focus lens (1 DOF) As more information is known about the motion control needs for these devices and subassemblies, individual requirements will be added to address them. |
| FR-2150 | Motion control coordination | LGS controls will coordinate the LGS facility motion control and other tasks needed to perform dithering, offsetting and chopping for science observations. |
| FR-3366 | Loss of Communications | In the event LGS Controls fails to establish communication or losses communication with any controller or service, other than a safety system, it shall generate an appropriate alarm and attempt to continue operations to the best of its ability or enter a safe state depending on the condition. The operator, upon receiving the alarm, will be able to take action as needed. |
| FR-2484 | Optical | This folder contains optical requirements |
| FR-2132 | Command Processor | LGS Controls shall provide a command processor (sequencer) with the following functionality: 1. Processing of simple commands. An example of a simple command is the setting of a single parameter. 2. Processing of complex commands. An example of a complex command is to configure the LGS facility for a particular mode of operation. This will require setting several parameters, configuring any optical devices, and checking for valid completions of all. 3. Processing of synchronous commands. A synchronous command is one where the latency of the task to be performed by the command is negligible. 4. Processing of asynchronous commands. An asynchronous command is one where the latency of the task to be performed is large enough to require spawning a thread to execute the task and monitor for a command completion callback. 5. The return status of every command (or sub-command in the case of complex commands) must be checked for valid completion or error. If an error occurs, appropriate actions must be taken to process the error, abort the command or task if required, notify the requestor of the error, and take action to recover from the error, protecting the LGS facility as required. 6. A command interface that supports graphical, command-line, and script modes. |
| FR-2133 | Health Monitoring | LGS Controls shall provide a health monitoring function that periodically checks all the LGS facility subsystems for status. If an error is detected, LGS Controls will take appropriate actions to protect the LGS facility and provide notification to the other NGAO subsystems and users. |
| FR-2135 | Configuration Control | LGS Controls will provide an automated capability to configure the LGS facility for all modes of the system: startup, operational, calibration, and test. This includes configuring all the required cameras or sensors with the appropriate default parameters and configuring any optical devices as required. The specifics of these configurations will be documented as use cases in the operations concept document. |
| FR-2136 | Calibration Control | LGS Controls will provide a semi-automated capability to calibrate the LGS facility, to include alignment of all movable LGS optical devices. The term semi-automated is used here to indicate that the calibration capability shall minimize the amount of user interaction required. |
| FR-2138 | Telemetry | LGS Controls shall provide appropriate telemetry streams for the other AO subsystems, instruments, telescope control system, and top-level MCS, as required. The list of telemetry streams required by these systems is currently unknown; as they become defined, individual requirements will be written for each stream to be generated. |
| FR-2141 | Copy of Interface to LTCS | LGS controls will respond appropriately to protect the LGS facility and recover gracefully from a laser traffic control event. |