AO bench setup and checkout

Pre-LGS

- Setup chopper

  - Verify “Chopper Setup” has been completed (http://www.oir.caltech.edu/twiki_oir/pub/Palomar/PalmLGS/CassCage/Chopper_Setup.doc)

  - On the delay generator in the coude lab, detach the laser trigger cable attached to the Master Input and connect the internal 500Hz signal to the the Master Input.

  - From a PC, open up a remote desktop

    - Enter 198.202.125.223  The username is aousr and the password is the same as the ao system.

      - If the connection doesn’t work, verify that the blue PC in the Coude room is powered up.

    - Startup DG100 from the desktop.

    - The most recent settings are listed in: http://www.oir.caltech.edu/twiki_oir/pub/Palomar/PalmLGS/AOOperationalProcedures/lgs_delay_generator.xls

    - Click “Update Settings” in the upper right-hand corner

  - Have the BTO operator verify the 500Hz signal

  - Open up the “chopper.ht” hyperterminal program on the WFS PC.

    - Set the chopper to run off an external trigger (type, r and x.  Type e to verify it works)

- Verify the laser trigger in the Cass cage is connected to the HOWFS trigger input

  - Check if HOWFS will trigger off delay generator signal

- Put the laser trigger cable back into the master input on the delay generator

AO Checkout:

- Co-focus system (http://www.oir.caltech.edu/twiki_oir/pub/Palomar/PalmLGS/TestPlans/PALAO_internal_focus_v4.doc)

- Test LOWFS acquisition off axis