Reviewer comments on Alignment, Calibration & Diagnostics mini-review

Don Gavel

Opto-mechanical design

- In your consideration of the refractive option, did you talk with Brian Bauman? He has designed a refractive projector for the GPI telescope simulator that is co-focal from mid-optical to 2 microns. There may be an advantage as it allows for two pupils in free space

[TS]:  I have not talked with him but I will.

- Will the pupil mask rotate? This would be to test the reconstructor’s ability to deal with a rotating pupil.

[TS]: This was not currently planned, but probably could be included.  The area around the pupil is tight in the current design.

- I assume the simulator is inside the cold enclosure (this was not explicitly stated). Is this a requirement?

[TS]:  Yes, it is inside the cold enclosure.  This was not a requirement, but a response to packaging problems.  Placing it outside would mean that it is farther away from the AO focal plane, resulting in either larger optics or the need for extra optics close to the rotator to collimate the beams.  Placing it inside in a temperature controlled environment also helps with the overall system stability.

Alignment Plan

- The AO optical bench could allow for the first turning mirror to be adjustable so as to correct for slowly varying telescope pupil misalignments.

[PW]: This is an interesting idea, unfortunately it also moves the beam on the 1st OAP and also moves the image.  We should determine if the resultant aberrations are worth the reduced pupil motion.

- Section 2.4: The instrument feed switchyard mirror could be used to adjust pupil alignment in the instrument.

[PW]: Same response as for the previous suggestion.

- End of section 2.4: putting the telescope simulator in the cold enclosure should mitigate its thermal drift.

[PW]: It is in the cold enclosure.

- 3.7: In the present design concept (from JPL), the LOWFS do not have a common mounting plate. Perhaps struts or metered bars could be used to force their co-alignment.

[PW]: Thanks for pointing this out.  We will have to come up with a solution to aligning all of these.

Calibration Plan

What is the provision for image sharpening given that Davinci will not be adjustable in focus?[CN]: The report mentions two possibilities 1) moving the calibration source mechanically in z 2) Using the LODM, currently 1) is the preferred choice.  The calibration source supports the z motion.

4.3: It’s not clear that image plane calibration will work in IFU mode – the spaxel size is large compared to the diffraction limit. Perhaps power in the bucket, but likely very insensitive.[CN]: KAON 739 presents two solution to the IFU problem that probably will need to be used together 1) In the case of undersampled images, a higher resolution image can be synthesized from sub pixel stepping of the image across the detector.  2) Phase diversity methods could be used on the monochromatic images within the “data cube” of an integral field spectrograph.  Marcos van Dam, Thomas Stalcup and Randy Campbell have done some work on this with OSIRIS.  It is probably an area for the DAVINICI team to follow up on.

- 4.5: Good point. The position of the input tip/tilt (for calibrating the WFS) must be well known. Davinci could be used in the narrow field to help calibrate the motion, LOWFS detectors in the wide field. .  [CN]: Thanks for the suggestion

- 4.2 Astrometric grid. Another possibility is to use diffraction from a wire grid in a pupil plane. [CN]: Thanks for the suggestion; we selected the NIRC2 style grid because it was a known technology that has been test by Brian Cameron.

- 4.5 Comments on LGS WFS calibration:

  - I suggest putting the LGS sources on z-translation to simulate “elongation”

[TS]:  Interesting.  Are you suggesting to move them back and forth at a high rate during exposures?

  - Calibrating the open-loop WFS for open loop measurement has been demonstrated in the lab, but not on the sky yet. [CN] I used this reference: Ammons, et al, “Integrated Laboratory Demonstrations of Multi-Object Adaptive Optics on a Simulated 10-Meter Telescope at Visible Wavelengths” I think this was an LAO experiment by Mark, is that incorrect?

  - It would be nice to have a set of aberrator plates that have known low-order aberrations for testing calibration.  This would be in addition to atmospheric aberrator plates. LAO may be able to make these. [CN] We proposed using the LODM as a way to be able to run the same test, it’s at a pupil so all the LGS WFS should “see” the same aberrations.  I believe we will have the ability to put some aberration plates in the calibration source.  Thomas plans to discuss this issue tomorrow.

4.7 DM Calibration: I favor the insertable Hartmann sensor idea (doesn’t need interferometric reference wave, can use the existing beam, and it’s a lot easier than taking the MEMS out). Hartmann has two n’s by the way. The AO relay optomechanical design needs to be revisited to make sure we can supply the kinematic fixtures. [CN] I’ll note that as an issue to follow up on before PDR, can you [DG] take the lead and talk to Chris Lockwood and Reni Kupke about adding the kinematic fixtures?

4.8 DM registration: I racked my brain trying to think of a way one could determine registration given only a focal plane image. No real sensitivity. The suggestion of putting a pattern on the LODM and correcting it with the HODM seems reasonable, but not sure if it would be sensitive enough at the 1/10’th high-order subaperture regime. This needs quantitative analysis. I highly suggest putting some sort of pupil viewer, even if mounted outside Davinci.  [CN] The team (PW, TS, & CN) suggested that NGS WFS have a “pupil view” mode it is requirement now (FR-3445), it is behind the HODM and LODM so it could be used to verify the registration.

- 6: Step-by-step process is a good start. I would like to see this section become more quantitative.  [CN] The team would like to expand this and the corresponding sections for Rich’s error budget in the early part of the Detailed Design phase.