Keck Next Generation Adaptive Optics

Opto-mechanical Design Pros and Cons

Donald Gavel

Original December 3, 2007

Updated December 6, 2007 (DG)

Updated December 10 & 12, 2007 (DG and PW)

The following is a tabular comparison of the two options for KNGAO opto-mechanical design, with pros and cons indicated. The 2-tier design is Brian Bauman’s initial layout as reported at the completion of the AO System Architecture phase.  The 1-tier design is a layout by Peter Wizinowich, using Brian’s 2nd relay.

Conclusion: We have decided to proceed with the 2-tier design primarily because of the smaller pupil size and the maintenance of the MCAO option.  The decision to proceed with the 2-tier design is contingent on finding a solution for the Interferometer (item 5 below) and on completing a layout that demonstrates that the science instruments can be accommodated.  The 1-tier design will be maintained as a fallback in the event that these conditions cannot be satisfied.

Refer to the AO System Design web page for Zemax prescriptions and details of these layouts.

| # | Item | 1-tier | 2-tier | Importance | Cost Impact | Performance Impact |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 1st fold mirror conjugate | Con: At ~20km. Note: OAP 1 is conjugate to 12km, so the mitigation is a curved DM | Pro: At ~10km, a possible location for a 2nd MCAO DM | Low to Medium (depending on import-ance of fallback) High (if a science case for MCAO) | 1-tier: High if need to make OAP1 a DM | No |
| 2 | Clearance between K-mirror derotator (M2) and 1st OAP | Pro: ~130mm | Con: Closer. There is an orientation of the K-mirror where the clearance is about 50mm. This assumes the K-mirror must rotate through a 360 degree angle | High (if doesn’t fit) | 2-tier: Medium if makes rotator design difficult | No |
| 3 | OSIRIS feed | Pro: The same feed, as from the current AO system, can be accomplished after the 1st relay Note for either: Fitting error larger if don’t go through 2nd relay (SR(J) = 40 vs 54% for on-axis 15th mag t/t star near zenith); partially mitigated by at least 5 more reflections for 2nd relay, including beam reformatting mirrors. | Con: beam from the AO relay is not at the same vertical height as the OSIRIS beam line Note for either: Could reformat to better match internal OSIRIS pupil (although this may be fixed independent of NGAO) | High | 2-tier: Medium due to additional relay & OSIRIS redesign | Yes for SR if only use 1st relay Yes for background if go through 2nd relay |
| 4 | Tip/Tilt Perfor- mance | Con: larger beam will require larger tip/tilt mechanism; may eat in to tip/tilt bandwidth. A custom design for the TT stage can be made to meet requirements. Mass interacts with bandwidth requirement. Note: Can meet performance requirements | Pro: smaller beam means smaller tip/tilt mechanism and less mass, easing the ability to meet the bandwidth requirement. Note: bandwidth requirement is high (~90Hz) driven by the need to control for telescope vibration. | High | 1-tier: Medium due to more expensive tip/tilt mechanism for 140 mm pupil | 1-tier:Yes |
| 5 | Interferometer feed | Pro: reflections/polarizations and field rotation into the interferometer are identical to the current system | Con: reflections/polarizations and field rotation angle are probably different and would require extra reflections for compensation | High DG: needs clarification | 2-tier: Medium to high due to additional relay & beam formatting or other mitigation option | 2-tier: Yes (for IF) |
| 6 | Overall footprint | Note: 2.3 x 2.2 meters. | Note: 2.2 x 1.8 meters. | Medium | Low | No |
| 7 | Beam size in 1st relay and deformable mirror actuator pitch | Note: Must be 140 mm beam. Note: Either 7mm pitch or 5mm pitch DM can be used (5mm pitch would require 2x actuators). Note: Pupil registration prefers 140 mm | Pro: Either 100mm or 140mm beam is ok. 140 mm beam will work, but layout will scale in size accordingly Note: Pupil registration prefers 140 mm | Medium | 1-tier: Medium for larger optics | 1-tier: Yes Larger dichroics – coating uniformity? Overall stability of larger mounts and path lengths? |
| 8 | Deformable mirror | Pro: Could re-use the existing Keck AO DM (349 actuator 7mm pitch), but this is mitigated by the considerable hysteresis which will have an unknown effect on the hybrid open-loop/closed-loop control law. Note: Could use newer technology low-hysteresis actuators. | Con: 100mm pupil forces the procurement of a new DM. Note: Could use newer technology low-hysteresis actuators. | Medium | PW: Low DG: Medium to High – cost of additional actuators in 140mm DM if using 5mm spacing | No |
| 9 | Space for d-IFS | Pro: 300 mm of more vertical height for an up-looking d-IFS. Pro: more space between dichroic and focus (~ 150mm more) for ADC, acquisition pickoff, etc. Note: Side looking preferred. | Con: 300 mm of less vertical height for an up-looking d-IFS in current design (could be 300 mm more if design different) Con: limited lateral space will demand d-NIRI have a “snout.” Note: Side looking preferred. | High | 2-tier: Medium if makes design difficult | No |
| 10 | Access-ibility | Pro: 1-tier design allows for easier assembly and alignment. Note: Unclear at this point which design offers the best operational access. | Con: 2-tier design will demand a space-frame or other additional structure to hold second level of optical elements. More costly design & potentially more difficult to align. Pro: 1st relay is more compact. Note: Unclear at this point which design offers the best operational access. | Medium | 2-tier: Low to medium for space-frame | No |
| 11 | Space for LGS WFS | Note: LGS WFS is located on a 2nd tier above 1st relay | Note: LGS WFS is located on a 2nd tier below part of 1st relay | High | Low | No |
| 12 | Simul-taneous Instruments | Pro: All instruments (d-IFS, narrow cameras, OSIRIS or DSM) can be fed without interference. | Not clear yet how instruments will be fit. | Medium |  |  |