Keck Adaptive Optics Note 971

Near-Infrared Tip-Tilt Sensor System:

Pre-Ship Review Report

R. Dekany, R. Smith, D. Hale, J. Fucik, H. Rodriguez, J. Cromer

DRAFT Version 0.2

July 31, 2013

# Document Revision History

| Revision Number | Revision Date | Summary of Changes | Author |
| --- | --- | --- | --- |
| 0.1 | 3/2013 | Draft outline of KAON 971 including assignments for engineer contributions | R. Dekany |
| 0.2 | 7/31/2013 | Interim draft capturing initial inputs from Fucik, Croner, and Hale | R. Dekany |

# Introduction

The purpose of this KAON is to summarize and assess the completeness and readiness of the Near-InfraRed (NIR) Tip-Tilt Sensor (TTS) opto-mechanical system for delivery to WM Keck Observatory in preparation of its subsequent summit installation.

The Acceptance, Integration, Test and Commissioning plan (KAON 855, section 2.1) states that each NIR TTS subsystem shall include the deliverable of

”An acceptance review document including:

- Compliance matrix for function requirements

- Compliance matrix for interface requirements

- Compliance matrix for system requirements

- Documentation to support the requirements and interface compliance”

This document constitutes this deliverable for the Caltech deliverable components of the Camera subsystem consisting of the following excluding the Filter Changer component which is provided by WKMO:

# Report Organization

We choose to organize this document according to broad areas of requirements compliance as follows:

Section 3	System Description Summary

Section 4	Optical Performance[Jason]

Section 5	Thermal Performance [Roger]

Section 6	Electronics Performance [Roger]

Section 7	Detector Performance [Dave / John]

Section 8	Software Performance [Dave / John]

Section 9	Shipping Readiness [Hector]

We include the formal compliance matrices as Appendices to this document and reference their corresponding technical subsection for each of cross-referencing.

The following items are not ready for shipment to WMKO, accompanied by a suggested course of action to bring each to shipment readiness:

- Pre-ship Exception #1:  [status and proposed plan]

- Pre-ship Exception #2 [ status and proposed plan]

- …..

# System Description Summary

For reference, the mechanical configuration of the NIR TTS is shown in Figure 1.  The basic construction of the NIRTTS camera consists of an optomechanical base holding field lenses and a fold mirror that directs light from the plane of the Keck AO bench upward into a lens tube relay and thus onto an infrared sensitivity detector array.  The lens tube and detector are cooled to cryogenic temperatures by a Joule-Thompson gas expansion nozzle closed-cycle cooler, as are the optics in a rotating wheel that selects for use the observing filter (onto which is defined the camera pupil.)

Figure 1: NIR TTS Mechanical Overview (from KAON 890)

For additional detail pertaining to the opto-mechanical design details of the NIR TTS see KAON 892 Optical Design and KAON 890 Camera Opto-Mechanical Design, which have been updated to accurately reflect the as-built system design.

# Optical Performance [Jason]

## Overview

The NIRTTS camera is an f/7 diffraction limited re-imager that operates in the H (1.5μm – 1.8μm) and Ks (2.0μm – 2.3μm) Infra-Red (IR) photometric bands.  The nominal camera design gives a mean Strehl ratio of 92% over the 120” Field of View (FOV) and >85% for an optical system aligned to better than ±25μm for lateral/axial displacements and ±2.4’ for angular displacements.

Figure 2: NIRTTS Camera Optical Layout.  FL1 and FL2 are assumed to be operating at a temperature of 273K at 0.6atm.  C1 is assumed to be at an operating temperature of 100K in vacuum.  C2, C3, and C4 are assumed to be at an operating temperature of 110K in vacuum.  There are three element groups in the optical prescription:  G1 is the field lens group that reforms the Adaptive Optics (AO) System pupil onto the back surface of the filter, G2 is the air space triplet camera that re-images the AO image plane that is 100mm behind the vertex of FL1 onto the Teledyne H2RG HgCdTe detector, and G3 is the field flattener that corrects for the field curvature in G2.

The camera is composed of three lens groups: field lens (G1), air-spaced triplet camera (G2), and field flattener (G3) with a total of six refractive optical elements with power. Figure 1 shows the optical layout of the NIRTTS camera where the optical elements are all spherical surfaces manufactured from Barium Fluoride (BaF2) or Infrasil fused silica glass material.  The change in the index of refraction and radius of curvature as a function of temperature for the various materials was accounted for in the design of the camera, which is detailed in KAON890.

## Functional Compliance

## ISP Optics Manufacturer’s Report

The six-element camera optics was polished and Anti-Reflection (AR) coated by ISP Optics Corporation.  Table 1 gives the Quality Assurance (QA) inspection report supplied by the vendor where the measurements where made in lab ambient conditions (i.e. T=293K and P=1atm).  Figure 2 shows the as measured AR coating reflectance curves for both the fused silica and BaF2 substrates.  The fused silica optical elements have an average reflectance of 0.3% and the BaF2 optical elements have an average reflectance of 0.4% across the instrument passband of 1490nm to 2310nm.  Both AR coatings meet the requirement of the reflectance to be less than 0.5%.  Both coatings were measured at room temperature, but it is expected that the edges of the AR coating will shift 60nm towards the red part of the spectrum at cryogenic temperatures.  This shift in the AR coating edges at operational temperatures was accounted for in the design by oversizing the passband of the AR coating.

| #### Element | Radius [mm] | CT [mm] | ET [mm] | Dia. [mm] | Glass | Wedge [arcsec] | Irregularity PV / RMS [λ@633nm] | RMS Roughness [nm] | Scatch/Dig |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| FL1 | S1: 288.615CX S2: 8361.572CX | 20.058 | 14.784 | 108.030 | BaF2 | 19.1 | S1: 0.23 / 0.03 S2: 0.50 / NaN | NaN | 40/20 |
| FL2 | S1: 178.411CX S2: 411.774CC | 15.134 | 10.058 | 108.026 | BaF2 | 19.1 | S1: 0.19 / 0.03 S2: 0.30 / 0.06 | S1: 4.5 S2: 2.5 | 40/20 |
| C1 | S1: 46.975CX S2: 355.217CC | 6.454 | 3.909 | 32.090 | BaF2 | 32.1 | S1: 0.21 / 0.02 S2: 0.15 / 0.03 | S1: 2.9 S2: 2.5 | 40/20 |
| C2 | S1: 139.337CC S2: 39.578CX | 13.056 | 8.855 | 38.204 | BaF2 | 43.2 | S1: 0.17 / 0.02 S2: 0.10 / 0.02 | S1: 4.3 S2: 1.7 | 40/20 |
| C3 | S1: 24.727CC S2: 34.535CX | 13.004 | 9.801 | 41.482 | Infrasil | 49.7 | S1: 0.06 / 0.01 S2: 0.15 / 0.03 | S1: 1.3 S2: 1.3 | 40/20 |
| C4 | S1: 41.686CC S2: 162.924CC | 9.834 | 17.623 | 45.990 | Infrasil | 26.9 | S1: 0.15 / 0.04 S2: 0.15 / 0.03 | S1: 1.5 S2: 1.9 | 40/20 |

Table 1: ISP Optics QA Inspection Report Summary

Figure 3: ISP Optics Corporation AR coating reflectance curves for fused silica and BaF2 substrates.  Both AR reflectance curves were measured at 6° Angle Of Incidence (AOI) at room temperature.  The edges of the AR reflectance curve are estimated to shift about 60nm to the red part of the spectrum at cryogenic temperatures, but this was accounted for in the coating design for which the AR coating passband was oversized.  The fused silica AR coating has an average reflectance of 0.3% and the BaF2 AR coating has an average reflectance of 0.4% across the instrument passband of 1490nm to 2310nm.

Suggested Actions: None

## Asahi Filter Manufacturer’s Report

Asahi Spectra supplied the two IR photometric filters, which are multilayer dielectric coatings on an Infrasil fused silica glass substrate that is 38mm in diameter and 5.0mm thick.  The H band IR photometric filter has a Full Width of Half Max (FWHM) of 287nm and a Center Wavelength (CWL) of 1634nm with an average transmission of 92% over the passband.  Similarly, the Ks band IR photometric filter has a FWHM of 320nm and a CWL of 2150nm with an average transmission of 94% over the passband.

Figure 4: Asahi Spectra H and Ks IR Photometric Filters Transmission Curves.  The dashed line curve is the atmospheric transmission at Mauna Kea at 1.6mm precipitable water vapor at 1.5 airmass.

Suggested Actions: None

## Off-the-Shelf Optical Components

There are two major off the shelf optical components namely the instrument window and the fold mirror.  The instrument window substrate was purchased from CVI Melles Griot part number PW-2025-INF, which has since been discontinued.  The window is an Infrasil fused silica substrate that is 50.8mm in diameter and 6.35mm in thickness.  It has a 10-5 scratch/dig surface quality with a λ/10 transmitted wavefront error at 633nm and less than 10” of wedge.  The substrate was AR coated by ISP Optics Corporation with the same coating used for the Infrasil camera elements in Section 1.1.1.  A spare window substrate was purchased and coated for the NIRTTS project.  The fold mirror was purchased from Edmunds Optics part number NT48-177-566.  The fold mirror is a Zerodur substrate that is 76.2mm in diameter and 19.1mm in thickness.  It has a 60-40 scratch/dig surface quality with a λ/20 wavefront error at 633nm.  The fold mirror has a protected (i.e. Silicon Monoxide overcoat) gold metallic mirror coating with a 97% average reflectance from 800nm to 3000nm.

Suggested Actions: None

## Interface Compliance

## System Compliance

## As-built Documentation

KAON 890 Camera Opto-Mechanical Design contains detail on the as-built mechanical design of the NIR TTS Camera.  KAON 892 contains detail on the as-built optical design.

| #### Element [#] | #### Name | #### Radius [mm] | #### CT [mm] | #### ET [mm] | #### Dia. [mm] | #### Material |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Field Lens 1 | #### 288.510 CX | #### 20.051 | 14.779 | #### 107.991 | #### BaF2 |
|  |  | #### 8358.537 CX | #### 12.829 | #### 21.372 | #### 107.991 | #### Al 6061 |
| 2 | Field Lens 2 | #### 178.346 CX | #### 15.129 | #### 10.055 | #### 107.987 | #### BaF2 |
|  |  | #### 411.625 CC | #### 190.000 | #### 186.705 | #### 103.958 | #### Al 6061 |
| 3 | Fold Mirror | #### Plano | #### 64.000 | #### 64.000 | #### 76.200 | #### Al 6061 |
| 4 | Window | #### Plano | #### 6.400 | #### 6.400 | #### 50.000 | Infrasil |
|  |  | #### Plano | #### 16.603 | #### 19.419 | #### 50.000 | #### Al 6061 |
| 5 | Camera 1 | #### 46.836 CX | #### 6.435 | #### 3.897 | #### 31.996 | #### BaF2 |
|  |  | #### 354.162 CC | #### 11.802a | #### 11.523c | #### 28.200 | #### SS 17-4 GP1 |
| 6 | Filter | #### Plano | #### 5.000 | #### 5.000 | #### 32.000 | #### Infrasil |
|  |  | #### Plano | #### 10.202b | #### 9.491d | #### 32.000 | #### Al 6061 |
| 7 | Camera 2 | #### 138.938 CC | #### 13.019 | #### 8.829 | #### 28.084 | #### BaF2 |
|  |  | #### 39.464 CX | #### 2.649 | #### 3.830 | #### 38.094 | #### SS 17-4 GP1 |
| 8 | Camera 3 | #### 24.727 CC | #### 13.004 | #### 9.801 | #### 26.084 | #### Infrasil |
|  |  | #### 34.535 CX | #### 98.568 | #### 99.192 | #### 41.482 | #### Invar |
| 9 | Camera 4 | 41.686 CC | #### 9.834 | #### 17.623 | #### 44.060 | #### Infrasil |
|  |  | #### 162.924 CC | #### BFL* | #### 3.232 | #### 43.990 | #### Invar |
| * BFL = 4.724 |  |  |  |  |  |  |
| a = 13.002 in K band |  |  |  |  |  |  |
| b = 9.002 in K band |  |  |  |  |  |  |
| c = 12.723 in K band |  |  |  |  |  |  |
| d = 8.290 in K band |  |  |  |  |  |  |
| NOTE: Physical outer diameter of Camera 4 is 45.990 |  |  |  |  |  |  |

Table 2: NIRTTS Camera Optical Prescription at Operating Temperatures

## Lens Prescription at Temperature

After receiving the “as-built” optical surfaces from ISP Optics Corporation the values in Table 1 were inserted into the Zemax optical model and then the air spaces were re-optimize to deliver a maximum Strehl ratio over the FOV of the camera.  Since all the optical surfaces were manufactured to within their allowed manufacturing tolerances, no significant changes were made to the optical model from the prescription delivered at the Detail Design Review (DDR).  Table 3 gives the optical prescription of the camera at operational temperatures for the “as-built” design, where the lens elements in Table 1 have been scaled for temperatures of 100K-110K.  Figure 5 shows the performance of the nominally aligned camera at operating temperature in the H and Ks bands when the camera is behind the Keck AO system.

Figure 5: Nominal Prescription Strehl Ratio Map when the NIRTTS Camera is behind the Keck AO system.  (Left Panel)  Gives the H band Strehl map as seen by the Teledyne H2RG detector with a peak of 95.5% and an rms over the FOV of 90.6%.  (Right Panel) Gives the Ks band Strehl map as seen by the detector with a peak of 98.6% and an rms over the FOV of 94.0%.

## Roll-pin Flexure Cryogenic Lens Mount

One of the major changes in the design since DDR was the lens mount design for the BaF2 lens elements at cryogenic temperatures.  Do to uncertainties in understanding the stresses involved in a lens element that is bonded to a metal mechanical cell with an epoxy joint the project decided that it would be less risky to go with an all mechanical mount design for the relatively small (i.e. ~50mm in diameter) optical elements.  The design of the mount was an adaptation from the cryogenic mounts used in the WIYN High Resolution Infrared Camera (WHIRC).  The detail design of the cryogenic lens mount as it is implemented in NIRTTS can be found in KAONTBD.  Solid Concepts using a Direct Metal Laser Sintering (DMLS) manufacturing method, which is a 3D printing of metal materials technology, fabricated the lens mounts.  DMLS manufacturing is a form of rapid prototyping of parts, which saved both in cost and time.  Also, do to the complex geometry of the mount flexures it was a suited candidate for the DMLS manufacturing method that could otherwise not be accomplished with traditional CNC fabrication methods.  It should be noted that the cells could have been manufactured with a combination of CNC fabrication and wire EDM manufacturing for the flexures themselves.  Currently, wire EDM manufacturing can achieve 2.5μm feature resolution, but at a higher cost and schedule time.  However, the 25μm feature resolution of the DMLS process was accurate enough for this particular design implementation.

After receiving the parts from Solid Concepts a few post manufacturing steps were taken prior to installation of the lens element. Since the DMLS is an additive process that is based on sintering fine powdered layers that are 30μm thick together it creates a part with a surface roughness of 15-40μm P-V, which is equivalent to a 320 – 600 grit.  The surfaces that come into contact with the lens element were polished to avoid scratching or chipping of the brittle BaF2 glass material during installation.  Also, the lens cells were designed to give a particular preload force onto the lens element, which is achieved by an interference fit between the mechanical cell and lens element.  In this particular case a 50μm radial interference fit was designed, such that the lens would be constrained from moving more than the required 25μm and the stress imparted to the lens be less than the required 5MPa.  However it is almost impossible to manufacture (by any method) the distance between two opposing springs to within 25% of the designed interference fit do to the uncertainties of how the spring will relax post machining.  For DMLS manufacturing, the relaxation of the spring is caused by the localized heating and cooling of the metal material on a thin wall section that is 380μm - 630μm thick.  In practiced it was found that the radial roll pin springs had to be pushed back in order to achieve the desired preload force.  A custom mandrel was fabricated with a shaft diameter 100μm greater than the as fabricated lens element diameter.  The mandrel was then inserted into the roll pin flexure mount and by turning the mandrel in the cell it burnishes the roll pin flexures.  The amount of burnishing was done experimentally till the desired spring separation between opposing springs was achieved for the designed interference contact between the cell and lens.

In order for the mechanical lens mounts to achieve the desired positional accuracy in the presence of frictional forces a low coefficient of friction coating needed to be added to the surfaces of the mount that come into contact with the lens (see KAON TBD for more detail).  The mechanical lens cells were fabricated from a 17-4 Stainless Steel alloy, which limited the coating options.  The WHIRC lens mount design (Smee 2011) was fabricated from an Aluminum alloy, which used a black Teflon impregnated anodized coating.  An anodization process is not possible with a Stainless Steel alloy material, but could be coated with a Polytetrafluoroethylene (PTFE) (i.e. Teflon) material, which would be vacuum compatible.  For cost and schedule reasons, the project decided to use Kapton tape with an acrylic adhesive that has a thickness of 75μm.  The Kapton tape provided two technical advantages over the PTFE coating, which was a lower coefficient of friction and uniform thickness.  Kapton tape comes in two varieties:  silicon adhesive and acrylic adhesive.  Silicon adhesive has been known to migrate over time in cryogenic vacuum environments, which is not known to be a problem with acrylic adhesive.  With no coating the radial and axial positioning accuracy of the lens mount was calculated to be ±40 - 60μm and ±14 - 25μm respectfully and ±2 - 9μm and ±0.5 - 3μm with a low coefficient of friction coating.  Camera #1 lens cell was tested for its positional accuracy by pushing the lens a distance of 150μm with a delrin tipped set screw and then releasing the set screw to see if the lens returned to its original position.  A microscope with a 50x magnification and a fiducial scale with a resolution of 25μm/1div was used to measure the displacements.  There was no noticeable change in the position of the lens after perturbing the lens, which indicated the lens cell was functioning as designed.

Suggested Actions: None

Figure 6: Roll Pin Flexure Mount for BaF2 Lens Elements at Cryogenic Temperatures.  Mount is composed of two elements: the roll pin flexure cell that constrains the radial motion and the leaf spring disk that constrains the axial motion.

Figure 7: Kapton Tape applied to Camera#2 Lens Cell as a Low Coefficient of Friction Coating.

## Metrology Report

## Profilometer Measurements

## Faro Arm (CMM) Measurements

## Image Quality

## Precision Projector Lab Overview

## Extra-Focal Imaging

## Data Reduction

## Strehl Map

## MTF Curves

## Instrument Throughput

## Optical Compliance Matrix

## Outstanding Issues

## Detector Offset from Camera FOV

## Stray Light

## Astigmatic Corrector

# Thermal Performance [Roger / Hector]

## Functional Compliance

Measurements were made in the lab to determine compliance with the following thermal functional requirements calling for experimental verification:

- Requirement #XXX

Measured Performance:

  - We measured…

Suggested Actions: None

## Interface Compliance

## System Compliance

## As-built Documentation

KAON 902 contains detail on the as-built thermal design.  KAON XXX contains detail on the Assembly and Alignment procedures including the thermal connections internal to the NIR TTS cryostat.

# Electronics Performance [Roger]

## Functional Compliance

Measurements were made in the lab to determine compliance with the following thermal functional requirements calling for experimental verification:

- Requirement #41: The NIR TTS system shall use less than or equal to 1200 W of electrical power.

  - See Table 3 for measured power usage.

| Component | Power Consumption [Watts] |
| --- | --- |
| ARC Detector Controller (as-shipped with one clock board, one timing board, and one video board) | TBC |
| Gamma Vacuum Ion Pump | TBC |
| ARS Orca Jule Thompson Cryo-Cooler | 630 |
| Lakeshore 336 Temperature Controller | TBC |
| Filter Motor | TBC |
| Sun Server | TBC |

Table 3: NIRTTS Power Consumption.

- Requirement #58: The time to restart the NIR TTS system shall not exceed 10 minutes.

  - The boot time of the system (not including cool-down of the dewar) has been measured to be TBC minutes.

Other tests have been made, as follows:

Hirose PCB / Flex Circuit / PGA:

  - We performed a continuity check from each PGA socket to the controller box patch panel and to the pin on the cable connector at the input to the appropriate Leach board (video, clock).

  - We verified that bias voltages and clocks reached their correct socket on the PGA by enabling them individually in software and seeing the correct action on the correct socket pin.

  - The Hirose PCB was initially found to be shorting to the aluminum clamp when light pressure was applied, so to each side of the PCB a piece of Kapton tape was applied.

  - The original power board would not work with large (ARC-80) power supplies. E.Croner found that pin “N” of the main input connector on the power board was not tied to ground, but needed to be. With added wire to ground pin N, power control board tests OK with ARC-80 power supply.

Suggested Actions: None

## Interface Compliance

## System Compliance

## As-built Documentation

KAON 888 contains detail on the as-built electronics design.  KAON 970 contains detail on the Assembly and Alignment procedures including the electronics connections internal and external to the NIR TTS cryostat.

# Detector Performance [Dave / John]

## Known Detector Anomalies

Detector SN222 exhibits several anomalies which have been attributed to the multiplexer, as follows:

  - spurious vertical line shifts

  - horizontal streaks at channel boundaries

  - erroneous window addressing

During testing the completed detector controller at room temperature, an intermittent problem was found reading in window mode. This problem went away when the instrument was assembled onto the COO Precision Projector, but was never properly corrected. There is a chance it could resurface.

## Functional Compliance

Measurements were made in the lab to determine compliance with the following detector functional requirements calling for experimental verification:

  - Requirement #8: The NIR TTS system shall come equipped with a means of identifying the location of tip-tilt stars on the camera..

    - A full-frame acquisition mode exists (“read mode 1”) which allows for reading the full frame and saving as a FITS file on the host computer.

  - Requirement #9: Goal: The NIR TTS system shall be capable of using up to three tip-tilt stars as input to the real-time controller tip-tilt determination, plus a single readout area for sky background measurements.

    - Up to eight (8) regions of interest (ROIs) can be programmed.

  - Other tests have been made, as follows:

We installed all boards and cables as will be used in the as-shipped system:

    - video board SN264

    - clock driver board SN387

    - timing board SN886

    - power board SN1197 + chassis SN548

    - power supply SN382 with long power cable

    - hirose PCB + flex circuit + PGA socket

    - detector mount + handling fixture.

Loopback Test:

    - Shorted Input: We jumpered the VBIASPOWER bias voltage into all four video+ inputs by installing a jumper directly into the PGA socket. Note that VBIASPOWER is used as the reference input voltage for the video amplifiers.

    - Bias Input: We jumpered another one of the bias voltages (VBIASGATE) into all four video+ inputs, but installing a jumper directly into the PGA socket.

Bare Multiplexer Test:

    - After verifying that the proper clock signals and bias voltages reached each pin on the PGA socket, a bare MUX was installed and operated at room temperature. Initially we had no FRAMECHK signal and no video output (IE, 0V). Found that FRAMECHK was being shorted on the Hirose PCB clamp, which was remedied by covering the PCB with Kapton tape. We also found that the video boards did not have pull-up resistors installed at the inputs (these were designated as “UI” for “user-installed” on the schematic). The pull-up resistors were added which pull up to a fixed bias output, BO7, so we re-assigned VBIASPOWER (3.3V) to BO7.

    - The bare MUX could be read out but showed no change in output with changes in VRESET but rather, output changed in response to DSUB, which should be impossible for a bare MUX. After all other possibilities were ruled out it was determined that the bare MUX itself was at fault. Later examination of the MUX under a microscope revealed a scratch in the indium near the edge had pulled indium material into the bus lines, causing a short and explaining the unusual behavior with DSUB.

Live Detector Test – room temperature

    - Checked for test outputs, FRAMECHK and LINECHK; these signals are generated by the MUX so they are only present if the MUX is being clocked properly. These signals were present.

    - Recorded full-frame test images at room temperature with permanent reset applied.

    - Measured MUX gain (see Error! Reference source not found.)

    - Initial tests showed inconsistent MUX gain. Furthermore, it was found that the VBIASPOWER was being pulled down during readout. VBIASPOWER is nominally 3.3V. During single channel readout it was being pulled down to 3.2V and during 4 channel readout it was being pulled down to 2.9V. We replaced the 1.0k ohm series RC filter resistors at the bias outputs (R345, R350, R356, R358, R362, R366, R371) on video board S/N 264 with 49.9 ohm resistors. This stiffened the bias outputs, correcting the drop in VBIASPOWER. The MUX gain was re-measured (see Error! Reference source not found.).

Live Detector Test – operating temperature

    - measured speed / noise curve (NOT YET COMPLETE – 7/31/13)

    - measured dark current (NOT YET COMPLETE)

    - measured conversion gain (NOT YET COMPLETE)

## Interface Compliance

The detector server interfaces to the detector controller as follows:  The timing board, inside the ARC chassis is connected via two fiber-optic fibers to a ARC PCI interface board, plugged into the PCI bus of a Sun Fire V240 server.  The software interface consists of three-letter (24-bit) commands that are exchanged between the PCI board and the timing board.  Basic commands are defined by ARC.  Custom commands have been defined by the Caltech team.  Pixel data are transmitted from the timing board to the PCI interface which writes the data to DMA memory, where the server access the data for FITS file creation.

The interface between the detector controller and the real-time computer (RTC) is also fiber optic, running from a second channel on the timing board to the RTC.  The software protocol is defined and discussed in KAON 875 – Detector Readout Scheme and Link Protocol Description.

Control of the detector is achieved from other parts of the AO system over Ethernet to the Sun Fire V240 through keywords and Channel Access using the Keck Tasking Library (KTL) software interface.

## System Compliance

## As-built Documentation

KAON 894 describes the extensive inherent detector noise performance evaluation made with the NIR TTS sensor upon receipt in the Caltech Optical Observatories Test Cryostat.  The results presented herein supersede those results, however, as they have been obtained with final electronics, mount, and cabling.

KAON 875 describes the readout modes and timing of the detector controller, the readout scheme for any number of regions of interest (ROIs), and the data protocol for the HOTLink fiber transceiver used for the video link.

NIR TTS control and detector software has been installed in the XXX version control system as WKMO observatory.  Build procedures have been delivered to WMKO and are available at XXX.

Suggested Actions: None

Figure 8: Oscilloscope trace for the live detector SN/220. LINECHK (bottom trace in green) is generated when the last cell of the horizontal shift register is reached. Example shown is for a 4x4 window.

Figure 9: Oscilloscope trace for the live detector SN/220. FRAMECHK (bottom trace in green) requires over-clocking. An extra VCLK at the end of a frame will generate a FRAMECHK pulse. Example shown is for a 4x4 window.

Figure 10: S/N 220 room temperature raw (unsubtracted) image. Permanent reset is applied. Read mode 1 reads through four output channels, which are apparent here.

| VBIASGATE [ volts ] | VRESET=DSUB [ volts ] | OUTPUT 7 [ volts ] | MUX gain (output/VRESET) |
| --- | --- | --- | --- |
| 1.8 | 0.3 | 2.32 |  |
|  | 0.4 | 2.40 |  |
|  | 0.5 | 2.56 |  |
|  | 0.6 | 2.64 | 1.2 |
| 2.0 | 0.3 | 2.16 |  |
|  | 0.4 | 2.24 |  |
|  | 0.5 | 2.32 |  |
|  | 0.6 | 2.40 | 0.8 |
| 2.2 | 0.3 | 1.92 |  |
|  | 0.4 | 2.00 |  |
|  | 0.5 | 2.16 |  |
|  | 0.6 | 2.24 | 1.2 |

Table 4: S/N 222 at room temperature with permanent reset applied, before replacing the series RC filter resistors at the bias output of the video board.

| VBIASGATE [ volts ] | VRESET=DSUB [ volts ] | OUTPUT 7 [ volts ] | MUX gain (output/VRESET) |
| --- | --- | --- | --- |
| 1.8 | 0.3 | 2.56 |  |
|  | 0.4 | 2.64 |  |
|  | 0.5 | 2.72 |  |
|  | 0.6 | 2.80 | 0.8 |
| 2.0 | 0.3 | 2.32 |  |
|  | 0.4 | 2.40 |  |
|  | 0.5 | 2.48 |  |
|  | 0.6 | 2.56 | 0.8 |
| 2.2 | 0.3 | 2.16 |  |
|  | 0.4 | 2.24 |  |
|  | 0.5 | 2.32 |  |
|  | 0.6 | 2.40 | 0.8 |

Table 5: S/N 222 at room temperature with permanent reset applied, after replacing the series RC filter resistors at the bias output of the video board.

# Software Performance [Dave / John]

## Introduction

The heart of the NIR TTS camera control software consists of three instances of the Keck-developed rpcKey_server, with Caltech-supplied libraries for control of the detector, the filter wheel, and the temperature controller.  A fourth instance of the server is run as a global server which acts as a single-point access server for the camera and also communicates with the AO system for real-time updates to the detector server.  All instances run under Solaris 10 on a Sunfire v240.

An emulator was supplied to Microgate not only to facilitate testing and demonstrate functional compliance of their part of the instrument but also to demonstrate functional compliance of our part of the instrument. With the exception of the single-step commands (SSS and SSF) the DSP firmware in the emulator is the same operational code as that used for the final instrument (only detector-specific aspects were stripped in order to comply with ITAR restrictions). Microgate reports success using the emulator.

## Functional Compliance

Measurements were made in the lab to determine compliance with the following software functional requirements calling for experimental verification:

- Requirement #XXX

The detector operates in several different read modes, as follows:

  - Read Mode 0 = full-frame 32-channel raw and CDS images to host. Not currently supported by the hardware, and hence not tested. This is a path for future upgrade.

  - Read Mode 1 = full-frame 4-channel raw and CDS images to host. We setup, exposed, and recorded FITS files on disk; performance was verified by examining the image files visually.

  - Read Mode 2 = Region(s) of Interest / Filmstrip to host. We setup, exposed, and recorded a FITS file on disk for a filmstrip image, which consists of a single ROI read multiple times with each successive read stored adjacent to the previous (thus creating a “filmstrip”). Performance was verified by both visual examination of the filmstrip image and by running a noise analysis Matlab procedure, which calculates RMS noise per frame per unit sampling frequency.

  - Read Mode 3 = Region(s) of Interest to Microgate AO System. We setup and initiated video mode for single and multiple ROIs. Performance was verified by monitoring with a logic analyzer the data bus that feeds the HOTLink fiber transmitter. A simulated data mode was used which enabled writing a specific pattern of data for each of three ROIs to the secondary fiber port, and observing these data in the correct sequence on the logic analyzer.

The Keyword Server was used to operate the detector in all of the available read modes, to verify not only that each read mode produced the correct data but also that the system could be sucessfully switched from one read mode to another.

Performance of the Filter Server:

  - The filter wheel was commanded to each of the allowed positions, blocking, open, H, and Ks, and full-frame images were recorded in those positions to verify that the expected filter was in place.

  - There appears to be a drift with one of the encoders, such that after some period of time of sitting idle, the encoder counts drift from the target counts used to drive the filter wheel into position. Although the filter wheel itself doesn’t move, when the server is queried for the filter position, if the encoder counts don’t match the target counts (within some tolerance) then the position reported could be “unknown”.

Suggested Actions: None

## Interface Compliance

The interface between the NIR TTS camera software and other software components of the AO system is done through the WMKO-developed KTL keyword interface, over Ethernet.  Additionally, real-time updates to certain detector parameters are obtained from the AO Supervisory Control (SC) system via Channel Access using “event functions” defined in the NIR TTS global server.  For each detector-server keyword that needs to be modified in real time (I E during pixel streaming through the second-channel fiber) a corresponding AO keyword is defined in the global server keyword configuration file.  Each of these keywords is associated with an event function, which matches the AO keyword to the NIR TTS detector keyword.  When the global server detects a change in any one of these AO keywords, the corresponding detector keyword is changed also.  In this way, real-time updates from the AO SC can be made in the NIR TTS detector server.

Interfaces from the servers to the temperature and filter controllers are done with TCP/IP over Ethernet using the socket library supplied with the rpcKey_server.

## System Compliance

Measurements were made in the lab to determine compliance with the following software functional requirements calling for experimental verification:

- Requirement #58, Tab 2, #10, Tab 4, #28 Tab 4, KAON 835

Measured time to start all servers:  22  seconds.  Note that any individual server can be stopped and restarted independently.  Individual server starts take about 5 seconds or less.  Hardware is not necessary for a server to start but the server will complain frequently in the log file if the hardware is not available.

Measured time to initialize the detector server:  3 seconds.  If starting in the middle of operations, neither the temperature server nor the filter server should require initialization.

Measured time to home the filter wheel:  42 seconds.  This is the maximum amount of time, for the wheel to make a complete revolution.  Normal times should be substantially less.

Measured time for the Sun Fire V240 to reboot:  1 min.  48 seconds.

Measured time for the Sun Fire V240 to reboot from a complete power outage:  9 minutes 33 seconds.  This lengthy time is because the computer does a number of self-tests before actually booting the operating system.  There may be some system settings that could be changed to eliminate or shorten these tests and thus decrease the time to power up.

Suggested Actions: Investigate possible ways to shorten the power-up time for the V240.

- Requirement #68, Tab 3, KAON 835.

No measurements required.  All code is archived in the WMKO CVS repository.  The camera system can be built by checking out the code and running the appropriate make files.

Documentation will be stored and maintained on the appropriate WMKO instrument page.

- Requirement #69, Tab 3, KAON 835.

No measurement required.  For any of the servers, to move from the current release to a previous version requires only the change of a soft link in the released area of the software environment to point to the desired version directory:

•	Detector server:  $RELDIR/Versions/kss/trick/server/tds/default

•	Filter server:  $RELDIR/Versions/kss/trick/server/tfs/default

•	Temperature server:  $RELDIR/Versions/kss/trick/server/ttcs/default

•	Global server:  $RELDIR/Versions/kss/trick/server/tgs/default

## As-built Documentation

KAON 875 describes all of the raw commands that are accepted by the detector controller. These commands are sent to the controller by the Keyword Server. The Keywords supported by the NIR TTS are described in KAON 857 Keyword Interface Spreadsheet, Tab #4.

Complete, detailed descriptions of the servers can be found in KAON 882, Camera Host Software.

NIR TTS control and detector software has been installed in the XXX version control system at WKMO observatory.  Build procedures have been delivered to WMKO and are available at XXX.

- Requirement OM FR-19: Each opto-mechanical stage shall be provided with an appropriate servo motor and encoder to ensure that the positioning requirements can be met.

Requirements…

  - The dichroic exchanger positioning requirement is  1 mm.  The focus stage positioning requirement is  0.05 mm (confirmed).

Measured Performance…

  - The positional repeatability of both stages was better than the measurement resolution of a dial indicator with 2.5 m accuracy when driven under servo control, dramatically exceeding the requirements.

# Shipping Readiness [Hector]

The NIR TTS Packing and Shipment plan was distributed on XXX to the following persons for review:

- Reviewer #1

- Reviewer #2

- …

We have incorporated feedback from these reviews in version XXX of this document.  Currently, the state of shipping readiness of the equipment described in Section 3 is summarized in Table 6.

Table 6.  Shipping Readiness [XXX insert table above]

# Appendix – Formal Compliance Matrices [Rich]

- [Formal compliance matrices corresponding to requirements documentation will be provided as appendices to this KAON].

Smee, S. A., Barkhouser, R. H., Scharfstein, G. A., Meixner, M., Orndorff, J. D. and Miller, T. 2011, ‘Design of the WIYN High Resolution Infrared Camera (WHIRC)’, PASP 123, 87–106