Keck Adaptive Optics Note XXX

Near-Infrared Tip-Tilt Sensor System

Remove/Install Cryo-Cooler Review Report

H. Rodriguez, R. Dekany, R. Smith, D. Hale, J. Fucik, J. Cromer

August 16, 2013

Draft Version 0.1

Document Revision History

| Revision Number | Revision Date | Summary of Changes | Author |
| --- | --- | --- | --- |
| 0.1 | 08/07/2016 | first draft | H. Rodriguez |

# Overview

The purpose of this procedure is to describe in detail the steps required to safely remove and install the cold head assembly from the Near-Infrared Tip-Tilt Sensor System (NIRTTS).

# Overall Components

Ion Pump

Indexing Pin Access Port

Filter Wheel Motor Assembly

Extension Tube (Shroud)

Cold Head Assembly

Pressure Relief Valve

Cryo-Cooler Assembly Access Port

ARC Electronics

Figure 1: Components that will be affected in this procedure.

Pre-requisite:

- Dewar must be at ambient pressure.

- Replacing or maintenance of any components inside the dewar are required to be done inside a clean room environment and follow clean room protocol.

Tools Required:

- 5/32” Extra-long T-handle hex key (23”) index pin

- 9/64” Extra-long T-handle Hex Key (18”) cold head

- 3mm Extra-long T-handle hex key (12”) coupler #3

- 3/32” Extra-long T-handle hex key (18”) filter wheel pedestal

- 3/32” allen wrench

- 5/32” allen wrench

- CCR open end wrenches (1”, 1-1/8”, 1-3/16”)

- Grounding strap

# Procedures:

Remove Cryo-Cooler Assembly

- Send filter wheel to the home position.

- Disconnect cables.

- Disconnect the detector cable and install shorting plug.

  - Use a grounding strap during the removal of the detector cable.

- Disconnect the filter wheel motor cable.

  - Power off the filter wheel controller before disconnecting the cable.

- Disconnect the thermal control cable.

- Remove electronics.

- The ARC detector controller is installed onto the shroud using (4) 10-32 socket head captive screws.  Remove electronics from shroud.

- Insert index pin through access port.

- Remove the ion pump by releasing the KF16 bulkhead clamp kit.  Use 5/32” allen wrench.  Ion pump is susceptible to contaminates when exposed to the environment.  Place ion pump in an evacuated chamber after removing from the dewar.

Figure 2: Ion pump connected to a KF16 valve and vacuum pump.

- Use index pin hex key to engage the index pin to the filter wheel.

  - The index pin is spring loaded, and will require a 90° turn to engage.

- Remove the shroud to get access to coupler #3.

- Remove (8) 8-32 socket head cap screws that fastens the shroud to the rear dewar cover.  Remove shroud to get access to coupler #3.

- Unclamp the bottom socket head cap screw from coupler #3 using the 3mm hex key.

- Loosen the pedestal from the rear dewar cover.

- Loosen (6) 4-40 x ½” socket head cap screws that secures the pedestal to the rear dewar cover.

- Remove filter wheel motor assembly.

- Rotate ferrofluidic feed through nut.

Figure 3: Position of nut before removing or installing cold head.

- Remove the pressure relief valve.

- Remove the KF16 bulkhead clamp kit for the pressure relief valve to get access to the cold head clamp fastener.

- Unclamp cold head from cryo cooler interface.

- Loosen the 8-32 x 1 ¼” socket head cap screw on the clamping linkage 4 ½ turns.  Use a 9/64” T-handle hex key wrench.

- Remove cold head assembly.

- Remove (12) 6-32 x ½” socket head cap screws on the cold head flange, and carefully pull up the cold head assembly.

- Disconnect K14 sensor connector.

- Disconnect the K14 temperature sensor that is mounted to the cold head tip at the quick disconnect connector.

Quick disconnect connector

Carbon Getter

K14 Temperature Sensor

Figure 4: K14 temperature sensor quick disconnect.  Constantan wire connections are very fragile, and can easily break.

- Remove carbon getter and indium seal.

- Replace indium seal if the carbon getter is removed or if the seal is damaged.

- Cover the cold head port.

- Use 6-32 x ½” socket head cap screws to cover the hole with the cold head blank plate.

- Cold head removal complete.

Install Cryo-Cooler Assembly

- Install carbon getter to cold head if applicable.

- Cut out indium seal to fit between the carbon getter and cold head tip.

- Fasten (5) 4-40 x ½” socket head cap screws to mount the carbon getter to the cold head tip.

- Connect K14 temp sensor connector

- Fasten the K14 temperature sensor to the cold head tip with a 4-40 x ¾” socket head cap screw and disc belleville washer.

- Connect the sensor quick disconnect connector.

Figure 5: K14 temperature sensor mounted on cold head tip.

- Lower cold head assembly into the cryo-cooler interface assembly.

Cold Head Tip/Carbon Getter

8-32 x 1 ¼ shcs

Clamping Linkage

Figure 6: Cryo-Cooler Interface Assembly

- Fasten the (12) 6-32 x ½” socket head cap screws on the cold head to the rear dewar cover.

- Clamp the cold head to the cryo cooler interface.

- After inserting the carbon getter into the cryo-cooler interface, access the cryo-cooler clamp assembly through the pressure relief valve port.

- Clamp the 8-32 x 1 ¼” socket head cap screw.  Using the 9/64” T-handle hex key turn the screw 4 ½ complete turns.

- Install pressure relief valve

- Use the KF16 bulkhead clamp kit to fasten the pressure relief valve to the rear dewar cover.  Fasten the #10 socket head cap screws with a 5/32” allen wrench.

- Use the 1” open end ccr wrench to tighten the ferrofluidic feed through nut.

- Install filter wheel motor assembly.

Figure 7: Filter Wheel Motor Assembly

- Lower coupler #3 mounted on the filter wheel motor assy. onto the ferrofluidic shaft.

- Home filter wheel motor encoder.

- Clamp coupler #3.

- Use the 3mm T-handle hex key to clamp coupler #3 to the ferrofluidic feed through shaft.

- Release index pin.

- Access the filter wheel index pin through the ion pump port.  See Figure 1.

- Use a 5/32” T-handle hex key to release the pin from the filter wheel.

- Test filter wheel for rotation.

- Install ion pump.

- The ion pump is susceptible to contaminants.   If the pump is in an evacuated environment, use nitrogen gas to pressurize pump.

- Fasten the ion pump to the rear dewar cover.

- Use KF16 bulkhead clamp kit to secure the pump to the rear dewar cover.

- Fasten the 10-32 screws with a 5/32 allen wrench.

- Installation of the cold head is complete.