Near-Infrared Tip-Tilt Sensor

Servicing the Filter Wheel

Version 0.1

August 08, 2013

Authors:

Hector Rodriguez

Document Revision History

| Revision Number | Revision Date | Summary of Changes | Author |
| --- | --- | --- | --- |
| 0.1 | 08/08/13 | First Draft | H. Rodriguez |

Contents

Overview:	6

Reference Model:  Add Picture	6

Pre-Requisite:	6

Tools Required:	6

Filter Wheel:	6

Filter Wheel Replacement:	7

Assemble Filter Wheel	9

Figure 1: Filter wheel layout.	7

Figure 3: Front dewar cover with 1” standoffs and window cover.  Current window cover is painted black.	7

Figure 4: Front Radiation Shield	8

Figure 5: Filter Wheel Cover	8

# Overview:

The purpose of this document is to describe the procedure to service the filter wheel for the Near Infrared Tip Tilt Sensor instrument built by Caltech.

# Reference Model:  Add Picture

# Pre-Requisite:

- Pressure and temperature inside the dewar must be at ambient.  Use nitrogen gas to purge dewar.

- All work inside the dewar must be done in a clean room environment and protocol.

# Tools Required:

- Gantry crane

- 3/16” allen wrench

- 3/32” allen wrench

- 5/64” allen wrench

- 9/64” allen wrench

- Air bulb

- Lint free towels

- Nitrogen gas

- Isopropanol alcohol

# Filter Wheel:

- The filter wheel consists of (2) filters, open pupil, and a blocking space.

H Band Filter

Ks Band Filter

Blocking

Open Pupil

Figure 1: Filter wheel layout.

# Filter Wheel Replacement:

- Separate dewar from pedestal by removing (6) 8-32 socket head cap screws.  Not all screws are the same length.

- Remove dewar using the gantry crane.

- All components on the rear side of the dewar should be below the shroud.  Clamp down dewar on perforated bench with the front side facing up.

- Unfasten (12) ¼-20 x 5/8” socket head cap screws.

- Remove cover and rest on the (4) 1” standoffs.

Window Mount

1” Standoffs

Figure 2:Front dewar cover with 1” standoffs and window cover.  Current window cover is painted black.

- Use a lint free cloth to cover the dewar window.

- Unfasten the (16) 4-40  socket head captive screw and remove the front radiation shield.

Figure 3: Front Radiation Shield

- Unfasten the (8) 4-40 socket head captive screws that secure the filter wheel cover to the optical bench.

Camera #1

Figure 4: Filter Wheel Cover

- Cover camera #1 with a lint free cloth.

- To remove the filter, unfasten the (3) 2-56 socket head cap screws and remove the filter wheel retainer.

- Replace or install new filter.

- Store replaced filter.

# Assemble Filter Wheel

- Fasten the replaced filter to the filter wheel with a retainer, (3) 2-56 x ¼” socket head cap screws, and #2 disc belleville washers.  Blow off dust from the filters with an air bulb.

- Use the (2) alignment pins on the filter wheel cover to mount to the optical plate (painted side), and fasten with (8) 4-40 x 3/8” socket head captive screws and #4 split washers.

- Install the front radiation shield to the dewar radiation shield, and fasten with (16) 4-40 x ¼” socket head captive screws and #4 disc belleville washers.

- Inspect and clean O-ring on dewar shell.

- Wipe down the O-ring surface on the cover with isopropanol alcohol and a lint free wipe.

- Align the alignment pin on the cover to the shell.  Note: The window on the cover should align with camera #1.

- Fasten the cover to the shell with (12) ¼-20 x 5/8” socket head cap screws in a star pattern.

- Rotate dewar over on 1” stand-offs.

- Using the gantry crane to lower the dewar onto the pedestal.

- Fasten the (6) 8-32 socket head cap screws to secure the dewar to the pedestal.