Keck Adaptive Optics Note 456

Next Generation Adaptive Optics:

System Requirements Document

Version 1.13

August 22, 2007

1	Introduction	1

2	Scope and Applicability	1

3	References	2

3.1	Related Documents	2

3.2	Referenced Drawings	2

4	Revision History	2

5	Background	4

5.1	Purpose	4

5.2	Motivation for the Development of NGAO	4

5.3	Overview	5

6	Overall Requirements	7

6.1	Science Requirements	7

6.1.1	Purpose and Objectives	7

6.1.2	Science Performance Requirements	7

Table 1. Asteroid companion survey driven requirements	8

Table 2 Asteroid companions orbit determination driven requirements	10

Table 3. Asteroid size and shape derived requirements	12

Table 4. Moons of giant planets derived requirements	14

Table 5. Planets around low mass stars derived requirements	16

Table 6. Measurement of general relativity effects in the Galactic center derived requirements	19

Table 7. Radial velocity measurements derived requirements	20

Table 8 QSO Host galaxy derived requirements	21

Table 9. High-z galaxy derived requirements	22

Table 10. Nearby AGN derived requirements	24

Table 11. Spectroscopic studies of distant galaxies lensed by galaxies	26

Table 12 Imaging studies of distant galaxies lensed by galaxies.	28

6.1.3	Science Instrument Requirements	31

Table 13. Science Instrument Requirements	31

6.1.4	Science Operations Requirements	32

Table 14 Science Operations Requirements, Raw Data Quality	33

Table 15 Science Operations Requirements, Data Products Quality	34

Table 16 Science Operations Requirements, Archiving and Retrieval	34

6.2	Observatory Overall Requirements	35

6.2.1	Purpose and Objectives	35

6.2.2	Facility Requirements	35

Table 17. Facility Requirements	35

6.2.3	Observatory Science Instrument Requirements	37

Table 18. Observatory Science Instrument Requirements	38

6.2.4	Observatory Operational Requirements	38

6.2.5	Observatory Implementation Requirements	41

7	Optical Requirements	42

7.1	Purpose and Objectives	42

7.2	Performance Requirements	42

7.3	Implementation Requirements	45

7.4	Design Requirements	45

8	Mechanical Requirements	52

8.1	Purpose and Objectives	52

8.2	Performance Requirements	52

8.3	Implementation Requirements	52

8.4	Design Requirements	53

9	Electronic/Electrical Requirements	55

9.1	Purpose and Objectives	55

9.2	Performance Requirements	55

9.3	Implementation Requirements	55

9.4	Design Requirements	55

10	Safety Requirements	56

10.1	Purpose and Objectives	56

10.2	Scope	56

10.3	Laser Safety Requirements	56

10.4	Laser Projection Safety Requirements	56

10.4.1	Aircraft Safety	56

10.4.2	Space Command	56

11	Software Requirements	57

11.1	Purpose and Objectives	57

11.2	Scope	57

11.3	Performance Requirements	57

11.4	Implementation Requirements	57

11.5	Design Requirements	57

12	Interface Requirements	58

12.1	Purpose and Objectives	58

12.2	Performance Requirements	58

12.3	Implementation Requirements	58

12.4	Design Requirements	58

12.4.1	Optical Interface	58

12.4.2	Mechanical Interface	58

12.4.3	Electrical/Electronic Interface	59

12.4.4	Software Interface	59

13	Reliability Requirements	61

13.1	Purpose	61

13.2	Scope	61

13.3	Performance	61

14	Spares Requirements	61

15	Service and Maintenance Requirements	61

16	Documentation Requirements	62

16.1	Documentation Package	62

16.2	Drawings	62

16.3	Electrical/Electronic Documentation	62

16.4	Software	62

17	Glossary	63

Figure 1 Keck telescope structure	5

Table 1 NGAO baseline Mauna Kea Cn2 Profile	30

Table 2 Glossary of Terms	63

Requirements Tables

Table 1. Asteroid companion survey driven requirements	8

Table 2 Asteroid companions orbit determination driven requirements	10

Table 3. Asteroid size and shape derived requirements	12

Table 4. Moons of giant planets derived requirements	14

Table 5. Planets around low mass stars derived requirements	16

Table 6. Measurement of general relativity effects in the Galactic center derived requirements	19

Table 7. Radial velocity measurements derived requirements	20

Table 8 QSO Host galaxy derived requirements	21

Table 9. High-z galaxy derived requirements	22

Table 10. Nearby AGN derived requirements	24

Table 11. Spectroscopic studies of distant galaxies lensed by galaxies	26

Table 12 Imaging studies of distant galaxies lensed by galaxies.	28

Table 13. Science Instrument Requirements	31

Table 14 Science Operations Requirements, Raw Data Quality	33

Table 15 Science Operations Requirements, Data Products Quality	34

Table 16 Science Operations Requirements, Archiving and Retrieval	34

Table 17. Facility Requirements	35

Table 18. Observatory Science Instrument Requirements	38

# Introduction

This document describes the requirements for the Next Generation Adaptive Optics (NGAO) system to be built for the W. M. Keck Observatory (WMKO).

The requirements in this document are intended to be at a level appropriate for the system design phase.  Further development of the requirements will take place in the next phase of the project (preliminary design).  In particular, parametric performance requirements given at this stage are intended to indicate the scope and format of the requirements, but do not in all cases establish final values for the specified parameters.  In some cases values for these parameters have yet to be established and are given as TBD.

A more generic set of requirements for new WMKO instrumentation is described in the Observatory’s “Instrumentation Baseline Requirements Document.”  These requirements are also applicable to NGAO.  The NGAO System Requirements Document will take precedence over the “Instrumentation Baseline Requirements Document” in the event of a conflict.

It is important to understand that at this stage of development the requirements provide a basis for identifying the parameters that will be part of the system’s specifications, but the values given are subject to change as the development continues.  During the next phases of the project work will be done to refine the requirements for review at the preliminary design review.  The final requirements to be reviewed at the detailed design review will form the basis for the acceptance test criteria for the instrument.

The purpose of this document is to define and communicate the requirements for the NGAO-specific design and implementation in terms of the needed scientific and technical performance.  The document also expresses specific requirements for implementation or design where those requirements are essential to satisfactory integration and interoperation of NGAO with the observatory systems.  The document avoids prescribing specific design or implementation solutions except for solutions that embody the Observatory’s unique knowledge or experience.  The document establishes requirements for the NGAO that will guide the design through the detailed design phase.

# Scope and Applicability

This document establishes requirements for all aspects of NGAO beyond those already specified in the “Instrumentation Baseline Requirements Document”.  This document also establishes requirements for changes to related Keck telescope subsystems and software where required.

This document does not address the requirements for the science instruments that will work with NGAO, although it does cover the NGAO interfaces to these instruments.  Separate system requirements documents will need to be prepared for each of these instruments as part of their design process.

# References

## Related Documents

- Instrumentation Baseline Requirements Document.

- KAON 153.  Coordination and Use of Laser Beacons for AO on Mauna Kea.

- KAON 399.  NGAO Proposal Executive Summary.

- KAON 400.  NGAO Proposal.

- KAON 428.  Implications and requirements for Interferometry with NGAO.

- KAON 455.  NGAO Science Case Requirements Document v1.0.

- KAON 476.  NGAO Science Operations Observing Model Trade Study

- ANSI Z136.1 Safe Use of Lasers Indoors (2000).

- ANSI Z136.6 Safe Use of Lasers Outdoors (2000).

- 070709_NGAO_RollUp_v1.xls (July 9, 2007) [We need a more official document for this. It is apparently a revised version of the figure in section 2.5 of KAON 455]

## Referenced Drawings

Table X lists the drawing numbers, revisions and date, source and title for all drawings referenced in this document.

# Revision History

| Version | Date | Author | Reason for revision / remarks |
| --- | --- | --- | --- |
| 0.1 | Jan. 16, 2007 | Wizinowich | Initial version |
| 0.3 | Feb. 1, 2007 | Wizinowich | Multiple edits |
| 0.4 | Feb. 6, 2007 | Wizinowich | Multiple edits. Included seeing & telescope environment in section 6.1.2 |
| 1.0 | Feb. 21, 2007 | Wizinowich | Added Dekany performance requirements input |
| 1.1 | Apr. 17, 2007 | Wizinowich | Edited section 6 |
| 1.2 | May 15, 2007 | Wizinowich | Incorporated input from NGAO team meeting 6 |
| 1.3 | May 22, 2007 | Wizinowich | Miscellaneous |
| 1.4 | May 22, 2007 | Wizinowich | Mods to the table in section 6.1.2 |
| 1.5 | May 25, 2007 | Wizinowich | Additions to 6.1.4 based on KAON 476 |
| 1.6 | May 30, 2007 | Wizinowich | Science requirement changes based on SRD telecom |
| 1.7 | May 31, 2007 | Wizinowich | Solar system science requirement input with Le Mignant & Marchis + minor mods to science operation requirements |
| 1.8 | June 1, 2007 | Wizinowich | Galactic Center & QSO host science cases. Mods to Observatory operational requirements |
| 1.9 | June 6, 2007 | Wizinowich | AGN science case from Adkins, Le Mignant, Max & McGrath |
| 1.10 | June 19, 2007 | Wizinowich | Added planets around low mass stars & gravitational lensing requirement tables. First cut at asteroid shape & size table. Added H-band to GR Galactic Center case. Minor edits to asteroid companions table. |
| 1.11 | June 29, 2007 | Wizinowich | Minor edits to science performance requirements |
| 1.12 | August 15, 2007 | Gavel | Review and edits, highlighting gaps in the flow down from the ScRD. Added a requirements numbering system. |
| 1.13 | August 22, 2007 | Gavel | Continued edits, through all the science requirements tables and instrument requirements tables (to Table 13) |

# Background

## Purpose

The purpose of the background section of this document is to provide context and related information for the requirements defined in later sections of this document.

## Motivation for the Development of NGAO

The Keck telescopes are the world’s largest optical and infrared telescopes.  Because of their large apertures the Keck telescopes offer the highest potential sensitivity and angular resolution currently available.  WMKO has already demonstrated scientific leadership in high angular resolution astronomy with the first NGS and LGS AO systems on 8-10 m telescopes.  The importance of achieving the full potential of the Keck telescopes is recognized in the Observatory’s strategic plan which identifies leadership in high angular resolution astronomy as a key long-term goal.

In order to maintain our leadership in this field we must pursue new AO systems and the instrumentation to exploit them.  We have examined, and are continuing to examine, a broad range of key science goals in order to identify the most compelling future science goals of our community and to determine what is needed to realize these goals.  As a result we have identified that NGAO should provide the following suite of capabilities:

- Near diffraction-limited performance at near infrared wavelengths, producing a point spread function with unprecedented precision, stability and contrast;

- Increased sky coverage and a multiplexing capability, enabling a much broader range of science programs; and

- AO correction in the red portion of the visible spectrum (0.6-1.0 µm), delivering the highest angular resolution images available for filled aperture telescopes.

NGAO will be a broad and powerful facility with the potential to achieve major advances in astrophysics.  It will provide dramatic gains in solar system and galactic science where AO has already demonstrated a strong scientific impact.  NGAO will also allow for extraordinary advances in extragalactic astronomy, far beyond the initial gains being made with the Observatory’s current AO systems.

To be clear NGAO need not be a single facility.  It may be that the requirements are best met with multiple AO systems.

The NGAO proposal (KAON 400) and NGAO proposal executive summary (KAON 399) provide more background on the motivation for the development of NGAO.  Further scientific motivation is provided in the NGAO science case requirements document (KAON 455).

## Overview

The scientific and technical requirements for NGAO result in the following basic systems:

- AO system.  The AO system will likely consist of an AO enclosure, an opto-mechanical system, and software and electronics for both non real-time and real-time control.

- Laser facility.  The laser facility will likely consist of a laser enclosure, the laser(s), the launch facility including a beam transport system and launch telescope, safety systems and laser system control electronics and software.

- Science operations facility.  The science operations facility will primarily include the software and computers required to support operation of the AO system and science instruments.  This includes operating the systems for nighttime observing as well as pre- and post-observing activities.

- Science instruments.  The three highest priority instruments are currently a near-IR imager, a visible imager and a deployable near-IR integral field unit (IFU).  Three lower priority instruments have also been identified including a near-IR IFU, a visible IFU and an L and M-band imager.  There is also a requirement that the NGAO project be designed so as to allow the continued AO support of the Interferometer and the fiber injection module used for the ‘OHANA (Optical Hawaiian Array for Nanoradian Astronomy) project.

The AO and laser facilities and the science instruments will have to interface with the telescope structure.  Figure 1 shows a schematic view of a Keck telescope.  The most likely location for the NGAO system and science instruments is on one of the Nasmyth platforms of the telescope.  Nominally we have chosen the left Nasmyth platform of the Keck II telescope as our starting point.  The most likely location for the projection telescope is behind the f/15 secondary mirror in the top end of the telescope.

Figure 1 Keck telescope structure

# Overall Requirements

## Science Requirements

### Purpose and Objectives

The purpose of the science overall requirements section is to summarize and convey requirements that apply generally to the overall NGAO system and its accessories.  These are based on the Science Case Requirements Document (SCRD) (KAON 455) and, in this initial release, also on the NGAO proposal (KAON 399).

### Science Performance Requirements

The performance requirements developed in the SCRD are summarized in the following tables.  These will be updated as the science case requirements and performance budgets become better defined.

The relevant source of the requirement from the Science Requirements Document (SCRD), KAON 455, is referenced by section and paragraph number. Some of the requirements have a source other than the SCRD, a spreadsheet summary file named 070709_NGAO_RollUp_v1.xls (July 9, 2007) which supersedes the SCRD. References to this document are denoted “RollUp_v1 x” in the tables, where x indicates the cell number. Some of the requirements appear to be recommendations from the Observing Model Trade Study, KAON 476 (?).

The requirements for the asteroid companions survey science case are summarized in the following table (see the Multiplicity, Size and Shape of Minor Planets section of KAON 455 (V1) and the Multiple Asteroids 1: Survey mode to find new systems Observing Scenario).

Table 1. Asteroid companion survey driven requirements

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 1.1 | The companion sensitivity shall be ΔH ≥ 5.5 mag at 0.5” separation for a V ≤ 17 asteroid (asteroid size < 0.2”) with a proper motion of ≤ 50 arcsec/hour [ScRD §2.1.4.6, RollUp_v1 F3] | Telescope guiding can be derived from the non-sidereal motion of the asteroid. The asteroid can be used as the tip/tilt guidestar. The AO system has sufficient field of view for both objects and their seeing disks (>3 arcsec (see requirement 8)) The AO system has sufficient throughput to allow tracking a V=17 object at 50 arcsec/hour (14 mas/sec) to an accuracy better than (100 mas? Diffraction limit? Astrometric accuracy, #1.3, will force stricter requirement) The AO system has a low enough emissivity to allow H sensitivity to (17+5.5 mag?) in (how much time? 25 minutes? (see requirement 1.5)) The AO system has sufficient Strehl to achieve this requirement: (need final input on SNR in given exposure time to state wavefront error requirement. ~170nm?) | Near-IR imager |
| 1.2 | H-band relative photometric (between primary and companion) accuracy of 5% at 0.6” for ΔH = 3 for a V ≤ 17 asteroid (asteroid size < 0.2”) with a proper motion of ≤ 50 arcsec/hour [ScRD §2.1.4.4, RollUp_v1 H3] | AO system has sufficient throughput and low enough emissivity to make it physically possible to achieve this photometric accuracy within (how much time? 25 minutes? (see requirement 1.5). System has facilities and tools for calibrating the PSF. | Near- IR imager (no coronagraph) |
| 1.3 | H-band relative astrometric accuracy of ≤ 10 mas for a V ≤ 17 asteroid (asteroid size < 0.2”) with a proper motion of ≤ 50 arcsec/hour [SCRD §2.1.4.5, RollUp_v1 I3] | Uncalibrated detector distortion < 1.5 mas Non-sidereal tracking accuracy of the guider sufficiently small (<10 mas) Thoughput and emissivity to the guider sufficient to achieve SNR ≥ 100 on the primary in <1 second. Requirement #2.4 is stricter. | Near-IR imager |
| 1.4 | Target sample size of ≥ 300 asteroids in ≤ 4 years [SCRD §2.1.3.2 ¶4] Not sure how this number was derived from this ref. |  |  |
| 1.5 | ≥ 25 targets in an 11 hour night [SCRD §2.1.5.2, RollUp_v1 M3] | Observing efficiency: Able to slew to target and begin non-sidereal tracking to specified accuracy within <10 minutes (goal) once every 26 minutes on average. | Observing efficiency |
| 1.6 | Observing wavelengths = J and H-band [SCRD 2.1.6.5, RollUp_v1 B3] | Transmit J and H to science instrument. (only necessary to pass a single band at a time(?)) | J and H-band filters and sensitivity. |
| 1.7 | Spatial sampling ≤ Nyquist at J-band [SCRD §2.1.6.1] |  | ≤ 12 mas pixels in J band imaging mode. Assume this implies Nyquist sampling is also required in H, which would be met by keeping this pixel size for H band imaging mode as well. |
| 1.8 | Field of view ≥ 3” diameter [SCRE §2.1.6.2] | AO system passes a >3” unvignetted field of view | ≥ 250 x 250 pixels at 12 mas pixels |
| 1.9 | The following observing preparation tools are required: guide star finder, field of view, PSF simulation. [This is not one of the science requirements. Are we deducing this requirement from others? KAON 476 §6.2.2?] | Guide star finder tool. PSF simulation tool. |  |
| 1.10 | The following data products are required: Calibrated PSF. Access to archive with proper identification in World Coordinate System. [Not sure where this came from] | Calibrated PSF capability (to what accuracy?) Further study of PSF calibration tools and capabilities is needed (CfAO study). At the very least we need the ability to collect AO telemetry data (at what frame rate?), and the AO sequencer must be capable of repeating an observational parameter setup on a PSF star. | FITS header system is capable of handling non-sidereal offsets in reporting object coordinates in the World Coordinate System |
| 1.11 | Observing requirements: Observer present. Need to know when night will be. [Not sure where this came from] | Classical observing mode or service mode with observer participation |  |

The requirements for the asteroid companions orbit determination science case are summarized in the following table (see the Multiplicity, Size and Shape of Minor Planets section of KAON 455 (V1) and the Multiple Asteroids 2: Orbits determination for the discovered system Observing Scenario).

Table 2 Asteroid companions orbit determination driven requirements

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 2.1 | The companion sensitivity shall be ΔH ≥ 5.5 mag at 0.5” separation for a V ≤ 17 asteroid (asteroid size < 0.2”) with a proper motion of ≤ 50 arcsec/hour [SCRD §2.1.4.6, RollUp_v1 F3] | Same as #1.1 | Near-IR imager |
| 2.2 | The companion sensitivity shall be ΔI ≥ 7.5 mag at 0.75” separation for a V ≤ 17 asteroid (asteroid size < 0.2”) with a proper motion of ≤ 50 arcsec/hour [no such requirement for orbit determination] |  | I-band imager |
| 2.3 | H-band relative photometric (between primary and companion accuracy) of 5% at 0.6” for ΔH = 3 for a V ≤ 17 asteroid (asteroid size < 0.2”) with a proper motion of ≤ 50 arcsec/hour [ScRD §2.1.4.4, RollUp_v1 H3] | Same as #1.2 |  |
| 2.4 | H-band relative astrometric accuracy of ≤ 1.5 mas for a V ≤ 17 asteroid (asteroid size < 0.2”) with a proper motion of ≤ 50 arcsec/hour [SCRD §2.1.4.5 RollUp_v1 I3] | Uncalibrated detector distortion < 1.5 mas Non-sidereal tracking accuracy of the guider sufficiently small (<10 mas) Thoughput and emissivity to the guider sufficient to achieve SNR ≥ 400 on the primary in <1 second. Note this is stricter than requirement #1.3 |  |
| 2.5 | Target sample size of ≥ 100 asteroids in ≤ 4 years [SCRD §2.1.3.2 ¶4] |  |  |
| 2.6 | ≥ 25 targets in an 11 hour night [SCRD §2.1.5.2, RollUp_v1 M3] | Observing efficiency: Able to slew to target and begin non-sidereal tracking to specified accuracy within <10 minutes (goal) once every hour on average. Will generally only observe at one wavelength (wavelength depends on Strehl) (where did this come from? Discussion with science group?) |  |
| 2.7 | Observing wavelengths = J and H-band [SCRD §2.1.6.5, RollUp_v1 B3] | Same as #1.6 (Note: R-band may become a future requirement if R-band Strehl > 15%) | Near-IR imager (R-band imager) |
| 2.8 | Spatial sampling ≤ Nyquist at J-band [SCRD §2.1.6.1, §2.5] |  | ≤ 12 mas pixels in J band imaging mode. (Same as #1.7, see comment there) |
| 2.9 | Field of view ≥ 2” diameter [SCRE §2.1.6.2] | AO system passes a >3” unvignetted field of view | ≥ 250 x 250 pixels at J-band Nyquist sampling |
| 2.10 | The following observing preparation tools are required: guide star finder, field of view, PSF simulation. [Not sure where this came from, KAON 476?] | Same as #1.9 |  |
| 2.11 | The following data products are required: Calibrated PSF. Access to archive with proper identification in World Coordinate System. [Not sure where this came from, KAON 476?] | See #1.10 |  |
| 2.12 | Observing requirements: 8 epochs per target per semester. [SCRD §2.1.3.2 ¶4] |  |  |

The requirements for the asteroid size and shape (characterize surface and orbital parameters) science case are summarized in the following table (see the Multiplicity, Size and Shape of Minor Planets section of KAON 455 (V1)).  In addition to the requirement of a high resolution visible imager, the slope of the visible spectrum is needed to determine the asteroid age or surface type.  This case requires a spectral resolution of R ~ 100 for 0.7 – 1.0 µm wavelength with Nyquist sampling.  Since the slope of the spectra is of importance this could be accomplished with a visible IFU, a visible slit spectrograph or narrow-band filters.

Table 3. Asteroid size and shape derived requirements

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 3.1 | The companion sensitivity shall be ΔI ≥ 7.5 mag at 0.75” separation for a V ≤ 17 asteroid (asteroid size < 0.2”) with a proper motion of ≤ 50 arcsec/hour [RollUp_v1 F2] | Tracking requirements same as #1.1 The AO system has a low enough emissivity to allow I sensitivity to (17+7.5 mag?) in (how much time?) The AO system has sufficient Strehl to achieve this requirement: (need final input on SNR in given exposure time to state wavefront error requirement. ~170nm?) | I-band imager |
| 3.2 | Target sample size of ≥ tbd asteroids in ≤ tbd years [needs to be specified] |  |  |
| 3.3 | ≥ 10 targets in an 11 hour night [SCRD §2.1.5.2, RollUp_v1 M2] | Same as #2.6. 2.6 is the stricter requirement. |  |
| 3.4 | Observing wavelengths = 0.7 – 1.0 µm [SCRD §2.1.6.5, RollUp_v1 B2] | AO system must pass 0.7 to 1.0 microns wavelengths with throughput and emissivity derived for requirement #2.2 (this is stated for I band, not R; what should the throughput be at R?) | Visible imager (R through I band) |
| 3.5 | Spatial sampling ≤ Nyquist at R-band [SCRD §2.1.6.1, §2.5] |  | ≤ 8 mas pixels in R band imaging mode. |
| 3.6 | Field of view ≥ 2” diameter [SCRD §2.1.6.2] | Same as #2.9 | ≥ 375 x 375 pixels at R-band Nyquist sampling |
| 3.7 | Ability to measure the spectral slope with R ~ 100 at 0.85-1.0 m [SCRD §2.1.6.6, §2.5] |  | Spectroscopic imaging at R ~ 100 across R and I-band |
| 3.8 | Ability to measure the SO2 frost bands at R=1000 (R=4000 is acceptable) 1.98 and 2.12 m [SCRD §2.1.6.6, §2.5] |  | Spectroscopic imaging at R ~ 4000 in the K band. Are there other lines of interest possibly in H or J? |
| 3.9 | The following observing preparation tools are required: guide star finder, field of view, PSF simulation. [Not sure where this came from, KAON 476?] | Same as #1.9 |  |
| 3.10 | The following data products are required: Calibrated PSF. Access to archive with proper identification in World Coordinate System. [Not sure where this came from, KAON 476?] | See #1.10 |  |
| 3.11 | Observing requirements: |  |  |

The requirements for the moons of giant planets science case are summarized in the following table (see the TBD section of KAON 455 (V1) and the Io Observing Scenario). Note: moons of giant planets science case is not written up in the ScRD, and is essentially blank in the roll-up tables in section 2.5 in the ScRD. I’ve seen subsequent roll-up charts, but I’m not sure where the “official” version is stored (070709_NGAO_RollUp_v1.xls posted on Meeting 8 page? Still a blank row on moons.)

Table 4. Moons of giant planets derived requirements

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 4.1 | Able to map intensity distribution with a spatial resolution ≤ 20 mas at 0.7 µm for a target with V = 5 to 8 with proper motion ≤ 30” per hour | Telescope guiding can be derived from the non-sidereal motion of the moon. The moon can be used as the tip/tilt guidestar. The AO system has sufficient throughput to allow tracking a V=8 object at 30 arcsec/hour to an accuracy better than 20 mas. The AO system has sufficient Strehl to achieve the resolution requirement: (need input on SNR in given exposure time to state wavefront error requirement. ~170nm?) |  |
| 4.2 | Able to acquire Io within 5” of Jupiter and to track it to within 2.5” |  |  |
| 4.3 | Relative photometric accuracy of ≤ 0.01 mag on target |  |  |
| 4.4 | Absolute photometric accuracy of ≤ 0.05 mag on target |  |  |
| 4.5 | Relative astrometric accuracy of ≤ 2 mas on target |  |  |
| 4.6 | Target sample size of ≥ 12 |  |  |
| 4.7 | Observing wavelengths = 0.7 to 2.4 µm |  | Visible imager. Near-IR imager. May require a number of different filters |
| 4.8 | Spatial sampling ≤ Nyquist at all wavelengths |  | ≤ 7 mas pixels at 0.7 µm |
| 4.9 | Do not allow saturation |  |  |
| 4.10 | Field of view ≥ 2” diameter |  | ≥ 280 pixels at 0.7 µm |
| 4.11 | The following observing preparation tools are required: guide star finder, field of view, PSF simulation, target ephemeris |  |  |
| 4.12 | The following data products are required: Calibrated PSF. |  |  |
| 4.13 | Observing requirements: ≥ 4 epochs per target per semester. Time domain target; requires ≤1 hour notification of activity. |  |  |
| 4.14 | Time domain requirement: Requires ≤1 hour notification of activity. Requires appropriate (TBD) responsiveness. |  |  |

The requirements for the planets around low mass stars science case are summarized in the following table (see the Imaging and Characterization of Extrasolar Planets around Nearby Stars section of KAON 455 (V1)).  The key area in which NGAO will excel is the detection of planets around low-mass stars and brown dwarfs, because Keck, unlike GPI, will be able to use a laser guide star.  NGAO will also be able to search for planets around young solar-type stars where dust extinction is significant.  JWST will have coronagraphic capability in the 3 to 5 m window, but will have significantly lower spatial resolution than Keck NGAO.  In terms of the types of solar systems that can be studied, this means that JWST will focus on older, nearby main sequence stars (since older giant planets will remain visible in 3 to 5 m for a longer time).  JWST may be more limited than NGAO in doing large surveys, because of its longer slewing time and possibly a lifetime limit on the total number of slews.

Table 5. Planets around low mass stars derived requirements

|  | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 5.1 | Target sample 1: Old field brown dwarfs out to 20 pc. Sample size several hundred with a desired maximum survey duration of 3 years. [where did this requirement come from?] | Typical integration times of ~20 min. with 10 min overheads, 10 hrs per night and 10 nights/yr gives a survey size of ~200 targets/ yr. | Near infrared imager (possibly with coronagraph). Survey primary at J- and H-band. Possible dual- or multi-channel mode for rejecting background objects |
| 5.2 | Target sample 2: Young (<100 Myr) field brown dwarfs and low-mass stars out to 80 pc. Sample size several hundred with a desired maximum survey duration of 3 years. [where did this requirement come from?] | Typical integration times of ~20 min. with 10 min overheads, 10 hrs per night and 10 nights/yr gives a survey size of ~200 targets/ yr. | Near infrared imager (possibly with coronagraph). Survey primary at J- and H-band. Possible dual- or multi-channel mode for rejecting background objects |
| 5.3 | Target sample 3: solar type stars in nearby star forming regions and young clusters @ 100 to 150 pc. Bright targets (on-axis self-guiding generally). Sample size several hundred with a desired maximum survey duration of 3 years. [where did this requirement come from?] |  | Possible dual- or multi-channel mode for speckle suppression. Need IR ADC. |
| 5.4 | Companion Sensitivity Sample 1: 15 AU cutoff – 20 to 30 pc; distribution peaks at 4 AU = 0.2"; this gets 2 MJupiter at a separation of 0.2" with a J-band delta magnitude of ΔJ = 10 [where did this requirement come from?] | Excellent (quantify: 5 nm rms?) calibration of nonstatic non-common path aberrations, especially at mid-frequencies. Small servo-lag error (another error contributor. Maybe we should refer to a published non-common path aberration error budget) Speckle suppression techniques (as they would apply to the DM control algorithm – e.g. dark hole algorithm) TBD | 6 to 10 λ/D general-purpose coronagraph Speckle suppression capability (multi-spectral imaging) |
| 5.5 | Companion Sensitivity - Sample 2: 1 MJupiter is 300K, J=22, ΔJ = 11 (2 MJupiter is 2.5 mags brighter). This distribution could have a wider distribution of binaries a) 0.1" separation, ΔJ = 8.5 b) 0.2" separation, ΔJ = 11 c) goal ΔJ = 11 at 0.1" separation [where did this requirement come from?] | Excellent calibration of nonstatic non-common path aberrations, especially at mid-frequencies. Small servo-lag error Speckle suppression techniques | a) 6 λ/D general-purpose coronagraph b) 6 or 10 λ/D general-purpose coronagraph c) Not achievable with a general purpose coronagraph ( May need small IWD coronagraph. Non-redundant aperture masking is an interesting approach for this, limits currently unknown, probably requires low read noise in science detector.) |
| 5.6 | Companion Sensitivity Case 3 placeholder: at 5 Myr , 1 Msun primary; a) goal ΔJ = 13.5 to see 1 MJupiter or b) goal ΔJ = 9 for 5 MJupiter. 0.2" is uninteresting, 0.07" is needed. Should put in apparent magnitudes in a new table (probably J = 22 or 23) [where did this requirement come from?] | Excellent calibration of nonstatic non-common path aberrations, especially at mid-frequencies. Small servo-lag error Speckle suppression techniques | a) Not achievable with a general purpose coronagraph (requires a very small IWA coronagraph, non-redundant masking – maybe have a conversation with Ireland, Lloyd or Tuthill). b) 6 λ/D general-purpose coronagraph |
| 5.6s1 | Sensitivity of H=25 at 5-sigma detection in 20 minutes, at 1 arcsec separation from primary star. What is the magnitude of the primary star? Anything too dim for ExAO (R=12)? [RollUp_v1 G5] | Sufficiently high throughput and low emissivity to permit detecting H=25 in 20 minutes at 5 sigma above background. (need to derive throughput and emissivity numbers) Sufficiently low wavefront error to permit the coronagraph to suppress scattered light from primary to this level. TBD. 140nm meets goal; simulations of 170nm underway. [SCRD §2.2.5.1] | Coronagraph to suppress scattered light from primary star |
| 5.6s2 | Contrast of J>8.5 at 0.1 arcsec separation and J>11 at 0.2 arcsec separation from parent star [RollUp_v1 F5] | Sufficiently low wavefront error to permit the coronagraph to suppress scattered light from primary to this level. TBD. 140nm meets goal; simulations of 170nm underway. [SCRD §2.2.5.1] Excellent (quantify: 5 nm rms?) calibration of nonstatic non-common path aberrations, especially at mid-frequencies. Small servo-lag error (another error contributor. Maybe we should refer to a published non-common path aberration error budget) Speckle suppression techniques (as they would apply to the DM control algorithm – e.g. dark hole algorithm) TBD | 6 to 10 λ/D general-purpose coronagraph Speckle suppression capability (multi-spectral imaging) |
| 5.7 | H-band relative photometry (between primary and companion) Better than 5% photometry relative to primary star [RollUp_v1 G5] <0.05 magnitude relative to primary star [SCRD §2.2.5.4] Inconsistent specs. At what wavelengths does this have to be met? accuracy of ≤ 0.1 mag for recovered companions; goal of measuring colors to 0.05 mags (0.03 mag per band) [where did this requirement come from?] | Diagnostics on AO data to measure Strehl fluctuations if it takes a while to go on and off the coronagraph (a possible more attractive solution is specialized coronagraph features that simultaneously image the primary) | Induced ghost images of primary; or rapid interleaving of saturated and unsaturated images; or a partially transparent coronagraph |
| 5.8 | Astrometric precision 2 mas (~1/10 PSF) relative between primary and planet, for initial rejection of background objects. [SCRD §2.2.5.5, RollUp_v1 I5] Specifications for orbit determination of planet not given in SCRD…: For measuring orbits of nearby field objects, want 0.5 mas to measure masses to 10%. Note this gives you primary mass. To get secondary you need absolute astrometry at very low level (0.1 mas?). This is a goal. It’s unclear if we can meet this goal. Could be combined with Doppler measurements if that’s practical for the brighter objects. … so where did this come from? | Position stability requirement TBD. | Distortion requirement TBD (similar to asteroid science case). Also want ghost images of primary (as for photometry #5.7) in order to locate it accurately relative to planet. |
| 5.9 | Efficiency: 10 minute acquisition, 20 targets per night (30 goal) [SCRD §2.2.6.2] | AO system must be able to absolutely steer objects so they land on the coronagraph, implies 0.005" reproducibility of field steering? –or lock the tip/tilt to this accuracy relative to coronagraph field stop. How do we calculate this accuracy number? |  |
| 5.10 | Observing wavelengths = J-K bands (also interested in Y and z for companion characterization) [SCRD §2.2.4, §2.2.7.1, RollUp_v1 B5] | Transmit J - K to science instrument. | Methane band filters for rapid discrimination, Y or a custom filter for early characterization. |
| 5.11 | Able to register and subtract PSFs (with wavelength, time, etc.) [not sure where this came from or what the accuracy requirement is] |  | 1.5 x Nyquist sampled at J (goal Y) |
| 5.12 | Be able to see companions at 100 AU scales at 30 pc (goal 20 pc) [SCRD §2.2.5.3] |  | Field of view 3" radius (goal 5" radius) |
| 5.13 | Characterization of companion [SCRD §2.2.4, §2.5] |  | R ~150 IFU, sub-Nyquist sampling spectrograph (or possibly Nyquist spatial sampling IFU, R ~ 4,000, OH suppressing). Sensitive to J = 22 or 23. Don’t see a call for R=4000 in the science case |
| 5.14 | Sky Coverage of >30% [SCRD §2.2.7.3, RollUp_v1 N5] | Technical field for low-order wavefront guidestar pickoff at least 150 (180?) arcseconds to achieve this at high galactic latitude. Ability to acquire and track 3 tip/tilt stars. |  |
| 5.14 | The following observing preparation tools are required: guide star finder, field of view, PSF simulation. [Not sure where this came from, KAON 476?] |  |  |
| 5.15 | The following data products are required: Calibrated PSF. Access to archive with proper identification in World Coordinate System. [Not sure where this came from, KAON 476?] |  |  |
| 5.16 | Observing requirements: Observer present. Need to know when night will be. [Not sure where this came from, KAON 476?] |  |  |

The requirements for the Measurement of General Relativity Effects in the Galactic Center science case on both precision astrometry and radial velocities are summarized in the following two tables, respectively (see the Precision Astrometry: Measurements of General Relativity Effects in the Galactic Center section of KAON 455 (V1)).

Table 6. Measurement of general relativity effects in the Galactic center derived requirements

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 6.1 | Astrometric accuracy ≤ 100 µas for objects ≤ 5” from the Galactic Center [SCRD §2.3.8, RollUp_v1 I6] | High Strehl to reduce confusion limit: rms wavefront error ≤ 170 nm. IR tip/tilt sensors. Lower background than current AO system (not quantified how much lower) |  |
| 6.2 | Observing wavelengths K-band [SCRD §2.3.11, RollUp_v1 B6] | Transmit K band to science instrument |  |
| 6.3 | Spatial resolution ≤ Nyquist at H and K Not called out. Perhaps derived from #6.1 |  |  |
| 6.4 | Spatial resolution stable over field of view to 100 mas [SCRD §2.3.9] | Means of aligning tip/tilt sensors so that their mean reference position distorts the field by less than 100 mas in closed loop operation. Means of preventing WFS-blind field-distortion modes (if multi-DMs are in series). Error budgets are needed for these. |  |
| 6.5 | Field of view ≥ 10” x 10” for imaging [SCRD §2.3.11] | Science path shall allow an unvignetted 10” x 10” field. Note: there is no specification on allowable isoplanatic error |  |
| 6.6 | Target drift should be ≤ TBD mas Not sure what this means. See #6.4. |  |  |
| 6.7 | Ability to construct 40x40” mosaic to tie to radio astrometric reference frame [SCRD §2.3.5] To what accuracy? |  |  |
| 6.8 | The following observing preparation tools are required: not specified |  |  |
| 6.9 | The following data products are required: Not specified |  |  |

Table 7. Radial velocity measurements derived requirements

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 7.1 | Radial velocity accuracy ≤ 10 km/sec for objects ≤ 5” from the Galactic Center [SCRD §2.3.10, RollUp_v1 J6] | High Strehl, low background, differential atmospheric refraction correction, PSF estimation [suggestions from SCRD §2.3.10] (need to derive required accuracies. Need to make an error budget for ADC accuracy across band) | Spectral resolution ≥ 4000 |
| 7.2 | Observing wavelengths K-band [SCRD §2.3.6] |  |  |
| 7.3 | Spatial resolution ≤ 20 or 35 mas [SCRD §2.3.6] |  | 20 and 30 mas plate scales, R=4000 IFU spectrometer with… |
| 7.4 | Field of view ≥ 1” x 1” [SCRD §2.3.6] |  | …1” x 1” spatial field of view |
| 7.5 | Target drift should be ≤ TBD mas Not sure what this means. See #6.4. |  |  |
| 7.6 | The following observing preparation tools are required: not specified |  |  |
| 7.7 | The following data products are required: IFU pipeline for wavelength/flux calibration [SCRD §2.3.6] |  |  |

The requirements for the QSO Host Galaxy science case are summarized in the following table (see the QSO Host Galaxy section of KAON 455 (V1)).  The typical QSO that we are considering is at redshift xxx, had a magnitude yyy point source in the center, with a host galaxy of magnitude zzz per square arc sec.  The galaxy is TBD arc sec across.  [Say something about the profile of the galaxy]  The scientific goals are the following: 1) measure colors and magnitudes for the point source; 2) measure morphology and surface brightness profile for the galaxy; 3) obtain spectrum of point source; 4) obtain spatially resolved spectrum of galaxy in order to study its kinematics and stellar populations.  In order to accomplish these things, PSF subtraction will be crucial.  [This is a derived requirement.]

***QSO Host galaxies are not covered in the SCRD, nor are there any entries in the AO Requirements RollUp spreadsheet. These specs should be put on hold until the science requirements are defined.***

Table 8 QSO Host galaxy derived requirements

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 8.1 | Number of targets required: sample size of TBD galaxies in TBD nights or years | Requirement on sky coverage fraction may be implied here, particularly if data from space missions or radio surveys is required |  |
| 8.2 | Required signal to noise ratio for imaging of central point source: TBD at TBD wavelength | AO Strehl ratio > TBD AO background level < TBD; over TBD wavelength range | Instrument background level; Need to be background limited (implications for detector) |
| 8.3 | Photometric accuracy required for imaging the central point source: TBD at TBD wavelengths. | PSF stability and knowledge, temporal and field of view [uniformity trade]; Target drift should be ≤ TBD mas per hour | Calibration stability and accuracy, Zero-point stability and knowledge, Quality of flat-fielding |
| 8.4 | Required SNR for spatially resolved spectroscopy of the point source > TBD at TBD wavelengths | AO Strehl ratio > TBD AO background level < TBD; over TBD wavelength range; PSF stability and knowledge, temporal and field of view [uniformity trade]; Target drift should be ≤ TBD mas per hour | [What spatial sampling is optimum for good PSF subtraction?], spectral resolution R ~ TBD; detector limited SNR performance; FOV = TBD arc sec [e.g. larger for sky subtraction] |
| 8.5 | Required signal to noise ratio for imaging of host galaxy in presence of point source: TBD at TBD wavelength at TBD distance from the point source | AO Strehl ratio > TBD AO background level < TBD; over TBD wavelength range; PSF stability and knowledge, temporal and field of view [uniformity trade]; | Spatial sampling TBD |
| 8.6 | Photometric accuracy required for imaging the host galaxy: TBD at TBD wavelengths at TBD distance from point source. | AO Strehl ratio > TBD AO background level < TBD; over TBD wavelength range; PSF stability and knowledge, temporal and field of view [uniformity trade];Target drift should be ≤ TBD mas per hour | Calibration stability and accuracy, Zero-point stability and knowledge, Quality of flat-fielding; FOV = TBD arc sec [e.g. larger for sky subtraction] |
| 8.7 | Required SNR for spatially resolved spectroscopy of the host galaxy > TBD; spatial resolution = TBD; velocity resolution and accuracy = TBD; wavelength range TBD | AO Strehl ratio > TBD AO background level < TBD; over TBD wavelength range; PSF stability and knowledge, temporal and field of view [uniformity trade]; AO 80% enclosed energy radius = TBD; Target drift should be ≤ TBD mas per hour | FOV = TBD; spectral resolution = TBD; spatial sampling (spaxel size); detector-limited performance; Calibration stability and accuracy, Zero-point stability and knowledge, Quality of flat-fielding; Requirement on persistence? Required minimum trade between single-shot spectral coverage and field of view |
| 8.8 | Required observation planning tools (e.g. guide stars); PSF simulation tools to plan for whether PSF subtraction will be good enough to see the host galaxy |  |  |
| 8.9 | Required data reduction pipeline for IFU |  |  |

The requirements for the high-z galaxy science case are summarized in the following table (see the Galaxy Assembly and Star Formation History section of KAON 455 (V1)).

Table 9. High-z galaxy derived requirements

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 9.1 | Sensitivity. SNR ≥ 10 for a z = 2.6 galaxy in an integration time ≤ 3 hours for a spectral resolution R = 3500 with a spatial resolution of 50 mas [SCRD §2.4.4] This is not explicitly called out like this in the SCRD, but is reasonably consistent with it. The SNR of 10 is not called out. | Sufficiently high throughput and low emissivity of the AO system science path to achieve this sensitivity. Background due to emissivity less than 20% of sky + telescope. [SCRD §2.4.5.1 and SCRD Figure 6] |  |
| 9.2 | Target sample size of ≥ 200 galaxies in ≤ 3 years (assuming a target density of 4 galaxies per square arcmin) [SCRD §2.4.3] Again, not explicitly stated in this manner, but consistent with 10 targets / night over 20 observing nights. | Multi-object AO system: one DM per arm, or an upstream MCAO system correcting the entire field of regard. 6-12 arms on 5 square arc minutes patrol field. Minimum Spacing: TBD | Multiple (6-12) IFUs, deployable on the 5 square arc minute field of regard |
| 9.3 | Observing wavelengths = J, H and K (to 2.4 µm) What portions of these bands must be detected in a given single observation? [SCRD §2.4.5.2, RollUp_v1 B11] | AO system must transmit J, H, and K bands | Infrared IFUs designed for J, H, and K. |
| 9.4 | Spectral resolution = 3000 to 4000 [SCRD §2.4.5.2] |  | Spectral resolution of >3000 in IFUs |
| 9.5 | Spatial resolution 50 mas [SCRD §2.4.5.1] RollUp_v1 refers to 70mas when stating encircled energy spec. Is this consistent? | See encircled energy spec. #9.5 | 50 mas IFU spaxal scale |
| 9.6 | Encircled energy 50% in 70 mas [RollUp_v1 E11, N11] Note SCRD §2.4.5.1 seems garbled. (50% instead of 50mas (typo?)? What is meant by optimal tip/tilt stars?) Is this consistent with #9.5. Is there an EE spec for 50 mas? | Wavefront error sufficiently low (~170 nm) to achieve the stated requirement in J, H, and K bands. |  |
| 9.7 | Velocity determined to ≤ TBD km/sec for spatial resolutions of 50 to 70 mas (not specified in SCRD) | PSF intensity distribution known to ≤ 10% per spectral channel. Spatial and spectral model fitting valid to ≤ TBD (how did this get derived?) |  |
| 9.8 | Field of view ≥ 1” x 3” [SCRD §2.4.5.1] | Each MOAO channel passes a 1”x3” field plus a seeing disk diameter (~1”), assuming uncorrected field is at multiobject pickoff plane. Can be tightened appropriately if plane is partially or fully AO corrected. | Each IFU unit’s field of view is 1” x 3” |
| 9.8opt | Field of view ≥ 1” x 1” [SCRD §2.4.5.1] | Each MOAO channel passes a 1”x1” field plus a seeing disk diameter (~1”), assuming uncorrected field is at multiobject pickoff plane. Can be tightened appropriately if plane is partially or fully AO corrected. An additional MOAO channel is dedicated to sky observations. | Each IFU unit’s field of view is 1” x 1” One extra IFU for dedicated sky observations |
| 9.9 | Simultaneous sky background measurements within a radius of 3” with the same field of view as the science field | See choice of #9.8 or #9.8opt above. |  |
| 9.10 | Relative photometry to ≤ 5% for observations during a single night [RollUp_v1 H11] (left blank in SCRD) | Knowledge of ensquared energy in IFU spaxel to 5%. Telemetry system that monitors tip/tilt star Strehl and other real-time data to estimate the EE vs time. |  |
| 9.11 | Absolute photometry ≤ 0.3 mag (can’t find this requirement) |  |  |
| 9.12 | Sky coverage ≥30% to overlap with data sets from other instruments and telescopes [SCRD §2.4.5.3, RollUp_v1 N11] | Infrared tip/tilt sensors with AO correction of tip/tilt stars (specify order of correction?) |  |
| 9.13 | Should be able to center a galaxy to ≤ TBD of science field of view |  |  |
| 9.14 | Should know the relative position of the galaxy to ≤ TBD of spaxel size |  |  |
| 9.15 | Target drift should be ≤ TBD of spaxel size |  |  |
| 9.16 | The following observing preparation tools are required: TBD |  |  |
| 9.17 | The following data products are required: calibrated spectral data cube in World Coordinate System with absolute TBD accuracy |  |  |

The requirements for the Nearby AGN science case are summarized in the following table (see the Nearby AGN section of KAON 455 (V1)) (there is no write-up on AGNs in the present version of the SCRD).  The typical AGN that we are considering is at redshift <0.05, and if a Seyfert 1 galaxy has a magnitude yyy point source in the center, with a host galaxy of magnitude zzz per square arcsec.

The region of interest for spatially resolved spectroscopy is within the gravitational sphere of influence of the central black hole: generally we will need at least two resolution elements across this distance.  [convert to arcsecs as function of z]

The scientific goals are the following: to measure the black hole mass using stellar kinematics in the cores of AGNs.  In order to accomplish this, PSF subtraction will be crucial for Seyfert 1 galaxies.  [This is a derived requirement.]

Table 10. Nearby AGN derived requirements

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| - 10.1 | Number of targets required: sample size of TBD galaxies in TBD nights or years | Requirement on sky coverage fraction may be implied here |  |
| - 10.2 | Required wavelength range 0.85 – 2.4 microns |  |  |
| - 10.3 | Required spatial sampling at least two resolution elements across gravitational sphere of influence | 80% enclosed energy radius < ½ gravitational sphere of influence. This implies a total wavefront error of not more than TBD nm over a range from TBD to TBD microns. | Spectral and imaging pixels/spaxels < ½ gravitational sphere of influence (in the spatial dimension) |
| - 10.4 | Required field of view for both spectroscopy and imaging > 10 radii of the gravitational sphere of influence [fill this in] |  | Will need to get sky background measurement as efficiently as possible. For IR, consider using d-IFU on the sky; for visible, need solution |
| - 10.5 | Required SNR for spatially resolved spectroscopy of the central black hole region using stellar velocities > 30 per resolution element | AO Strehl ratio > TBD at 0.85 microns (Ca infrared triplet). This implies a total wavefront error of TBD nm at 0.85 microns. PSF stability and knowledge, temporal and field of view [uniformity trade]; | Spectral resolution R ~ 3000-4000 (TBD) with two pixels per resolution element; detector limited SNR performance; Spatial sampling at least two resolution elements across the gravitational sphere of influence |
| - 10.6 | Required signal to noise ratio for imaging of the region around the central black hole [is this a contrast requirement?] | AO Strehl ratio > TBD, this implies a total wavefront error of not more than TBD nm over a range from TBD to TBD microns. PSF stability and knowledge, temporal and field of view [uniformity trade]; | Spatial sampling at least two resolution elements across the gravitational sphere of influence |
| - 10.7 | Photometric accuracy required for imaging the central point source and possible cusp: TBD at TBD wavelengths | AO Strehl ratio > TBD, this implies a total wavefront error of not more than TBD nm over a range from TBD to TBD microns. AO background level < TBD; over TBD wavelength range; PSF stability and knowledge, temporal and field of view [uniformity trade]; | Calibration stability and accuracy, Zero-point stability and knowledge, Quality of flat-fielding; |
| - 10.8 | Velocity determined to ≤ TBD km/sec for spatial resolutions of TBD mas | PSF intensity distribution known to ≤ xxx% per spectral channel. Spatial and spectral model fitting valid to ≤ TBD |  |
| - 10.9 | Required observation planning tools (e.g. guide stars); PSF simulation tools to plan for observations of Seyfert 1 galaxies which have strong central point sources |  |  |
| - 10.10 | Required data reduction pipeline for IFU |  |  |

The requirements for the gravitational lensing science case are summarized in the following two tables.  (there is no write-up on gravitational lensing in the present version of the SCRD).

Table 11. Spectroscopic studies of distant galaxies lensed by galaxies

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 11.1 | SNR ≥ 10 for a z = 1 – 2 galaxy in an integration time ≤ 3 hours for a Gaussian width 20 km/sec Gaussian width (50 km/sec FWHM) with a spatial resolution of 50 mas |  | R ~ 5000 (or whatever is needed to achieve 20 km/sec sigma on these targets) |
| 11.2 | Target sample size of ≥ 50 galaxies, with density on the sky of 10 per square degree. Survey time ~ 3 years. | Number of IFUs: at least one, plus preferably one to monitor the PSF and one to monitor the sky. The extra two IFUs could be dispensed with if there were other ways to monitor the PSF and the sky background. |  |
| 11.3 | Observing wavelengths = J, H and K (to 2.4 µm). Emphasis is on shorter wavelengths. Would use z and I bands if available. |  |  |
| 11.4 | Spectral resolution: whatever is needed to get 20 km/sec radial velocity Gaussian sigma |  | Spectral resolution |
| 11.5 | Spatial resolution 50 mas | [need to make the spatial resolution and the enclosed energy requirements consistent with each other] |  |
| 11.6 | Velocity determined to ≤ 20 km/sec Gaussian sigma for spatial resolutions of 50 mas | PSF intensity distribution known to ≤ TBD% per spectral channel. |  |
| 11.7 | Field of view > 4” diameter (lenses are bigger than galaxies). Typical lens is 3 to 4 arc sec diameter. |  |  |
| 11.8 | Simultaneous sky background measurements |  |  |
| 11.9 | Relative photometry to ≤ 0.1 mag for observations during a single night |  |  |
| 11.10 | Absolute photometry ≤ 0.3 mag |  |  |
| 11.11 | Sky coverage at least 50% with enclosed energy radius within TBD mas. Sky coverage should be better than with current LGS. |  |  |
| 11.12 | Should be able to center a galaxy to ≤ TBD of science field of view |  |  |
| 11.13 | Should know the relative position of the galaxy to ≤ TBD of spaxel size. (Whatever works for high z galaxy case is OK here as well) |  |  |
| 11.14 | Target drift should be ≤ TBD of spaxel size (Whatever works for high z galaxy case is OK here as well) |  |  |
| 11.15 | The following observing preparation tools are required: |  |  |
| 11.16 | The following data products are required: calibrated spectral data cube |  |  |

Table 12 Imaging studies of distant galaxies lensed by galaxies.

Goal: screen potential lensed-galaxy targets for more detailed and lengthy spectroscopic study. (there is no write-up on distant lensed galaxies in the present version of the SCRD).

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 12.1 | SNR ≥ 3 per pixel (100 per source) for a z = 1 – 2 galaxy in an integration time ≤ 1/2 hour. Strehl > 0.3 at J band. |  |  |
| 12.2 | Target sample size of ≥ 200 galaxies, with density on the sky of 10 per square degree. Survey time ~ 3 years. | Overhead less than 10 min between targets. | 10 per square degree implies that you will only be able to observe one target at a time – average of 1 in every ~19x19’ patch. |
| 12.3 | Observing wavelengths = J, H and K (to 2.4 µm). Emphasis is on shorter wavelengths. Could use z if available. Thermal part of K band less important. |  |  |
| 12.4 | Spatial resolution better than 50 mas | Need a good model of the PSF or a simultaneous image of a PSF star. Need a figure of merit for goodness of the PSF: how well the model fits the “real” PSF in two dimensions. | Need Nyquist sampling of pixels at each wavelength. |
| 12.5 | Field of view > 15” diameter for survey. Bigger is better. Some degradation between center and edge of field is tolerable. (Need to quantify.) |  |  |
| 12.6 | Relative photometry to ≤ 0.1 mag for observations during a single night |  |  |
| 12.7 | Absolute photometry ≤ 0.3 mag |  |  |
| 12.8 | Sky coverage at least 50% with enclosed energy radius within 0.1 arc sec at H or K. Sky coverage should be better than with current LGS. |  |  |
| 12.9 | Should know the relative position of the galaxy to ≤ TBD of spaxel size. (Whatever works for high z galaxy case) |  |  |
| 12.10 | The following observing preparation tools are required: |  |  |
| 12.11 | The following data products are required: accurate distortion map (to 1% of the size of the galaxy, or 0.01 arc sec rms) |  |  |

Notes:

Want to repeat this exercise with galaxies lensed by clusters.

Wide field: Typical size of the highly magnified region of a cluster is 1 arc min.  Need low background: lens arcs from z ~ 7 are at most Vega magnitude 23 or 24 in H (brightest arcs).  Typical size small (half light radii 0.1 arc sec).

Closer galaxies with giant arcs: deployable IFU application.  Several arc sec long.  Want field of regard of about an arc min.  Usually 3 to 5 multiple arcs within a square arc min.  (But each might be long, and require more than 1 IFU unit.)  Less than 10 IFU units needed in a square arc min.

Is a requirement needed on achieving a particular astrometric accuracy in a particular time needed for survey-type science (i.e., 1 mas in 10 min).

Aren’t encircled energy requirements needed for the IFUs?

The atmosphere and telescope parameters assumed for achieving these numbers are summarized below.

Atmospheric Seeing Assumptions: (perhaps this should be an appendix?)

The NGAO system shall provide its nominal performance when the atmospheric seeing is characterized by the following conditions.  An evaluation of existing seeing data has been performed (KAON 303). The KAON 303 profile was modified to include a stronger ground layer and the standard r0 value was lowered from 20 to 18 cm. The resultant baseline median Cn2 profile is presented in Table 1.  From this model we calculate the following turbulence parameters for 0.5 µm wavelength (note that r0, θ0 and 1/fG increase as λ6/5):

- Fried’s seeing parameter 	r0 = 18 cm

- Isoplanatic angle 	θ0 = 2.5 arcsec

- Turbulence characteristic frequency 	fG = 39 Hz

In addition, we have adopted a standard deviation for r0 of r0  = 3 cm with a characteristic evolution time of t = 3 min.

Table 1 NGAO baseline Mauna Kea Cn2 Profile

| Altitude (km) | Fractional C | Wind Speed (m/s) |
| --- | --- | --- |
| 0.0 | 0.471 | 6.7 |
| 2.1 | 0.184 | 13.9 |
| 4.1 | 0.107 | 20.8 |
| 6.5 | 0.085 | 29.0 |
| 9.0 | 0.038 | 29.0 |
| 12.0 | 0.093 | 29.0 |
| 14.8 | 0.023 | 29.0 |

Telescope and Dome Environment: (perhaps this should be in an appendix?)

The NGAO system shall provide its nominal performance when the telescope and dome environment can be characterized by the following conditions.

Dome & telescope seeing: The Keck dome and telescope environment degrades the intrinsic seeing by less than 0.1 arcsec, in quadrature, as measured from the effective increase in image FWHM (this change corresponds to decreasing the r0 parameter from 18 cm to 17.8 cm).

Phasing: The phasing errors will be 10 nm rms wavefront or less before NGAO correction.  Standard performance is 60 nm rms currently.  Currently available algorithms have demonstrated 10 nm rms.  This error interacts with the segment figure error discussed next.  We may want to place an error on the overall telescope wavefront figure PSD instead.

Segment figure:  The wavefront error of the 36 segments will be less than 80 nm rms wavefront after warping, but before NGAO correction; this number is an average over all 36 segments segment.  As a goal the wavefront error shall be 80 nm rms over each segment.

Stacking: The segment stacking errors will contribute less than 20 nm rms wavefront to the overall wavefront error before NGAO correction.

Line of sight jitter: The aggregate line of sight jitter (wavefront tip and tilt) resulting from motion of the primary, secondary and tertiary mirrors will be less than 0.020 arc seconds rms before correction by the NGAO.  This vibration is known to currently be largely dominated by a narrow resonance at ~29 Hz.

Segment motion: The motion of each segment as a solid body will be less than 0.015 arc seconds rms before correction by the NGAO. This vibration is known to currently be largely in a narrow resonance at ~29 Hz.

### Science Instrument Requirements

The NGAO system must be capable of supporting the following science instruments (in rough order of priority), based on the NGAO proposal and SCRD:

- Visible imager.  Wavelength range = 0.6 to 1.1 µm.  Field of view = 20”x20”.  Image sampling = 6 mas pixels.

- Near-IR imager.  Wavelength range = 1.0 to 2.45 µm.  Field of view = 20”x20”.  Image sampling = 10 mas.

- Deployable near-IR Integral Field Unit (IFU).  Wavelength range = 1.0 to 2.45 µm.  Field of Regard = 1.5’x1.5’ with 3x3.4” fields of view for each IFU.  Image sampling = 100 mas.

- Near-IR IFU.  Wavelength range = 1.0 to 2.45 µm.  Field of view from 2”x1.25” to 16”x5”.  Image sampling = 20 to 100 mas.

- Visible IFU.  Wavelength range = 0.6 to 1.1 µm.  Field of view from 1.2”x1.36” to 12”x6.8”.  Image sampling = 20 to 100 mas.

- L and M-band imager.  Wavelength range = 3.0 to 5.3 µm.  Field of view = 25”x25”.

- Image sampling = 25 mas.

- Future science instruments from the above list or completely new instruments.  These future science instruments would need to be designed so as to fit at a movable port or to replace a fixed first generation instrument.

Table 13. Science Instrument Requirements

| # | Science Instrument Requirement | Discussion | Based on Red means unspecified or vague |
| --- | --- | --- | --- |
| 13.1 | Visible Imager: the field of view shall be ≥ 2 arcsec (diameter?) |  | SCRD §2.1.6.2, §2.5: asteroid shapes and companions |
| 13.2 | NIR Imager: the field of view shall be ≥ 2 arcsec |  | SCRD §2.1.6.2, §2.5: Io, Titan, debris disks and QSO host galaxies |
| 13.3 | NIR Imager: the field the field of view shall be ≥ 3 arcsec |  | SCRD §2.2.5.3: planets around low mass stars |
| 13.4 | NIR Imager: the field the field of view shall be ≥ 10 x 10 arcsec |  | SCRD §2.3.11: Galactic Center |
| 13.5 | NIR Imager: shall provide a coronagraph |  | SCRD §2.2.4, RollUp_v1 G5 : planets around low mass stars, debris disks and QSO host galaxies |
| 13.6 | NIR Imager: wavelength coverage shall be at least 0.9 to 2.4 µm |  | SCRD §2.2.4, RollUp_v1 B5: planets around low mass stars |
| 13.7 | NIR IFU: field of view shall be ≥ 2 arcsec | Only 1” is required for the Galactic Center | SCRD §2.1.6.2: Asteroids, Titan |
| 13.8 | Visible IFU: field of view shall be ≥ few (specify) arcsec |  | SCRD §2.5: Nearby AGNs |
| 13.9 | Visible IFU: spectral resolution shall be R ~ 4000 | What range is acceptable? Note: Asteroid size and shape prefers R=1000 but tolerates 4000 [SCRD §2.1.6.6] | SCRD §2.3.6: Galactic center radial velocity, Nearby AGNs |
| 13.10 | NIR deployable IFU: field of view shall be ≥ few arcsec | 1”x3” or larger required for high z galaxies. (see #9.8opt for an option to this) Galactic Center needs a dIFU.don’t see this requirement for GC | SCRD §2.4.5.1: High z galaxies, gravitational lensing |
| 13.11 | NIR deployable Imager: field of view shall be ≥ few (specify) arcsec |  | SCRD §2.4.5.1:: high z galaxies (no discussion of deployable imager here, although I know we talked about it) |
| 13.12 | NIR dImager: field of view shall be ≥ 10 arcsec |  | KAON 455 (v1) Table 5: resolved stellar populations |

It is TBD whether the NGAO system will need to support any of the existing science instruments (NIRC2, NIRSPEC or OSIRIS).  The Interferometer and ‘OHANA requirements are discussed in section 6.2.3.

It is a goal for NGAO to support visitor science instruments.

### Science Operations Requirements

The top-level science operations goals for the NGAO system including the science instruments are the following (see KAON 476):

- Science-grade quality of the raw data (i.e., image quality and completeness of observations).

- Science-grade quality of the data products (i.e., photometry, astrometry, etc.).

- Science impact from a given data product (i.e., number of publications and citations).

The requirements that support these top-level goals are defined in the following tables.

Science-grade quality of the raw data (see KAON 476, section 6.2.1):

Table 14 Science Operations Requirements, Raw Data Quality

| # | Science Operations Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 14.1 | Provide an extensive set of tools for instrument performance simulation and observing preparations | These tools should be designed with the end user in mind. | KAON 476, section 6.2.1 |
| 14.2 | Document the instrument performance at an appropriate level to support observing preparations | This is both a development and operations requirement, since a continued effort will be required to characterize, monitor and document the performance | KAON 476, section 6.2.1 |
| 14.3 | Provide a semi-real-time level 1 data reduction pipeline for each instrument to at minimum perform background subtraction, cosmetic correction and shift-and-add of images. |  | KAON 476, section 6.2.1 |
| 14.5 | Provide semi-real-time tools to perform an assessment of the image quality on the level 1 data including SNR, Strehl and encircled energy. |  | KAON 476, section 6.2.1 |
| 14.6 | Provide a science operations paradigm that optimizes the completion rate for a significant fraction (TBD) of observing programs. |  | KAON 476, section 6.2.1 |

Science-grade quality of the data products (see KAON 476, section 6.2.2):

Table 15 Science Operations Requirements, Data Products Quality

| # | Science Operations Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 15.1 | Provide the required calibration methods and tools to achieve the astrometry performance requirements |  | KAON 476, section 6.2.2 |
| 15.2 | Provide the required calibration methods and tools to achieve the photometry performance requirements |  | KAON 476, section 6.2.2 |
| 15.3 | Provide the required calibration methods and tools to achieve the PSF characterization requirements |  | KAON 476, section 6.2.2 |

Science impact from a given data product (see KAON 476, section 6.2.2):

Table 16 Science Operations Requirements, Archiving and Retrieval

| # | Science Operations Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 16.1 | Develop a plan for data archival |  | KAON 476, section 6.2.3 |
| 16.2 | Develop a plan for data retrieval from the data archive |  | KAON 476, section 6.2.3 |
| 16.3 | Document the on-sky science performance of each science instrument with NGAO | This should be based on an extensive science verification phase. | KAON 476, section 6.2.3 |

## Observatory Overall Requirements

### Purpose and Objectives

The purpose of the overall requirements section is to convey requirements that apply generally to the overall instrument and its accessories based on the Observatory’s requirements.

Note that the Observatory’s standard requirements for all new instrumentation are summarized in the Instrumentation Baseline Requirements Document.  Additional Observatory requirements specific to NGAO are listed in the following sections.

### Facility Requirements

The following are requirements imposed by the nature of the existing facility.

Table 17. Facility Requirements

| # | Facility Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 17.1 | The NGAO system & science instruments should be located on the Nasmyth platform of one of the Keck telescopes | The Keck telescope foci and Nasmyth deck storage locations are already heavily utilized. The current AO systems occupy the left Nasmyth platform locations of both telescopes. HIRES occupies the right Nasmyth on Keck I while DEIMOS and NIRSPEC share the right Nasmyth on Keck II. The Cassegrain foci are occupied by LRIS (and MOSFIRE in the future) on Keck I and by ESI on Keck II. The bent Cassegrain ports are believed to have inadequate space and weight capacities. The prime focus could potentially be available but there would be many constraints on an instrument at this location. The most viable option is in the location of an existing AO system. Alternatives would be to decommission HIRES or for the existing AO system and the NGAO system to be able to share the same platform. |  |
| 17.2 | The NGAO system should accommodate the entire Keck pupil | The Keck primary has a maximum edge-to-edge diameter of 10.949 m. |  |
| 17.3 | If the existing f/15 or f/40 secondary mirrors are used then the NGAO system will be constrained by the resultant f/#, focal plane and pupil location | Both telescopes have f/15 secondary mirrors, as well as chopping secondary units that can accommodate f/25 and f/40 secondary mirrors. The choice of f/15 secondary mirrors for the current AO systems was largely driven by the resultant reduced size of the AO systems and the availability of PCS (Phasing Camera System) via a rotation of the tertiary mirror. The inability of the current systems to chop at the telescope pupil has been a limitation for thermal IR observations. The focal length of the telescope with the f/15 secondary mirror is 150 m. The 10.949 m primary corresponds to an f/13.66 beam with an exit pupil diameter of 1.460 m located 19.948 m in front of the focal plane. |  |
| 17.4 | The NGAO facility must not compromise the performance of a non-NGAO instrument when that instrument is being used for science or engineering | This requirement is intended to ensure that the NGAO facility, when not in use, does not introduce vibrations or stray light that might impact the performance of another science instrument or an engineering instrument such as PCS. |  |
| 17.5 | The NGAO facility must not compromise the performance of the telescope when the telescope is used for non-NGAO observations. | The NGAO system should not impact the dynamic performance of the telescope through vibrations or different telescope dynamics. If NGAO hardware is to be mounted in the top-end then it must be designed/implemented not to compromise the secondary mirror performance. |  |
| 17.6 | The NGAO facility / science instrument combination should provide compensation for science field rotation | The Alt-Az telescope design requires compensation for field rotation in order to keep the science field fixed on the science instrument. |  |
| 17.7 | The NGAO facility / science instrument combination should provide compensation for pupil rotation | The Alt-Az telescope design and the irregular shape of the Keck primary mirror require that NGAO system provide appropriate compensation for pupil rotation. Examples: The existing Keck AO system updates the reconstructor to compensate for pupil orientation. The NIRC2 coronagraph mask rotates to match the rotating pupil. |  |
| 17.8 | The NGAO facility should provide compensation for LGS projector rotation | The Alt-Az telescope design will cause the laser projector (and the resultant LGS asterism) to rotate with respect to the Nasmyth platform. Compensation will be required to maintain the off-axis LGS on the corresponding wavefront sensor. |  |
| 17.9 | The NGAO facility should accommodate access for routine maintenance of the telescope | For example, access to the elevation journal, elevation wrap, bent Cassegrain platform and stairwell to the mirror cell |  |

### Observatory Science Instrument Requirements

In addition to the science instrument requirements specified in section 6.1.3 the NGAO facility must allow the Observatory to continue supporting Interferometer science with the two Keck telescopes.  The requirements for these instruments are developed in KAON 428 and are further specified in later sections.

Table 18. Observatory Science Instrument Requirements

| # | Observatory Instrument Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 18.1 | The NGAO facility should support the Keck Interferometer (KI) with performance as good or better than provided by the pre-NGAO Keck AO systems | The KI dual star modules (DSM) currently move into both AO enclosures on rails to feed the KI. The requirement to feed the KI requires that collimated and f/15 light can be fed to the DSM and that the field rotation, pupil rotation, longitudinal dispersion and polarization from the NGAO system and the AO system on the other telescope be identical. See KAON 428 Implications and Requirements for Interferometry with NGAO. |  |
| 18.2 | The NGAO system should support the ‘OHANA interferometer with performance as good or better than provided by the pre-NGAO Keck AO systems | Injection modules are currently placed on each AO bench to feed an optical fiber that goes to the KI. In future the output from these fibers will be interfered with those from multiple telescopes. |  |

### Observatory Operational Requirements

The purpose of this section is to document the Observatory’s overall requirements for the support of science operations.  These requirements can be divided into the following categories:

- Percent of time collecting science quality data.

- Capability to support a certain number of nights per year of science operations.

- Operational costs.

- Impact on daytime and nighttime operations.

- Compliance with regulations, including safety, Mauna Kea policies, FAA and U.S. Space Command.

Each of these categories is represented in one of the following tables.  In addition, all of the above categories require a facility-class NGAO system and science instruments, and this requirement is reflected in the final table.

| # | Observatory Operational Requirement | AO Derived Requirements | Instrument Derived Requirements |
| --- | --- | --- | --- |
| 1 | More than 80% of the allocated observing time is spent on collecting science-quality data | Observing model. Observing overhead and efficiency. Observing reliability. Observing tools. |  |

| # | Observatory Operational Requirement | AO Derived Requirements | Instrument Derived Requirements |
| --- | --- | --- | --- |
| 1 | The NGAO system must be capable of supporting 240 nights/year of science operation | Observatory currently willing to support 140 nights/year. |  |

| # | Observatory Operational Requirement | AO Derived Requirements | Instrument Derived Requirements |
| --- | --- | --- | --- |
| 1 | The NGAO facility should be designed to keep the total annual operational personnel and non-personnel costs below the following budgets: - Four FTEs and $200k (FY07 dollars) in non-personnel costs for the first 100 nights/year of science operations An additional cost of 0.5 FTEs and $25k (FY07 dollars) in non-personnel costs for each additional 50 nights of science operations. |  |  |

| # | Observatory Operational Requirement | AO Derived Requirements | Instrument Derived Requirements |
| --- | --- | --- | --- |
| 1 | The NGAO facility should not require more than two engineering nights per year for system maintenance |  |  |
| 2 | The NGAO facility should not routinely require more than 30 minutes of daytime telescope restriction on an NGAO science night |  |  |

| # | Observatory Operational Requirement | AO Derived Requirements | Instrument Derived Requirements |
| --- | --- | --- | --- |
| 1 | The Mauna Kea laser projection requirements must be satisfied | This includes requirements on laser power, wavelength, laser traffic control participation, aircraft safety and space command. See KAON 153. The current policy only accepts sodium wavelength lasers, and requires that a single laser beacon not exceed 50W and that a maximum of 200 W be projected from a single facility, and that laser beacons not be projected below 70º zenith angle. | Mauna Kea laser projection requirement |

| # | Observatory Operational Requirement | AO Derived Requirements | Instrument Derived Requirements |
| --- | --- | --- | --- |
| 1 | The NGAO system must be facility-class | Facility-class has many implications on safety, operability, reliability, maintainability, lifetime, documentation, configuration management, etc. |  |
| 2 | The NGAO system must operate within specifications under the normal summit temperature and humidity conditions | See the conditions specified in the Instrument Baseline Requirements Document. |  |

### Observatory Implementation Requirements

| # | Observatory Implementation Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | The time between decommissioning an AO capability on the telescope where the NGAO system is to be installed and making NGAO available for limited shared-risk science must not be longer than 6 months | Down time impact on Interferometer and AO science at the Observatory. |  |
| 2 | The telescope downtime required to implement NGAO must not be longer than 5 days |  |  |
| 3 | The NGAO system and instruments must complete the Observatory standard design review process. |  |  |
| 4 | The NGAO system must complete an operations transition review where operational responsibility is transferred from development to operations | This has implications on defining transition requirements and on training. |  |

# Optical Requirements

## Purpose and Objectives

The purpose of this section is to describe optical requirements for the performance, implementation and design of the NGAO optical system.

## Performance Requirements

The following performance requirements are duplicated from the Science Performance Requirements in section 6.1.2 since these are direct optical performance requirements.

| # | Performance Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | Telescope plus NGAO transmission to the input of the science instruments ≥ 70% at 0.7-2.4 µm. | The wavelength range is explicitly identified in the SCRD. The transmission is a placeholder. Need to determine if we should extend this down to H (656.3 nm). May be better to write in terms of SNR. | KAON 455 (v1) Table 4 |
| 2 | Goal: Telescope plus NGAO transmission to the input of the science instruments ≥ 70% at L-band. |  | KAON 455 (v1) Table 4 low mass stars & Galactic Center science cases |
| 3 | NGAO background, including the science instrument shall be ≤ 100% of the sky plus telescope at K-band. | The measured NIRC2 K-band background is 12.24 mag/arcsec2. The predicted sky background is 13.46. | KAON 455 (v1) Table 4: for the asteroid science cases the background should be ≤ the current LGS background |
| 4 | Goal: NGAO background including the science instruments shall be ≤ 20% of the sky plus telescope at K-band. |  | KAON 455 (v1) Table 4: High z galaxies. |
| 5 | Wavefront error ≤ 140 nm rms for V ≤ 17 on-axis guide star |  | KAON 455 (v1) Table 4: asteroid shape & companions. |
| 6 | Wavefront error ≤ 140 nm rms for V ≤ 16 guide star ≤ 30” from science object |  | KAON 455 (v1) Table 4: planets around low mass stars. |
| 7 | Wavefront error ≤ 170 nm rms for objects ≤ 5” from the Galactic Center |  | KAON 455 (v1) Table 4: Galactic Center. |
| 8 | Encircled energy ≥ 50% within a 0.05” diameter circle at K-band for sky coverage ≥ 5% |  | KAON 455 (v1) Table 4: High z galaxies. |
| 9 | Encircled energy ≥ 50% within a 0.075” diameter circle at K-band for sky coverage ≥ 30% |  | KAON 455 (v1) Table 4: High z galaxies. |
| 10 | The companion sensitivity shall be ΔH ≥ 5.5 mag at 0.5” separation for a V ≤ 17 on-axis guide star |  | KAON 455 (v1) Table 4: asteroid companions. |
| 11 | The companion sensitivity shall be ΔH ≥ 13 at 1” separation for a V ≤ 16 guide star ≤ 30” from science object |  | KAON 455 (v1) Table 4: planets around low mass stars. |
| 12 | Sensitivity? Does this just drive transmission? |  | KAON 455 (v1) Table 4: planets around low mass stars. |
| 13 | H-band photometric accuracy of ≤ 0.05 mag at 0.6” for ΔH = 3 for a V ≤ 17 on-axis guide star |  | KAON 455 (v1) Table 4: asteroid companions. |
| 14 | H-band photometric accuracy of ≤ 0.05 mag relative to the primary star at 1” separation for ΔH = 13 for a V ≤ 16 guide star ≤ 30” from science object |  | KAON 455 (v1) Table 4: planets around low mass stars. |
| 15 | “Uncalibrated detector distortion < 1.5 mas” | Is this really the way we want to write this? Shouldn’t we put the high level requirement here? | KAON 455 (v1) Table 4: asteroid companions. |
| 16 | “1/10 of the PSF FWHM” | Wouldn’t be better off writing this as mas? | KAON 455 (v1) Table 4: planets around low mass stars. |
| 17 | Astrometric accuracy ≤ 100 µas at K-band for objects ≤ 5” from the Galactic Center |  | KAON 455 (v1) Table 4: Galactic Center. |
| 18 | Radial velocity accuracy ≤ 10 km/sec at K-band for objects ≤ 5” from the Galactic Center |  | KAON 455 (v1) Table 4: Galactic Center. |
| 19 | Overheads between targets ≤ 10 min |  | KAON 455 (v1) Table 4: Asteroid size & companions, planets around low mass stars. |
| 20 | Strehl or PSF stability requirement? | TBD |  |

The following performance requirements are derived from the Science Performance Requirements in Section 6.1.2 and the relevant performance budgets.

| # | Performance Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | The residual static wavefront errors from the NGAO system shall be ≤ 50 nm rms as delivered to the visible and NIR imagers |  | Companion sensitivity requirement & performance budget. |
| 2 | The NGAO system shall introduce ≤ TBD % polarization in the path to the TBD science instruments |  |  |
| 3 | The stability of the focal ratio produced by the NGAO system shall be TBD as delivered to the TBD science instruments | The purpose of this requirement is to have a fixed plate scale. |  |
| 4 | Instability must not decrease the Strehl on the visible and NIR imagers by more than 5% during any 1-hour period or over a 3 ºC temperature change (Should look at changing the instability requirements to rms error.) | The image improvement provided by the facility must be stable over periods of hours, under stable atmospheric conditions. The wavefront sensor calibrations must be stable with time, temperature and telescope position. This requirement does not include performance degradations due to changing atmospheric conditions, or different magnitude and location of NGS or LGS. This requirement could be tested using a point source (i.e., single mode fiber) at the NGAO system input. |  |
| 5 | Instability must not decrease the Strehl on the visible and NIR imagers by more than 10% during any 24-hour period or over a 5 ºC temperature change |  |  |
| 6 | The maximum intensity of any ghost images (of a science object) produced by the NGAO system shall be no more than TBD the maximum intensity of the source within the science field. |  | Derived from the companion sensitivity requirement. |
| 7 | The maximum intensity of any ghost images (of a science object) produced by the imaging science instruments shall be no more than TBD the maximum intensity of the source within the science field. |  | Derived from the companion sensitivity requirement. |

## Implementation Requirements

| # | Optical Implementation Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | The NGAO optical axis shall be coincident to the telescope’s elevation axis to ≤ TBD. | This is intended to minimize pupil and image motion. |  |

## Design Requirements

NGAO is required to provide an AO corrected beam to each of the science instruments, including the Interferometer and ‘OHANA fiber injection module.  This may be accomplished with one or more AO systems.  The following requirements are valid for all of these science instruments.

| # | Optical Design Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | The focal ratio provided to the science instruments should be TBD |  |  |
| 2 | The exit pupil location provided to the science instruments should be TBD |  |  |
| 3 | The NGAO system and science instrument combination should be capable of keeping the field or pupil fixed on the science instrument. |  |  |
| 4 | The NGAO system shall have ≤ TBD of non-common path aberration delivered to the science instruments. |  |  |
| 5 | The NGAO system should be capable of correcting ≥ TBD nm of low spatial frequency (Zernikes 4 to 15) non-common path aberration in the science instruments. | Image sharpening can be used to correct for aberrations in the science instruments. New science instruments should be designed to have small optical aberrations. The interferometer and possible legacy instruments such as NIRC2 or OSIRIS may be allowed larger aberration budgets. |  |
| 6 | The peak-to-peak range of tip/tilt correction provided by NGAO shall be ≥ 3” on sky. | The existing Keck AO systems have a peak-to-peak range of 1.6” which has proven to be inadequate in windy conditions. |  |

The following design requirements are duplicated from the Science Performance Requirements in section 6.1.2 since these are direct optical performance requirements.

| # | Science Instrument Optical Design Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | Unvignetted contiguous fields shall be provided to the NIR and visible science imagers and single field IFUs. These fields should be centered to within 1” of the telescope’s optical axis. The maximum field size is 20”x20”. | Driven by the science requirements on the contiguous field science instruments. |  |
| 2 | Multiple unvignetted contiguous fields shall be provided to the NIR d-IFU. Each field should be ≥ 4” in diameter. | Driven by science requirements on the NIR d-IFU. |  |
| 3 | The unvignetted field of regard provided to the NIR d-IFU shall have a total area of ≥ 6 arcmin2 and shall have a maximum off-axis distance of ≤ 1.5’ with respect to the telescope’s optical axis. | Driven by science requirements on the NIR d-IFU. |  |
| 4 | NGAO shall provide appropriate outputs to ≥ TBD science instruments with a switching time of ≤ 5 min. |  |  |

The following design requirements are imposed by the non-interferometric science instruments.

| # | Science Instrument Optical Design Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | A wavelength range of 0.7 to 1.0 µm must be provided to the visible science instruments (imager and IFU) |  |  |
| 2 | A wavelength range of 1.0 to 2.45 µm must be provided to the NIR science instruments (imager, IFU and deployable IFU) |  |  |
| 3 | A wavelength range of 3.0 to 5.3 µm must be provided to the thermal NIR imager | This is currently low priority and should be discussed if this drives the design. |  |
| 4 | An unvignetted field of view ≥ 20”x20” must be provided to the science imagers (visible and NIR) | Should this only be 10x10” for the visible imager? |  |
| 5 | An unvignetted field of view ≥ 12” diameter must be provided to the visible IFU |  |  |
| 6 | An unvignetted field of view ≥ 16” diameter must be provided to the NIR IFU |  |  |
| 7 | A field of regard of ≥ 1.5’ diameter must be provided to the NIR deployable IFU with vignetting by all sources ≤ TBD over TBD % of the field of regard |  |  |
| 8 | A field of view of 25”x25” must be provided to the L and M-band imager | Low priority. |  |
|  | An unvignetted science target field ≥ 3” x 3.4” shall be provided to each channel of the NIR deployable IFU |  |  |
|  | The NGAO + deployable IFU system shall support simultaneous observations of at least two science targets separated by ≤ 5” | There is no requirement for simultaneous nearest-neighbor target distances less than 5”. |  |
|  | The NGAO + deployable IFU system shall support simultaneous observations of at least six science target fields inscribed within a 30” diameter field |  |  |
|  | NGAO shall be capable of compensating for focus changes due to changing filters or modes |  |  |

The following design requirements are imposed by the Interferometer and/or ‘OHANA support requirements.

| # | Interferometry Optical Design Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | A wavelength range of 1.1 to 14 µm must be provided to the Interferometer |  |  |
| 2 | A wavelength range of 1.1 to 2.45 µm must be provided to the OHANA injection module |  |  |
| 3 | A field of view of ≥ 1’ diameter must be provided to the Interferometer |  |  |
| 4 | A field of view of 5” diameter must be provided to the OHANA injection module |  |  |
| 5 | NGAO must be able to support a chopping mode for the interferometer. | The nuller requires small amplitude chopping with the AO loops closed at each end of the chop for alignment purposes. |  |
| 6 | The interferometer output of NGAO must be polarization matched to the interferometer output of the AO system on the other telescope in order to produce ≤ 3º of differential s-p phase shift | The current KI achieves polarization matching by keeping the number, angle and coatings of all reflections the same in the beam trains from each telescope. The differential s-p phase shift in the current KI is measured at 6º resulting in a loss in V2 of 0.003. |  |
| 7 | The interferometer output of NGAO must have the same image rotation as the interferometer output of the AO system on the other telescope |  |  |
| 8 | The interferometer output of NGAO must have the same pupil rotation as the interferometer output of the AO system on the other telescope |  |  |
| 9 | The interferometer output of NGAO must have the same longitudinal chromatic dispersion as the interferometer output of the AO system on the other telescope | Transmissive optics fabricated from different materials can have different amounts of longitudinal chromatic dispersion resulting in the loss of fringe visibility |  |
| 10 | The ratio of the Strehls from the interferometer output of NGAO and the interferometer output of the AO system on the other telescope must be ≤ 1.2 and ≥ 0.9. | A Strehl mismatch of 22% or an intensity ratio of 1.22 results in a V2 loss of 0.010. |  |
| 11 | NGAO must be able to accommodate the accelerometers needed to support the Interferometer | On the current AO bench one accelerometer is placed near the telescope focus and another near the output to the DSM. These are used to measure vibration along the optical path and are used in the fringe tracker control system. The accelerometer acquisition system is housed in an electronics rack in the AO electronics vault. |  |
| 12 | NGAO or NGAO in combination with a modified DSM must be capable of supporting the laser metrology beams from the interferometer | These metrology beams are a potential source of background light on the wavefront sensors |  |
| 13 | NGAO must incorporate the required tools and tolerances to support alignment to the interferometer | For example, the current AO bench hosts a corner cube to aid in aligning the interferometer to the optical axis of the AO system and telescope |  |
| 14 | NGAO or NGAO in combination with a modified DSM must provide a collimated 100 mm diameter beam to the interferometer | In the current AO system a removable (on a translation stage) dichroic beamsplitter, located between the deformable mirror and second off-axis parabola, folds the collimated beam to the DSM |  |
| 15 | The rms residual tilt at the NGAO system output to the interferometer should be ≤ 0.007” for TBD guide star. |  |  |

# Mechanical Requirements

## Purpose and Objectives

The purpose of this section is to describe mechanical requirements for the performance, implementation and design of the NGAO mechanical systems.

## Performance Requirements

| # | Mechanical Performance Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | The NGAO AO system shall not exceed a thermal dissipation budget, into the dome environment, of 100 W |  |  |
| 2 | The NGAO laser system shall not exceed a thermal dissipation budget, into the dome environment, of 100 W |  |  |
| 3 | The NGAO AO system shall not exceed a thermal dissipation budget at the top-end of the telescope, into the dome environment, of 50 W |  |  |

## Implementation Requirements

| # | Mechanical Implementation Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | The NGAO facility must allow a means to install new science instruments | For example, if a new instrument can only be installed with the overhead crane then a means of removing the appropriate portion of the AO enclosure roof would be required. |  |

## Design Requirements

| # | Mechanical Design Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | The maximum weight of the AO system on the Nasmyth platform shall not exceed 10,000 kg | The weight requirements are imposed by limits on what the telescope can support at various locations without changing its performance. |  |
| 2 | The maximum weight of the laser facility on the azimuth rotating part of the telescope shall not exceed 10,000 kg |  |  |
| 3 | The maximum weight of the beam transport system on the elevation portion of the telescope shall not exceed 150 kg |  |  |
| 4 | The maximum weight of the laser launch facility in the top-end module shall not exceed 150 kg |  |  |
| 5 | If mounted behind the f/15 secondary mirror, the launch telescope facility must allow for the removal, storage and installation of the f/15 secondary module |  |  |
| 6 | If mounted behind the f/15 secondary mirror, the launch telescope facility must not extend beyond the module in the x,y-directions and must not extend more than 1 m beyond the top of the telescope structure. |  |  |
| 7 | The NGAO facility must fit within the mechanical constraints of a Nasmyth platform (nominally the Keck II left Nasmyth platform) |  |  |
| 8 | The NGAO facility must provide access and space for the installation of each science instrument and a mechanical interface on which to mount each instrument |  |  |
| 9 | The NGAO facility must provide access for routine maintenance of the elevation bearing, elevation wrap, bent Cassegrain platform and mirror cell stairwell |  |  |
| 10 | The required glycol flow rate and pressure for cooling the NGAO facility shall not exceed TBD and TBD, respectively |  |  |

# Electronic/Electrical Requirements

## Purpose and Objectives

The purpose of this section is to describe electronic and electrical requirements for the performance, implementation and design of the NGAO electronic and electrical systems.

## Performance Requirements

## Implementation Requirements

## Design Requirements

| # | Electrical Performance Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | The entire NGAO facility must not exceed a total electrical power requirement of 30 kW |  |  |
| 2 | The NGAO facility cabling through the azimuth wrap must not require an area of more than TBD |  |  |
| 3 | The NGAO facility cabling through the left Nasmyth elevation wrap must not exceed an area of more than TBD |  |  |
| 4 | The NGAO facility cabling through the right Nasmyth elevation wrap must not exceed an area of more than TBD |  |  |

# Safety Requirements

## Purpose and Objectives

Safety is the paramount concern for all activities at the observatory.  The purpose of this section is to provide requirements related to specific safety concerns during the operation and handling of NGAO.

## Scope

The general Observatory safety requirements that are also applicable to NGAO are already contained in the Instrumentation Baseline Requirements Document.  This section covers the additional laser safety and laser projection safety requirements.

Adequate earthquake restraints are also required for all systems.

## Laser Safety Requirements

A safety system will be implemented to ensure the safe use of the laser.  This safety system will include both engineering and administrative/procedural controls to assure safe operations.  The system will apply ANSI Z136.1 and Z136.6 standards for safe use of laser for indoor and outdoor.

The NGAO system will conform to OSHA and local codes in addition to codes specified for each subsystem.

## Laser Projection Safety Requirements

This section covers the additional requirements on safety for projecting the laser beams outside the dome.

### Aircraft Safety

An aircraft safety system compliant with FAA requirements must be implemented and approved by the FAA.

### Space Command

A system must be implemented to facilitate effective communication with U.S. Space Command of projection dates and targets, and to ensure that no projection occurs on a target without Space Command approval.  This requirement may be deleted if it is determined that Space Command approval is no longer required.

# Software Requirements

## Purpose and Objectives

The software requirements section describes requirements for performance, implementation and design.

## Scope

Unless otherwise indicated all of the requirements of this section apply to all software components of NGAO.

## Performance Requirements

## Implementation Requirements

## Design Requirements

| # | Software Design Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | The NGAO system must be able to support the external interfaces supported by the existing AO systems | The external keywords and EPICS channels used by the existing AO systems are documented in KAON 315, Summary of External Interfaces in the Current WFC and Implications for the NGWFC Design” |  |
| 2 | The NGAO system must be able to accept tip/tilt offloads from the science instrument |  |  |
| 3 | The NGAO system must be able to acquire and configure the AO system based on the star magnitude and spectral type provided in the target list |  |  |
| 4 | NGAO must be able to support a chopping mode for the Nuller mode of the interferometer |  |  |

# Interface Requirements

## Purpose and Objectives

This section is reserved for interface requirements that are not addressed by other portions of the document.

## Performance Requirements

## Implementation Requirements

## Design Requirements

### Optical Interface

### Mechanical Interface

| # | Mechanical Interface Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | An agreed upon kinematic interface between the NGAO opto-mechanical systems and the telescope structure must be provided | Kinematic interfaces are proposed so that thermal changes do not distort optical benches. |  |
| 2 | NGAO must be compatible with the interferometer dual star module (DSM) or replicate its functionality | The current Keck AO systems interface to the interferometer via an opto-mechanical system known as the DSM. |  |
| 3 | An agreed upon mechanical interface between the NGAO electronics and the telescope must be provided |  |  |
| 4 | An agreed upon mechanical interface between any NGAO enclosures and the telescope must be provided. |  |  |
| 5 | An agreed upon mechanical interface between any NGAO glycol cooled systems and the telescope instrument and/or facility glycol systems must be provided |  |  |
| 6 | An agreed upon mechanical interface between any NGAO CCR-cooled systems must be provided |  |  |

### Electrical/Electronic Interface

| # | Electrical Interface Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | An agreed upon electrical interface for power between NGAO systems and the Observatory /telescope must be provided | The Observatory is responsible for providing power to the various NGAO system locations. |  |
| 2 | An agreed upon interface for communication between NGAO systems and the Observatory/telescope must be provided | The Observatory is responsible for the implementation of all cables between NGAO system locations. The Observatory must in particular approve all cabling required to go in the elevation or azimuth wraps. |  |

### Software Interface

| # | Electrical Interface Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | An interface should be provided to offload tip/tilt errors to telescope pointing through the telescope drive and control system (DCS) |  |  |
| 2 | An interface should be provided to offload focus errors to the secondary mirror piston through DCS |  |  |
| 3 | An interface should be provided to offload coma errors to the secondary mirror tilt through DCS |  |  |
| 4 | An interface should be provided to offload segment stacking errors to the active control system (ACS) |  |  |
| 5 | An interface should be provided to write all appropriate NGAO keywords to the science instrument FITS header |  |  |
| 6 | An interface should be provided to allow NGAO to read/write NGAO science instrument keywords |  |  |

# Reliability Requirements

## Purpose

A process should take place to confirm that the NGAO system will provide a high level of reliability for a 10 year lifetime.

## Scope

Unless otherwise indicated all of the requirements of this section apply to all components of NGAO.

## Performance

System downtime should be minimized by a combination of component reliability, ease of repair, maintenance and appropriate sparing.

| # | Reliability Performance Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | ≤ 5% of observing time lost to problems | This includes any loss to an exposure in progress and the time to start the next exposure after recovering from a fault |  |
| 2 | The median time between faults during observing time should be ≥ 3 hrs | Frequent short duration faults are not acceptable since they have a high impact on science productivity |  |

# Spares Requirements

TBD

# Service and Maintenance Requirements

TBD

# Documentation Requirements

The documentation requirements are defined in the Instrumentation Baseline Requirements Document.

## Documentation Package

## Drawings

## Electrical/Electronic Documentation

## Software

# Glossary

Table 2 defines the acronyms and specialized terms used in this document.

Table 2 Glossary of Terms

| Term | Definition |
| --- | --- |
| ACS | Active Control System |
| ANSI | American National Standards Institute |
| AO | Adaptive Optics |
| DCS | Drive and Control System |
| DSM | Dual Star Module |
| FAA | Federal Aviation Administration |
| FOV | Field Of View |
| FWHM | Full Width at Half Maximum. |
| IFU | Integral Field Unit |
| KAON | Keck Adaptive Optics Note |
| KI | Keck Interferometer |
| LGS | Laser Guide Star |
| MTBF | Mean Time Between Failures |
| NGAO | Next Generation Adaptive Optics |
| NGS | Natural Guide Star |
| NIR | Near InfraRed |
| NIRC2 | NIR Camera 2 |
| NIRSPEC | NIR SPECtrometer |
| ‘OHANA | Optical Hawaiian Array for Nanoradian Astronomy |
| OSHA | Occupational Safety and Health Administration |
| OSIRIS | OH-Suppression InfraRed Integral field Spectrograph |
| TBC | To Be Completed |
| TBD | To Be Determined |
| WMKO | W. M. Keck Observatory |