## Key and Notes
KNGAO Real-Time Controller (RTC)
Requirements spreadsheet and design conformance report
Conformance Color Code
|  | Green | Design conforms to requirement
|  | Yellow | There is a problem with the requirment (noted)
|  | Red | Push-back
|  | White | Outside of scope of RTC design
Notes
Blank traceability entry indicates that the requirement is not traceable to a higher level requirement

## General
General
Short Name | ID | Priority | Description | Rationale | Traceability | Conformance
Real Time Controller | FR-1401 | Essential | The AO Real Time Controller (RTC) system provides the main system for correction of atmospherically distorted wavefronts. In its primary mode, when used with laser guider stars, the RTC analyzes the LGS wavefront from 7 laser beacons. In addition, inputs from three NGS wavefront sensors will be used to estimate tilt and other low order aberrations such as focus and astigmatism. The totality of these wavefront sensor inputs is used to estimate the three dimensional volume of atmospheric turbulence over the telescope. This estimation step is commonly called tomography. The RTC is responsible for using the tomographic solution to determine the shape applied to the various deformable mirrors in the AO system. The RTC also receives input on the difference between LGS and NGS measurement of atmospheric turbulence from a NGS truth wavefront sensor (TWFS). Additionally, the RTC will support a NGS mode (no LGS) where natural guider stars will provide all the wavefront information. | The RTC is the wavefront estimator and controller in the NGAO system. | KAON 511: NGAO System Design Manual
KAON 642: NGAO Design Changes in Support of the Build to Cost Guidelines
Source of Control | FR-1403 | Essential | The RTC shall be under the supervisory control of the AO Control System. With the exception of the items listed below, the RTC shall change state or take actions only in response to commands from the AO Control System. Exceptions: | The RTC is a slave subsystem of the AO Control System. | KAON 511: NGAO System Design Manual
|  |  | 1. The RTC shall take whatever action is required to protect itself and its system components from damage. |  | KAON 642: NGAO Design Changes in Support of the Build to Cost Guidelines
2. The RTC shall perform an automated sequence upon power up and load a set of default operating parameters.
RTC Control functions | FR-1405 | Essential | RTC control functions: | These are all required functions for the RTC. | KAON 511: NGAO System Design Manual
|  |  | a) The RTC shall provide time-averaged telemetry to the AO Control System for use in low-order mode offloading loops. |  | KAON 642: NGAO Design Changes in Support of the Build to Cost Guidelines
b) The RTC shall prevent the build up of un-sensed modes, such as piston and global waffle.
c) The RTC shall compute commands for deformable mirror actuators that fall outside the illuminated foot print of the deformable mirror, i.e. slaved actuator commands.
d) If needed the RTC will provide functionality for determining the shape of the common deformable mirror in the first stage of the cascaded relay.
Real Time Computer | FR-1425 | Essential | The RTC shall be implemented using a dedicated real-time computer system. | Standard practice for AO wavefront control. | KAON 511: NGAO System Design Manual
KAON 642: NGAO Design Changes in Support of the Build to Cost Guidelines

## Interface
Interface
Short Name | ID | Priority | Description | Rationale | Traceability | Conformance
TWFS interface | FR-1412 | Essential | The RTC shall receive NGS Truth Wavefront Sensor (TWFS) updates on a periodic basis from the AO Control System. The update rate shall be no less than every 1 second. | The RTC will receive information about the NGS wavefront in order to correct for structure in the sodium layer and elongation of the LGS. This is the equivalent of the LBWFS in the current AO system. | KAON 511: NGAO System Design Manual
[5 empty] | KAON 642: NGAO Design Changes in Support of the Build to Cost Guidelines | this is slow data that should be handled by a non-real-time system that is part of the AO Control scope
AO Control System Interface | FR-1418 | Essential | The RTC shall be under the supervisory control of the AO Control System and shall have an interface with the control system for this purpose. The AO Control System shall be responsible for updating all RTC parameters, checking the health of the RTC, and marshalling slow speed telemetry and diagnostics from the RTC. | The RTC is a slave subsystem of the AO Control System. | KAON 511: NGAO System Design Manual
KAON 642: NGAO Design Changes in Support of the Build to Cost Guidelines
KAON 679: NGAO Control Software Architecture
Data Server Interface | FR-1419 | Essential | The RTC shall have an interface to the NGAO Data Server for the recording of telemetry and parameter data. The interface shall be via ethernet with a minimum speed of 100 Base-T. The definition of the telemetry stream is TBD, but will likely include the following: | The Telemetry Recorder/Server (TRS) in the current Keck AO system has proven to be extremely valuable. We want to implement a similar recording capability for NGAO. | KAON 511: NGAO System Design Manual
KAON 642: NGAO Design Changes in Support of the Build to Cost Guidelines
|  |  | 1. Average raw camera frames |  | KAON 679: NGAO Control Software Architecture
2. Average WFS centroids
3. Average wavefront inputs (wavefronts produced by the WFSs)
4. Average wavefront outputs (high-order and low-order wavefront outputs produced by the RTC).
5. All RTC parameter changes
Note: This does not include any full frame rate data, which is handled in another requirement.
DM and TTM interface | FR-1422 | Essential | The RTC will have electrical interfaces with the AO deformable mirrors and tip tilt correctors listed below: | The RTC must control all DMs and TTMs. | KAON 511: NGAO System Design Manual
|  |  | 1 wide-field relay DM |  | KAON 642: NGAO Design Changes in Support of the Build-to-Cost Guidelines.
1 wide-field tip/tilt stage (carries DM above)
1 LOWFS MEMs DM on t/t stage in each LOWFS wavefront sensor (3 total)
1 narrow field DM (MEMS)
1, 4, or 7 LGS WFS TT stages
WFS detector and camera interface | FR-1423 | Essential | The RTC shall receive pixel data directly from the detectors used in the LGS WFS, the LOWFS, and the NGS WFS. Control of these devices (frame rates, etc.) will be performed by the AO Control System. Note that some of these devices may have a shared data/control interface. In those cases, direct control of the device will be via the RTC, but the RTC will expose those control features to the AO Control System, so that the AO Control System is the one that initiates the control of these devices. | The RTC must receive pixel data directly to minimize latencies. The AO Control System is responsible for controlling the entire AO system, including the cameras and detectors. If the control interface for a detector is shared with the data interface (which is fed directly to the RTC), then the RTC must expose appropriate control functionality for the cameras and detectors.
High Speed RAID Interface | FR-2253 | Essential | The RTC will have an interface to a high-speed RAID system for the recording of both high-speed telemetry and parameter data. This will require a specialized hardware interface to the RAID system to handle the high throughput requirements (e.g., Fibre Channel, or similar). The data rate is sufficient to capture telemetry data listed in FR-1419 at the highest frame rate. | A high speed disk farm is called out in the RTC design. | NGAO Real Time Controller Processing Engine Design Document, v1.1.134, August 9, 2009.

## Performance
Performance
Short Name | ID | Priority | Description | Rationale | Traceability | Conformance
Bandwidth and Latency | FR-1406 | Essential | 1. The high-order bandwidth of the RTC in LGS mode shall be at least 100 Hz. | The control bandwidth and latency of the RTC must be sufficient to meet the error budget requirements. | Bandwidth / Latency Flowdown Budget:
|  |  | The tip-tilt bandwidth of the RTC in LGS mode shall be at least 30 Hz. |  | http://www.oir.caltech.edu/twiki_oir/pub/Keck/NGAO/FlowdownSummary/RTC_errorBudget_per_KAON644.xls
The maximum latency measured from the receipt of the final WFS pixel data to the start of the sending of commands to the various DMs in LGS mode shall be no greater than 1 ms
The maximum latency measured from the receipt of the final WFS pixel data to the start of the sending of commands to the various TTMs in LGS mode shall be no greater than 1 ms
2. The high-order bandwidth of the RTC in NGS mode shall be at least 100 Hz.
The tip-tilt bandwidth of the RTC in NGS mode shall be at least 30 Hz.
The maximum latency measured from the receipt of the final WFS pixel data to the start of the sending of commands to the various DMs in NGS mode shall be no greater than 1 ms
The maximum latency measured from the receipt of the final WFS pixel data to the start of the sending of commands to the various TTMs in NGS mode shall be no greater than 1 ms
Numeric Precision | FR-1408 | Unassigned | The RTC numerical precision including digitization of the drive signals sent to the deformable mirrors must be consistent with a final wavefront error of 1 nm rms. | From Don Gavel: We should state the required numerical precision and tie it to the error budget for justification. Otherwise a potential vendor might attempt to do it with 8-bit integer arithmetic, which would introduce unacceptable error. This is a more relevant issue for FPGA implementation than it would be for a "conventional" floating point processor, where we might get away with assuming it uses standard high precision arithmetic. | Bandwidth / Latency Flowdown Budget:
http://www.oir.caltech.edu/twiki_oir/pub/Keck/NGAO/FlowdownSummary/RTC_errorBudget_per_KAON644.xls
Accuracy of wavefront reconstruction | FR-1434 | Essential | The accuracy of the wavefront information to be applied to the deformable mirrors shall be 5 nm rms. | Wavefront reconstruction accuracy must be specified. | Wavefront reconstruction error budget
Accuracy of tip/tilt reconstruction | FR-1435 | Essential | The accuracy of the RTC tip-tilt mirror commands shall be 1 milliarcsecond (on sky) or better. | This is an assignment from an error budget, and controllable via design of the RTC | Wavefront reconstruction error budget
Update rate | FR-1436 | Essential | The RTC shall be capable of operating at a maximum loop rate of at least 2 kHz when in full laser guide star mode. | High Strehl requirement drives loop rate. | Bandwidth / Latency Flowdown Budget:
|  |  | The RTC shall be capable of operating at a maximum loop rate of at least 2 kHz when operating in single natural guide star mode. | Lower limit is an estimate of what is required for engineering and troubleshooting purposes. | http://www.oir.caltech.edu/twiki_oir/pub/Keck/NGAO/FlowdownSummary/RTC_errorBudget_per_KAON644.xls
|  |  | The RTC shall be capable of operating in a range of loop rates from 100 Hz to 2 kHz as observing conditions require. Lower limit is 100 Hz | No observing requirement is driving the lower limit that we know of.
Note that it is possible to implement low bandwidth control using a high loop rate, so a fast loop rate does not preclude slow updates based on longer WFS integration times.
Wavefront sensor input data rate | FR-1437 | Essential | The RTC shall be able to accept wavefront sensor camera data on each camera input interface at a maximum rate of 2000 frames per second. The maximum frame size is expected to be 256 x 256 pixels. | The RTC must be capable of running at a 2 kHz sustained rate. Not all wavefront sensor cameras may run at this rate simultaneously, but each input port must be able to accept a 2 kHz sustained input rate. | Bandwidth / Latency Flowdown Budget:
http://www.oir.caltech.edu/twiki_oir/pub/Keck/NGAO/FlowdownSummary/RTC_errorBudget_per_KAON644.xls
Startup shutdown time | FR-1454 | Essential | Startup of the RTC from a powered off state shall take no more than 20 minutes. | Need reasonable startup and shutdown times to meet system uptime and efficiency requirements.
Shutdown of the RTC from an operational state to power off shall take no more than 10 minutes.
DM data rate | FR-1456 | Essential | The RTC shall be able to send commands to each of the DM's in the system at a minimum rates of 4kByte/sec/actuator. | The RTC must be able to send commands to all of the DMs at the fastest loop rate (2kHz). The DM command interfaces are 16 bits wide (even though the commands may use fewer bits). | Bandwidth / Latency Flowdown Budget:
|  |  | Note: there are 5 DMs in the system: Woofer (20x20 actuators), Tweeter (64x64 actuators), LOWFS correctors (32x32). |  | http://www.oir.caltech.edu/twiki_oir/pub/Keck/NGAO/FlowdownSummary/RTC_errorBudget_per_KAON644.xls
The corresponding rate on a per DM basis is:
Wide-field DM (woofer): 1.6MB/Sec
Narrow-field DM (tweeter): 16.4MB/Sec
LOWFS DMs: 4.1 MB/Sec
Atmospheric information update rate | FR-1438 | Essential | The RTC shall be able to accept updates to Cn2, layer height, layer wind velocity from the AO Control System at a minimum rate of once per 90 seconds. | The AO Control System must forward atmospheric profiler data to the RTC on a periodic basis. | See KAON 552 on requirements for the atmospheric profiler. Ultimately, this is tied to an open question of how a-priori knowledge of the atmospheric profile impacts performance at any given rate.

## Physical
Physical
Short Name | ID | Priority | Description | Rationale | Traceability | Conformance
Size | FR-1440 | Essential | The RTCshall occupy a space not greater than 8 cubic meters volume. | This needs to be based on a real space constraint for the AO enclosure on the Nasmyth platform. This is a number the RTC designers feel comfortable with. | AO enclosure design
Weight | FR-1441 | Essential | The weight of the RTC shall be less than 10,000 kg. | The weight must be such that the instrument can be mounted on the Nasmyth platform of the Keck facility. | See SR-702. Nasmyth platform weight limits.
[5 empty] | This is the total weight allocated to the Nasmyth per KAON 456. There is no published weight budget with allocations to subsystems | no published weight budget
Power Dissipation | FR-1442 | Essential | The power dissipation of the RTC shall be less than 20 kW. | The power dissipation must be such that the power can be removed from the RTC without causing problems in dome seeing.
Location | FR-1451 | Important | If possible, the RTC shall be located on the Nasmyth platform of the telescope. However, if required (due to size, weight, power, or heat restrictions), the RTC may be located in the computer room. This assumes the increased signal transmission time due to the cable distance between the Nasmyth platform and the computer room (appx 260 ft) does not adversely impact the overall AO control loop latencies. If the RTC is located in the computer room, fiber optic cables will be required to transmit the sensor data from the WFS cameras to the RTC, and to transmit the DM and TT mirror command data from the RTC to the mirrors. | The RTC must comply with the volume, weight, power and heat dissipation limits in order to be installed in the AO enclosure on the Naysmyth platform.

## Reliability
Reliability
Short Name | ID | Priority | Description | Rationale | Traceability | Conformance
Single Event Upset (SEU) rate | FR-1443 | Essential | The RTC shall be able to detect and automatically recover from cosmic or gamma ray events. The design of the detection and recovery scheme shall be consistent with the system uptime requirements. | The RTC must be able to operate in the environment with minimum impact to science observing. Moderate shielding may need to be included to reduce SEU's to an acceptable level. | SR-766, SR-731, SR-730.
Uptime | FR-1449 | Essential | The Mean Time Between Failures (MTBF) shall be consistent with the overall NGAO MTBF and uptime requirements, 140 hours | The RTC must not impact NGAO operations. | SR-766, SR-731, SR-730.
Maintenance | FR-1450 | Essential | The RTC shall be designed to support routine maintenance by Keck Observatory personnel: | The RTC must be maintaineable and must not adversely impact NGAO operations. | See SR-766, SR-730, SR-731.
1. It shall be designed using easily replaceable and swappable modules so that bad modules may easily be identified and replaced.
2. It shall be designed to allow easy access to all modules and sub assemblies.
3. It shall be designed to function at normal loop rates with diagnostic equipment installed, such as card extenders, bus/logic analyzers, or other test equipment.
4. The Mean Time To Repair (MTTR) shall be consistent with the system uptime requirements, 4 hours.

## Functional
Functional
Short Name | ID | Priority | Description | Rationale | Traceability | Conformance
LGS WFS Processing | FR-1409 | Essential | LGS WFS processing: | The RTC must implement LGS wavefront sensor computations for input to the tomography engine and implement laser tip-tilt control to stabilize the beams on the LGS WFSs. | KAON 511: NGAO System Design Manual
|  |  | a) The RTC will receive WFS pixels from multiple LGS sensors and calibrate them for sensor and sky backgrounds. This includes a provision for subtraction of Rayleigh background if CW sodium lasers are used. |  | KAON 642: NGAO Design Changes in Support of the Build to Cost Guidelines
b) The RTC will estimate wavefront slopes from WFS pixel intensities using a choice of center-of-mass, weighted center-of-mass, and quad-cell algorithms.
c) The RTC will update its estimate algorithm of the wavefront slopes as atmospheric and sodium layer conditions change.
d) The RTC will estimate the full aperture tip-tilt from each LGS wavefront and export this information to the appropriate subsystem for LGS tip-tilt control, also know as uplink tip tilt. (This subsystem is located in each LGS WFS channel). The specific update rate is 2 kHz
e) If the selected laser is pulsed, to reduce background form Rayleigh scattering, the LGS WFS CCD should be synchronized with laser pulse rate. The timing error and clock source are 1 microsecond, respectively.
f) The RTC shall allow adjustable frame rates for LGS observations.
LOWFS processing | FR-1410 | Essential | Low-order wave front sensor (LOWFS) high-bandwidth NGS processing (LGS AO mode): | Design decision based on current AO practice. | KAON 511: NGAO System Design Manual
|  |  | a) The RTC shall receive WFS pixels from multiple low order NGS and calibrate them for sensor and sky backgrounds. |  | KAON 642: NGAO Design Changes in Support of the Build to Cost Guidelines
b) The RTC shall estimate wavefront slopes from pixel intensities using a center-of-mass algorithm
c) The RTC shall update its estimate of the wavefront slopes as atmospheric conditions and AO performance change.
d) The AO Control System shall optimize the setting of the NGS WFS based on NGS star brightness, atmospheric conditions, and AO performance.
e) The AO Control System shall adjust frame rates of the tip-tilt sensors between 25 Hz to 2.0 kHz as the brightness of the NGS changes. The rate of each LOWFS sensor will be independently selectable.
Truth WFS processing | FR-1411 |  | Truth WFS Processing (TWFS) LGS AO mode: | Design decision based on current AO practice. | See above | not in the scope of the RTC
a) Calibrate slow variations and biases in AO control from changes in sodium layer profile, flexures, and drifts. At this time, the expected update rate of the Truth wavefront sensor maybe slow enough that the information could be processed by either an RTC subsystem or a dedicated TWFS software that would perform wavefront reconstruction on the TWFS input and not the RTC. For the purposes of this document, we assume that the TWFS software performs the reconstruction.
|  |  | b) The TWSF software will provide low order NGS wavefront information | See also requirements FR-196, FR-213, FR-216, and FR-219.
c) The focus information from the TWFS will be used to change the focus of the LGS wavefront sensors.
High-order WFS Processing | FR-1413 | Essential | High-order NGS WFS Processing NGS (NGS AO): | Design decision based on current AO practice. | KAON 511: NGAO System Design Manual
|  |  | a) The RTC shall receive WFS pixels from a single NGS and calibrate them for sensor and sky backgrounds. |  | KAON 642: NGAO Design Changes in Support of the Build to Cost Guidelines
b) The RTC shall estimate wavefront slopes from the CCD pixel intensities using a choice of center-of-mass, weighted center-of-mass, or quad cell algorithm.
c) The AO Control System shall update the slope estimation algorithm for seeing and AO performance at a rate of once per minute.
d) The AO Control System shall adjust the frame rate of the wavefront sensor between 100 Hz and 2000 Hz as NGS guide star brightness varies.
Atmospheric Correction | FR-1414 | Essential | Atmospheric Correction function of the Real-Time Controller: | Design decision based on current AO practice. | KAON 511: NGAO System Design Manual
|  |  | a) In LGS mode, the RTC shall take data from LGS WFS cameras and the LOWFS cameras to: |  | KAON 642: NGAO Design Changes in Support of the Build to Cost Guidelines
a.1) produce a tomographic estimate of the atmospheric turbulence above the telescope using a Fourier-domain preconditioned back-projection tomography algorithm.
a.2) produce commands, in volts, to drive all the DMs and fast Tip/Tilt mirrors in the AO system
a.3) produce offloading information (focus, astigmatism, coma, and telescope guiding), flowed up to the AO Control System.
b) The algorithm for tomography will compute the minimum variance estimate of the volume-distributed turbulence, accounting for the a-priori statistics of the atmospheric turbulence and statistics of wavefront sensor measurement noise.
c) The RTC shall, in real time, incorporate parametric information about the LGS constellation, the NGS constellation for the LOWFS, the science mode and instrument, the atmospheric turbulence profile, wavefront sensor signal levels, the orientation of the telescope pupil, the orientation between wavefront sensor measurements, and AO deformable mirror actuators.
d) In NGS mode, the RTC shall estimate the wavefront given data from the NGS WFS camera, and distribute the solution, in the form of voltage commands, to the woofer, tweeter, and fast tip/tilt mirror.
Stable Compensator | FR-1415 | Essential | The RTC shall implement a stable compensaotor in its servo loops for all the deformable and tip-tilt mirrors to ensure control loop stability. | Design decision based on current AO practice. | KAON 511: NGAO System Design Manual
KAON 642: NGAO Design Changes in Support of the Build to Cost Guidelines
Telemetry data | FR-1416 | Essential | The RTC shall provide the following types of telemetry products: | Telemetry of various types is required for normal system operation (acquisition, offloading), calibration, engineering purposes, debugging and troubleshooting. | KAON 511: NGAO System Design Manual
KAON 642: NGAO Design Changes in Support of the Build to Cost Guidelines
1. High-speed telemetry: telemetry that is posted at the full frame rate or loop rate at which it is generated.
2. Time averaged telemetry: telemetry that is time averaged over a user specified interval and posted at a (possibly different) user specified interval.
3. Single instance: telemetry that is provided in response to a single user request. The telemetry may be from either high-speed data or time averaged data.
The high-speed telemetry shall flow directly to a high-speed RAID (disk farm).
The time-averaged and single instance telemetry shall be provided to the AO Control System.
The high-speed telemetry products shall include:
1. Raw Camera Data
2. Centroids
3. Wave Fronts
4. Tomographic Layers
5. Science on-axis High Order Wave Front
6. Science on-axis Woofer Wave Front
7. Science on-axis High Order Commands
8. Science on-axis Woofer Commands
9. RTC Current Parameter Settings
The time-averaged and single instance telemetry products may include any of the data products listed above.
Calibration | FR-1417 | Essential | The RTC shall provide the appropriate data to the AO Control System for use in calibrating the wavefront sensors and correctors in the AO system. Calibration functions include: | We must be able to calibrate the AO system.
a) The deformable and tip-tilt mirror actuator gains.
b) The WFS linearity.
c) The DM to lenslet registrations.
d) The determination of the pupil orientation and illumination of various correctors and wavefront sensors in the NGAO system.
Control Parameter Information | FR-1426 | Essential | The AO Control System shall provide the RTC with all parameter data required for normal operation. The RTC shall incorporate any new parameter data into its real time processing within one second of receipt from the AO COntrol System. The parameters shall include the following: | The AO Control System must upload the required parameter data to the RTC. These parameters are detailed in "NGAO Hard Real Time and Soft Real Time Controller Interface", M. Reinig. | KAON 511: NGAO System Design Manual
|  |  | 1) Layer information (Height and Cn2) |  | KAON 642: NGAO Design Changes in Support of the Build to Cost Guidelines
2) Gain(s) for reconstruction an other system functions
3) Pre conditioning (and other) matrices
4) Kolmogorov profile
5) Guide star cone effect scaling factors
7) Tomography, centroid, and any other algorithm selection
8) All look up tables and matrixes for DM or Centroid processing
Standard units for reported data | FR-1427 | Essential | All atmospheric data computed by the RTC shall be reported in nanometers. The RTC shall provide tomography over altitude at 5 altitudes, from ground up to 15 km. The RTC shall provide a turbulence estimate with a height resolution of at least 8-bits. | Design decision: currently standard method of reporting atmospheric turbulence information in adaptive optics | KAON 621: The Noise Propagator for Laser Tomography Adaptive Optics
LGS WFS inputs | FR-1428 | Essential | The RTC shall receive wavefront sensor data from 7 laser guide stars: 4 central asterism laser guide stars and 3 point-and-shoot laser guide stars. | Need to support 4 LGSs in the central asterism and 3 point-and-shoot LGSs. | KAON 511: NGAO System Design Manual
KAON 642: NGAO Design Changes in Support of the Build to Cost Guidelines
Natural guide star (Tip/Tilt/Focus/Astigmatism) processing in LGS mode | FR-1429 | Essential | The RTC shall take input from 3 NGS sensors. Two of the sensors are imagers from which tip/tilt information shall be extracted. One of the sensors is a low-order wavefront sensor from which tip/tilt/focus/astigmatism information shall be extracted. The RTC shall take the raw pixel data from these sensors and produce the appropriate tip/tilt/focus/astigmatism measurements. The RTC shall take the tip/tilt/focus/astigmatism measurements and process them into the overall tomography calculations. | Tip/tilt/focus/astigmatism is required in order to compensate for low-order blind modes due to the fact that LGS does not provide tip/tilt information. | KAON 511: NGAO System Design Manual
KAON 642: NGAO Design Changes in Support of the Build to Cost Guidelines
Tip-tilt mirror control | FR-1432 | Essential | The RTC shall provide control information to 11 tip/tilt stages: | Laser guidestars need tip/tilt correction in order to stay nominally centered on the linear region of the LGS wavefront sensor. It was decided to put these in the downlink path as opposed to the uplink path in order to minimized round-trip delay response. Note: the number of TT mirrors in the LGS WFS assembly has not been finalized. It may be 1 (for all 7 WFSs), 4 (1 for central asterism and 3 for the point-and-shoot), or 7 (1 for each LGS). | This supports the present system architecture, described in KAON 642, "Design Changes in Support of Build-to-Cost".
7 downlink laser guidestar tip/tilt stages in LGSWFS units
|  |  | 3 natural guidestar tip/tilt stages in the LOWFS units | The current desing for the LOWFS sensors has them operating in tip/tilt nulling mode, therefore they need tip/tilt control mirrors.
|  |  | 1 tip/tilt stage which mounts the woofer DM | The tip/tilt stage of the woofer provides all of the tip/tilt correction for the science light.
WFS design | FR-1444 | Essential | The wavefront sesning computations in the RTC shall be of Shack Hartmann design. | The NGAO wavefront sensors will be of Shack Hartmann design. | KAON 511: NGAO System Design Manual
KAON 642: NGAO Design Changes in Support of the Build to Cost Guidelines
WFS sub aperture size in pixels | FR-1445 | Essential | The RTC shall implement a WFS subaperture size of 4 x 4 pixels.
DM influence function and non linearity compensation | FR-1446 | Essential | The RTC shall calculate the DM commands necessary to best fit to the desired wave front. This includes compensating for the DM influence function as well as any other non-linear response functions. | The RTC must compensate for the DM influence function and any actuator non-linearities in order to achieve a best fit to the desired wavefront.
Tomography algorithm | FR-1447 | Essential | The RTC shall be implement the tomography algorithm described in KAON XXX (The RTC Algorithm Design Document). | The RTC design must implement the desired tomography algorithm; the two are inseperable and must be compatible with each other. | KAON 511: NGAO System Design Manual
KAON 642: NGAO Design Changes in Support of the Build to Cost Guidelines
Automated operation | FR-1448 | Essential | The RTC shall implement automated operation under the supervision of the AO Control System. This shall include automatic startup and verification of status upon initialization as well as an orderly shutdown in response to a standby command. The RTC shall provide diagnostics at the request of the AO Control System for verification of data and operation. The RTC shall inform the AO Control System in the event of corrupt data. | Operational model for TMT site testing systems. See "Characterizing the TMT Site Selection Equipment", Matthias SchÃ¶ck for the TMT Site Selection Team, Symposium on Seeing, Kona Hawaii, March 20, 2007, http://mkwc.ifa.hawaii.edu/symposium.
Time Stamp data | FR-1452 | Essential | The RTC shall implement industry standard timing referenced to Coordinated Universal Time (UTC). All data products produced by the RTC shall have timestamps associated with them. The accuracy of the timestamps shall be 10 usec or less. Timestamps shall include the start and end times of data acquisition for each sample. | A UTC reference is needed for normal operations as well as troubleshooting and debugging. | Needed for debugging the timing of the AO hardware and RTC.
|  |  |  | 10 usec is at the approximate resolution needed for troubleshooting and debugging. | Needed for PSF post-processing (e.g., ScR-489)
Archiving of data | FR-1453 | Essential | The RTC shall archive data to both the the high-speed RAID system and the NGAO Data Server. The data streams to be recorded are listed in FR-1416 and include calibration data, telemetry, and diagnostics. | Need for later post processing of NGAO data for turbulence effects (example, PSF estimation across field of view)
Wind layer computation output | FR-1457 | Important | The RTC shall have the capability to transmit the entire volume estimate of turbulence at the full frame rate to an external computer for wind layer computation. This external computer is designed to estimate the wind at each layer. This requires a high-speed dedicated HW/SW interface specifically for this purpose. Note this is in addition to the requirement (FR-1416) that the full volume of estimates of turbulence be transmitted to a telemetry recording system. | The external wind computation is an option. The wind computation takes the volume estimate of turbulence as an input and produces parametric wind velocity data that can be used by the RTC.
Parameter readback | FR-1458 | Essential | The RTC shall return the actual parameter data in use when a parameter readback is requested by the AO Control System. These data shall be the actual values in use and not merely cached values. Examples of parameter data are lookup tables, system matrices, Cn2 profiles, etc. | Previous experience has shown the need for actual values to support engineering and troubleshooting.
Wind layer computation input | FR-2246 | Important | The RTC shall have the capability to receive parametric wind velocity data (in the form of a Fourier shift factor for each Fourier component in each volume voxel element) from the wind layer computation computer. This is to implement wind prediction / look-ahead in the controller. The parametric input should be updatable once per second. The parameter input shall be via the standard AO control system input to the RTC. | The external wind computation is an option. The wind computation takes the volume estimate of turbulence as an input and produces parametric wind velocity data that can be used by the RTC.
Startup | FR-2254 | Essential | Upon power-up, the RTC shall perform an automated startup sequence, to include a self test and loading a default set of parameters, and report its status to the AO Control System | The RTC must come up in a known usable state.
Status and Health Reporting | FR-2255 | Essential | The RTC shall provide the AO Control System with periodic updates on its status and health. The AO Control system shall initiate this process. | The RTC must report to the overall NGAO system health monitor (part of the NGAO Control System).
|  |  | The RTC shall provide control information to 5 DMs: |  | KAON 511: NGAO System Design Manual
|  |  | The wide-field relay DM (woofer) (1 DM, 20 actuators across beam diameter) |  | KAON 642: NGAO Design Changes in Support of the Build to Cost Guidelines
The narrow-field relay DM (tweeter) (1 DM, 64 actuators across beam diameter)
The LOWFS wavefront correctors (3 DMs, each 32 actuators across beam diameter)