Keck Adaptive Optics Note 461

Wavefront Error Budget Predictions and Measured Performance

for Current and Upgraded Keck Adaptive Optics

P. Wizinowich, R. Dekany, M. van Dam

March 5, 2007

# Introduction

The purpose of this note is multifold:

- To summarize the measured versus predicted performance for the existing Keck AO systems and our understanding of the reasons for the differences.

- To provide an anchor for the Next Generation AO (NGAO) wavefront error budget tool.

- To predict the performance of the planned and potential upgrades to this system.

Note that item 2 is in support of NGAO System Design phase WBS 3.1.1.2.3 Anchor to Keck II LGS AO PSFs and item 3 is in support of WBS 3.1.2.1.2 NGAO versus Keck AO Upgrades trade study.

# Measured versus Predicted Performance for Current Keck AO

  - Measured versus Predicted Comparison

The originally predicted performance, and assumptions, for the Keck AO systems are defined in the Keck AO Blue Book (Keck Observatory Report 208).  The measured performance of the Keck NGS and LGS AO systems are summarized in two refereed publications by van Dam et al.  Table 1 compares the predicted and measured performance from these two sources for the NGS bright star case and the LGS case with a 10th magnitude NGS.  For the LGS case a column is also included for the NGAO tool predictions that will be discussed in section 3.  The final column shows the measured LGS AO performance for an 18th mag NGS.

The various error terms in the rows of this table are explained in the Blue Book and two van Dam et al. papers.  The “Miscellaneous” row was included for the measured NGS case to include terms that could not be explicitly assigned to the other rows such as dynamic calibration errors, additional telescope errors, poorly sensed modes and atmospheric dispersion.  The “Miscellaneous (NGAO)” row was included to represent the sum of seven additional small error sources in the NGAO error budget tool, each of about 20 nm in magnitude.  The percentile seeing should be interpreted using the following example: 75% means that 75% of the nights have worse seeing and the listed performance can be achieved or exceeded on the 25% of the nights with the best seeing. The reason we use percentiles rather than atmospheric parameters is that the LGS AO performance depends on a large number of atmospheric parameters, such as r0, τ0, Cn2 profile and the sodium density structure.

It is clear from the “Total wavefront error” and “K-band Strehl” rows that the measured performance is significantly lower than the Blue Book predicted, or rather planned, performance.  This discrepancy is discussed in the next section.  The Strehl in the measured column is in reality higher than the value calculated by adding the terms in quadrature, since the the high spatial frequency terms are double counted (focal anisoplanatism and atmospheric and fitting error).

One point should be made before we move to this explanation.  That is that the impact of the tip/tilt errors is different than that of the higher order terms.  The tip/tilt errors broaden the diffraction-limited core while the higher order errors move energy from the diffraction-limited core into the seeing limited halo.

Table 1 Keck AO Wavefront Error Budget:

Planned (Blue Book), Measured and Predicted (NGAO tool)

  - Measured versus Predicted Explanation

The discrepancy between the Blue Book values and the measured performance is explained by the delivered system not achieving some of its requirements and the Keck environment being different than the assumptions in the Blue Book.  These factors are further explained as follows:

- The science instrument did not achieve its wavefront quality requirements.  Note that the image quality can be improved by image sharpening, but only over narrow fields and wavelengths. In addition, scattering and vibrations on the AO bench limit the ability to obtain diffraction-limited images with a calibration source.

- The delivered AO system did not achieve its bandwidth requirement.  Both the DM and TT bandwidth terms are significantly higher than the predicted performance due to the fact that the achieved bandwidth (30 Hz for NGS AO, 20 Hz for LGS AO) is significantly less than the expected bandwidth (90 Hz).  The DM bandwidth error discrepancy is due to latency in the real-time control system and our overly optimistic expectations about the ratio of sample rate to bandwidth.  The same latencies impact the NGS tip/tilt bandwidth.  Tip/tilt bandwidth latencies and errors are reduced in the LGS case (the higher numbers in this case are due to a fainter star).  The DM measurement error discrepancy in the LGS case is also partly due to the lower bandwidth since these two terms are traded against each other in operation (the system is run slower to reduce the measurement error at the expense of the bandwidth error). In addition, the tip-tilt mirror exhibits a large resonance when driven at large bandwidths, limiting the ability to correct for high frequency aberrations.

- The delivered laser system did not achieve the planned laser power.  A 20 W system was planned versus the actually achieved ~ 13 W system (only 10 W at the time that the Strehl ratios presented in Table 1 were measured). This results in increased DM bandwidth and measurement errors, as the frame rate is reduced due to the reduced power.

- The Keck telescopes contribute significant tip/tilt vibrations in the 20-40 Hz range that was not accounted for in the Blue Book.  The low bandwidth of the delivered system does not allow for significant correction of these vibrations.

- The value for the focal anisoplanatism parameter, d0, assumed (based on very little data) in the Blue Book appears to have been too optimistic.  More recent measurements of the Cn2 profile above Mauna Kea indicate that this parameter is smaller and the resultant amount of focal anisoplanatism is therefore significantly larger.

# Anchoring the NGAO Wavefront Error Budget Tool

The measured Keck II LGS AO performance is being used as an anchor to test and improve the NGAO wavefront error budget spreadsheet developed by Dekany et al (v1.02, with some modifications).  The comparisons of the NGAO tool predictions to the measured performance in Table 1 show good agreement.

The following assumptions were used in the NGAO tool for the Keck AO case:

- The atmosphere and sodium abundance in the NGAO proposal and a zenith angle of 10º (r0 = 0.178 m and fG = 41.5 Hz at this zenith angle, wind speed = 8 m/s, outer scale = 50 m, Na abundance = 4E9 atoms/cm2).  d0 is calculated from the atmospheric profile used in the NGAO proposal.  However, in order to better simulate the observed focal anisoplanatism the d0 value was multiplied by 0.8 (using the single laser back projection benefit factor).

- A high order wavefront sensor rate of 670 Hz in the NGS case and 500 Hz in the LGS case.  The high order servo decimation factor used was 25.

- A tip/tilt sensor rate of 670 Hz in the NGS case.  For the LGS case, 1000 Hz for the 10th magnitude case and 400 Hz for the 18th mag case.  The tip/tilt decimation factor used was 25 in the NGS case and 40 in the LGS case.

- In order to account for the observed telescope jitter a 0.2 arcsec rms jitter at 29 Hz was assumed.

- A CCD39 was used for the high order wavefront sensor and STRAP for the LGS tip/tilt sensor.  The noise performance of the CCD39 is better than the MIT/LL-64, however this should only have a very minor impact for the bright NGS and the LGS cases.  The assumed charge diffusion was 0.25 pixels.

- The wavefront sensor subaperture size and corresponding actuator spacing was 562.5 mm.

- A higher laser power, 16W, was used to reduce the measurement error.  The Keck laser was likely only projecting between 12 and 14W during the actual measurements.

- A rather large LGS spot size is assumed for centroiding: 2.24 arcsec.

- For the bright NGS case V = 8.0.  For the LGS cases, the NGS V = 10.0 for the bright case and 18.4 for the faint case.  An 18.4 mag star was assumed for the faint LGS case to more closely match the measured errors.  This was done to compensate for the fact that the current NGAO calculation predicts too high an SNR due to a multiplication of the single exposure SNR by the decimation factor in order to get the -3db SNR.

# Predicted Performance for Planned Keck AO Upgrades

  - Next Generation Wavefront Controller (NGWFC)

The NGWFC upgrade will provide the following improvements:

- Reduced read noise and dark current due to the new wavefront sensor camera.  The LL-64 had a read noise of 7 electrons at all rates with a high dark current.  The new CCD39 has a read noise of 4.4 electrons at 500 Hz and negligible dark current.  The new CCD should provide a minor performance improvement for the LGS case and a very significant improvement for the faint NGS case.

  - Since the CCD39 is already in the NGAO tool baseline tool we have not yet tried to determine this effect.

- Increased DM bandwidth.  The previous maximum frame rate of 670 Hz has been increased to > 2 kHz.  This implies that the delay between sampling and correcting has been reduced by a factor of about 2000/670 = 3.  In the bright NGS case we can increase the frame rate from 670 to 2000 Hz.   In the LGS case we are currently limited by laser power so we probably cannot run the camera significantly faster.  Note that the bandwidth error goes as f5/6, where f is the system bandwidth.

  - To simulate this with the NGAO tool the NGAO tool high order servo decimation factor was reduced from 25 to 20 in addition to increasing the frame rate to 2000 Hz.

- Increased TT bandwidth.  The tip/tilt bandwidth is currently limited by the lag time of the real-time system.  However, it can also be limited by the bandwidth of the tip/tilt mirror and/or the sample rate of the STRAP sensor (~ 1000 Hz).  Even with somewhat improved bandwidth there will still be relatively limited improvement due to the inability to correct for the ~ 30 Hz vibrations.

  - To simulate this with the NGAO tool the tip/tilt servo decimation factor was reduced from 40 to 30.

- Improved image quality to the wavefront sensor camera due to the implementation of new lenslet arrays.

  - To simulate this we could arbitrarily reduce the miscellaneous errors by subtracting a 30 nm rms term in quadrature.

The above effects are summarized in Table 2.

Table 2 NGAO Tool Predictions for NGWFC versus Current System

The predicted and actual NGS improvement has been measured to represent a 1.5 magnitude increase in the limiting magnitude.  The LGS performance improvement is not yet tested, but is likely to be modest.  The prediction is for a 10% K-band Strehl improvement.

  - Keck I LGS AO with NGWFC

The Keck I LGS AO system is largely a copy of Keck II’s with the following performance improvement exceptions:

- The new laser is 20 W versus the current 13 W laser.  This provides a 50% increase in sodium return.  In the photon-limited case, that we should be in, the DM measurement error (or SNR) goes as the square root of this power increase.

- The new laser is mode-locked CW versus the current CW laser.  Mode-locked is predicted to have a significant better coupling to the sodium layer.  The measured improvement from a Keck-Gemini experiment was a factor of two (see KAON 419).  Instead of further reducing the measurement error we will assume that this factor of two will allow us to double the sample rate and hence reduce the bandwidth error.

  - The above power and coupling improvements were simulated in the NGAO tool by increasing the power from 16 to 40 W.

- The new laser will be center projected versus side projection for the current system.

  - This will reduce the average LGS elongation by approximately a factor of two.  The observed LGS image size is a function of this elongation as well as other sources such as the seeing (double pass), laser beam quality and diffraction and aberrations from the projection telescope.  The observed laser spot size for the current Keck system ranges from 1” close to the projector to 2.6” on the other side of the telescope in good seeing conditions.  A reasonable range for center projection might therefore by 1” to 1.3”.  The DM measurement error goes linearly with the spot size and therefore would be expected to decrease by a factor of ~ 1.5.

    - The elongation improvement was simulated in the NGAO tool by reducing the laser launch telescope off-axis distance from 7 to 0 m.

  - A symmetric elongation pattern could also reduce the rotation-dependent aberrations seen by the low bandwidth wavefront sensor in the current Keck LGS AO system.  These are represented by the high order LGS aberrations term in the error budget.

    - To simulate this effect in the NGAO tool the dynamic wavefront sensor zero-point calibration error was reduced from 80 to 50 nm (arbitrarily).

Table 3 summarizes the above changes to the wavefront error budget.  The bolded items are the terms that were effected.

Table 3 NGAO Tool Predictions for K1 LGS with NGWFC versus Current System

  - Minor Planned Upgrades

An IR atmospheric dispersion corrector will be installed on Keck II for use with NIRC2.

A low spatial mode of the low bandwidth wavefront sensor is being implemented with 5x5, instead of 20x20, subapertures.  This should allow for a 1-magnitude increase to 19th magnitude in the limiting magnitude of the NGS used for LGS AO observations.

# Predicted Performance for Potential Keck AO Upgrades

We now look at the largest terms in the error budget and what could potentially be done to reduce these terms.

5.1 More Laser Power

LMCT is building a 50W laser for Gemini-S.  A first step, which would be needed to support additional upgrades, would be to upgrade the planned Keck I 20W laser with this 50W capability.  Since the DM measurement error goes as 1/SNR this would allow a 1.6 reduction in this term.  This would be balanced with a reduction in the DM bandwidth error.

- This is simulated in the NGAO tool by increasing the laser power from 40 to 100W.

5.2 More DM Actuators

Another obvious upgrade is to increase the number of actuators.  Xinetics has a DM with 3.5 mm actuator spacing versus the 7 mm spacing on the existing DM.  Doubling the number of actuators would reduce the atmospheric fitting error by a factor of d5/6 = 1.8.  Additional actuators would also reduce the telescope static errors.

- This is simulated in the NGAO tool by halving the subaperture width from 562.5 to 281.3 mm.

- Neyman, Flicker and Panteleev have calculated the telescope static errors for their poster entitled “Keck segment wavefront errors and their impact on AO performance.  Their conclusion was that the AO corrected wavefront error would be 66, 59 and 51 nm, respectively, for 20x20, 32x32 and 48x48 subapertures.  Doubling the number of actuators would therefore decrease the telescope static error from 66 to 55 nm.

Note that this upgrade would require some modest optical changes to the existing bench.  The lenslets would need to be changed.  We would need to upgrade the SciMeasure CCD39 wavefront sensor camera to one with more pixels; presumably the CCID-56.

- Inserting the CCID-56 at this point in the upgrades reduces the LGS measurement error from 61 to 54 nm.  Increasing the high order bandwidth from 1111 to 1250 Hz then balances the bandwidth and measurement errors at 57 and 58 nm, respectively.

The Microgate real-time system is modular enough that it could be expanded to handle the extra subapertures and actuators.  We would likely need a NGS mode that could still work with the current subaperture size; although we would use the larger number of subapertures in the bright star case.

5.3 New Science Camera(s)

A more capable AO system would warrant a new science instrument(s) to take advantage of its capabilities.  It is reasonable to assume that these instruments would be built to meet a more stringent wavefront requirement.  For our purpose we will assume the same 35 nm rms as assumed in the NGAO budget.

- This is included in the NGAO tool by simply reducing the uncorrectable instrument aberrations from 110 to 30 nm.

5.4 Tip/tilt Bandwidth

The tip/tilt bandwidth term is large due to vibrations in the telescope environment.  This could be reduced either by finding and correcting the source of these vibrations, and/or isolating the telescope/AO system from these vibrations and/or active correction.  We have demonstrated the use of a parametric oscillator to notch out up to 3 narrow vibration lines.  Another approach would be to have a separate sensor looking at a bright star.

- For the purpose of the NGAO tool we will assume we can effectively reduce the amplitude of the telescope vibrations from 0.2 arcsec to 0.1 arcsec.

5.5 Simplified Tomography

The tomography error could be reduced if we had some means of measuring the Cn2 profile above the telescope.  Assuming that we can implement some external Cn2 profiler (i.e., a MASS or SLODAR system) then we would need to implement a tomography portion to the real-time controller that used this information to determine the optimal correction to apply to the wavefront sensor centroid offsets and hence to the DM.

- Neyman estimates that this approach could reduce the effective d0 by as much as a factor of two, but on average by a factor of 1.3.  This factor of 1.3 was used in the NGAO tool.

5.6 Predicted Performance Summary

Table 4 includes the performance improvements assuming all of the above upgrades have been successfully implemented.  To summarize the changes that have been made with respect to the current system include the following:

- NGWFC implemented with CCID-56.

- Center laser projection (versus off-axis).

- 50 W of CW mode locked CW laser power (versus 13 W of CW).

- 3.5 mm actuator spacing DM (versus 7 mm; pupil size remains unchanged).

- Lower wavefront error from the science camera (30 nm valid for OSIRIS or a new instrument versus 110 nm for NIRC2).

- A factor of two reduction in the amplitude of tip/tilt vibrations.

- Implementation of a Cn2 profiler and a tomography capability in the real-time controller.

Table 4 Predictions for an Upgraded Keck AO System versus Current System

# Comparison of Upgraded Keck AO to NGAO

NGAO would only have a very modest wavefront error budget reduction for the bright NGS case due to the presence of more DM actuators for the reduction of telescope static wavefront errors.

The primary advantage of NGAO over an upgraded Keck AO system for the bright LGS case is the further significant reduction of the tomography error.  The DM bandwidth and measurement errors would also be reduced due to more laser power.

In the faint LGS case one or more IR tip/tilt sensors would also offer a significant advantage.  An IR tip/tilt sensor could also be implemented in an upgraded Keck AO system.  However, due to the lower Strehl, in the upgraded Keck system versus NGAO, the faint star performance would not be increased by as much.

The Keck upgrade (without an IR tip/tilt sensor and the NGAO case are compared in Table 5.

Table 5 Performance comparison for a Keck AO upgrade versus NGAO.

# Appendices

# Three wavefront error budget appendices are attached; one each for the NGS bright star, LGS bright star and LGS faint star cases.  Each column represents a case the NGAO tool was used to evaluate.  Each column builds on the improvements in the previous improvements; with the exception of the last column which was used to evaluate the NGAO system.  So, the 2nd last column is the Keck AO upgrade including adding the NGWFC to the current system, implementing the K1 LGS upgrade, replacing the CCD39 with a CCID56, implementing a new DM with ½ the actuator spacing, implementing a higher wavefront quality science instrument, using an atmospheric profiler to do simplified tomography, reducing the vibrations at 29 Hz by a factor of two and upgrading from a 20W to a 50W laser.

# References

Adaptive Optics for Keck Observatory.  Keck Observatory Report 208 (January, 1996).

Dekany et al., NGAO Error Budget Excel Spreadsheet v1.02 (February, 2007).

van Dam, M.A., Le Mignant, D., & Macintosh, B.A., Performance of the Keck Observatory adaptive-optics system.  Applied Optics 43: 5459-5467 (2004).

van Dam, M.A. et al., The W.M. Keck Observatory Laser Guide Star Adaptive Optics System: Performance Characterization.  Publications of the Astronomical Society of the Pacific 118: 310-318 (2006).

Appendix 1.  NGS (8th mag) Wavefront Error Budget versus Improvements

| NGS (8th mag) Case | Current | NGWFC | CCID56 | 2x DM | Sci. Inst. | Vibes | NGAO |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Atmospheric Fitting Error | 110 | 110 | 110 | 62 | 62 | 62 | 62 |
| Bandwidth Error | 115 | 38 | 38 | 38 | 38 | 38 | 30 |
| High-order Measurement Error | 16 | 29 | 28 | 38 | 38 | 38 | 43 |
| LGS Focal Anisoplanatism Error | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Asterism Deformation Error | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Multispectral Error | 22 | 22 | 22 | 22 | 22 | 22 | 22 |
| Scintillation Error | 14 | 14 | 14 | 14 | 14 | 14 | 14 |
| WFS Scintillation Error | 10 | 10 | 10 | 10 | 10 | 10 | 10 |
| Uncorrectable Static Telescope Abs | 66 | 66 | 66 | 55 | 55 | 55 | 43 |
| Uncorrectable Dynamic Telescope Abs | 7 | 2 | 2 | 2 | 2 | 2 | 1 |
| Static WFS Zero-point Calibration Error | 25 | 25 | 25 | 25 | 25 | 25 | 25 |
| Dynamic WFS Zero-point Calib Error | 80 | 80 | 80 | 80 | 80 | 80 | 80 |
| Go-to Control Errors | 0 | 0 | 0 | 0 | 0 | 0 | 43 |
| Residual Na Layer Focus Change | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| DM Finite Stroke Error | 16 | 16 | 16 | 16 | 16 | 16 | 12 |
| DM Hysteresis | 13 | 13 | 13 | 13 | 13 | 13 | 13 |
| High-Order Aliasing Error | 37 | 37 | 37 | 21 | 21 | 21 | 14 |
| DM Drive Digitization | 1 | 1 | 1 | 1 | 1 | 1 | 1 |
| Uncorrectable AO System Aberrations | 30 | 30 | 30 | 30 | 30 | 30 | 30 |
| Uncorrectable Instrument Aberrations | 110 | 110 | 110 | 110 | 30 | 30 | 30 |
| DM-to-lenslet Misregistration | 15 | 15 | 15 | 15 | 15 | 15 | 15 |
| DM-to-lenslet Pupil Scale Error | 15 | 15 | 15 | 15 | 15 | 15 | 15 |
| Angular Anisoplanatism Error | 24 | 24 | 24 | 24 | 24 | 24 | 24 |
| Tilt Measurement Error (one-axis) | 5 | 12 | 9 | 12 | 12 | 12 | 17 |
| Tilt Bandwidth Error (one-axis) | 17 | 6 | 6 | 6 | 6 | 6 | 3 |
| Tilt Anisoplanatism Error (one-axis) | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Residual Centroid Anisoplanatism | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Residual Atmospheric Dispersion | 26 | 26 | 26 | 26 | 26 | 26 | 1 |
| Science Instrument Mechanical Drift | 6 | 6 | 6 | 6 | 6 | 6 | 6 |
| Long Exposure Field Rotation Errors | 6 | 6 | 6 | 6 | 6 | 6 | 6 |
| Residual Telescope Pointing Jitter (1-axis) | 89 | 30 | 30 | 30 | 30 | 15 | 9 |
| Total Wavefront Error (nm) = | 250 | 210 | 210 | 185 | 152 | 149 | 148 |
| Strehl at K-band = | 0.57 | 0.69 | 0.69 | 0.75 | 0.82 | 0.83 | 0.84 |
| Zenith angle (deg) | 10 |  |  |  |  |  |  |
| Guide star magnitude | 8 |  |  |  |  |  |  |
| HO WFS Rate (Hz) | 670 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 |
| TT Rate (Hz) | 670 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 |
| WFS camera | CCD39 | CCD39 | CCID56 | CCID56 | CCID56 | CCID56 | CCID56 |
| HO decimation factor | 25 | 20 | 20 | 20 | 20 | 20 | 15 |
| TT decimation factor | 25 | 25 | 30 | 30 | 30 | 30 | 15 |
| Subaperture diameter (m) | 0.5625 | 0.5625 | 0.5625 | 0.28125 | 0.28125 | 0.28125 | 0.28125 |
| Science Instrument | NIRC2 | NIRC2 | NIRC2 | NIRC2 | TBD | TBD | TBD |
| Amplitude of vibrations (arcsec) | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.1 | 0.1 |

Appendix 2.  LGS (10th mag) Wavefront Error Budget versus Improvements

| LGS (10th mag) Case | Current | NGWFC | K1 LGS | CCID56 | 2x DM | Sci. Inst. | Tomog | Vibes | 50W las | NGAO |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Atmospheric Fitting Error | 110 | 110 | 110 | 110 | 62 | 62 | 62 | 62 | 62 | 62 |
| Bandwidth Error | 146 | 121 | 68 | 68 | 68 | 68 | 68 | 68 | 55 | 30 |
| High-order Measurement Error | 150 | 150 | 76 | 71 | 83 | 83 | 83 | 83 | 59 | 32 |
| LGS Focal Anisoplanatism Error | 208 | 208 | 208 | 208 | 208 | 208 | 167 | 167 | 167 | 76 |
| Asterism Deformation Error | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 29 |
| Multispectral Error | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 |
| Scintillation Error | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 |
| WFS Scintillation Error | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 |
| Uncorrectable Static Telescope Abs | 66 | 66 | 66 | 66 | 55 | 55 | 55 | 55 | 55 | 43 |
| Uncorrectable Dynamic Tel Abs | 9 | 7 | 4 | 4 | 4 | 4 | 4 | 4 | 3 | 1 |
| Static WFS Zero-point Calib Error | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 |
| Dyn WFS Zero-point Calib Error | 80 | 80 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 |
| Go-to Control Errors | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 43 |
| Residual Na Layer Focus Change | 36 | 27 | 27 | 27 | 27 | 27 | 27 | 27 | 34 | 7 |
| DM Finite Stroke Error | 16 | 16 | 16 | 16 | 16 | 16 | 16 | 16 | 16 | 12 |
| DM Hysteresis | 13 | 13 | 13 | 13 | 13 | 13 | 13 | 13 | 13 | 13 |
| High-Order Aliasing Error | 37 | 37 | 37 | 37 | 21 | 21 | 21 | 21 | 21 | 21 |
| DM Drive Digitization | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |
| Uncorrectable AO System Aberrations | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 |
| Uncorrectable Instrum Aberrations | 110 | 110 | 110 | 110 | 110 | 30 | 30 | 30 | 30 | 30 |
| DM-to-lenslet Misregistration | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 |
| DM-to-lenslet Pupil Scale Error | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 |
| Angular Anisoplanatism Error | 24 | 24 | 24 | 24 | 24 | 24 | 24 | 24 | 24 | 24 |
| Tilt Measurement Error (one-axis) | 11 | 13 | 13 | 13 | 13 | 13 | 13 | 13 | 13 | 0 |
| Tilt Bandwidth Error (one-axis) | 18 | 13 | 13 | 13 | 13 | 13 | 13 | 13 | 13 | 3 |
| Tilt Anisoplanatism Error (one-axis) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Residual Centroid Anisoplanatism | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11 |
| Residual Atmospheric Dispersion | 26 | 26 | 26 | 26 | 26 | 26 | 26 | 26 | 26 | 1 |
| Science Instrument Mechanical Drift | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 |
| Long Exposure Field Rotation Errors | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 |
| Residual Tel Pointing Jitter (1-axis) | 96 | 72 | 72 | 72 | 72 | 72 | 72 | 36 | 36 | 9 |
| Total Wavefront Error (nm) = | 372 | 357 | 311 | 310 | 296 | 276 | 247 | 239 | 229 | 155 |
| Strehl at K-band = | 0.30 | 0.34 | 0.43 | 0.44 | 0.49 | 0.51 | 0.61 | 0.63 | 0.64 | 0.82 |
| Zenith angle (deg) | 10 |  |  |  |  |  |  |  |  |  |
| Guide star magnitude | 10 |  |  |  |  |  |  |  |  |  |
| HO WFS Rate (Hz) | 500 | 500 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1300 | 2000 |
| TT Rate (Hz) | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 2000 |
| Laser | LP | LP | LP | LP | LP | LP | LP | LP | LP | CW |
| Laser power (W) | 16 | 16 | 40 | 40 | 40 | 40 | 40 | 40 | 100 | 150 |
| Off-axis projection distance (m) | 7 | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| WFS camera | CCD39 | CCD39 | CCD39 | CCID56 | CCID56 | CCID56 | CCID56 | CCID56 | CCID56 | CCID56 |
| HO decimation factor | 25 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 15 |
| TT decimation factor | 40 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 15 |
| Subaperture diameter (m) | 0.5625 | 0.5625 | 0.5625 | 0.5625 | 0.28125 | 0.28125 | 0.28125 | 0.28125 | 0.28125 | 0.28125 |
| Science Instrument | NIRC2 | NIRC2 | NIRC2 | NIRC2 | NIRC2 | TBD | TBD | TBD | TBD | TBD |
| Amplitude of vibrations (arcsec) | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.1 | 0.1 | 0.1 |
| Single laser backprojection benefit | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 |
| d0 (m) | 2.41 | 2.41 | 2.41 | 2.41 | 2.41 | 2.41 | 3.14 | 3.14 | 3.14 | 2.41 |
| TT sensor | STRAP | STRAP | STRAP | STRAP | STRAP | STRAP | STRAP | STRAP | STRAP | SNAP |

Appendix 3.  LGS (18th mag) Wavefront Error Budget versus Improvements

| LGS (18.4 mag) Case | Current | NGWFC | K1 LGS | CCID56 | 2x DM | Sci. Inst. | Tomog | Vibes | 50W las | NGAO |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Atmospheric Fitting Error | 110 | 110 | 110 | 110 | 62 | 62 | 62 | 62 | 62 | 62 |
| Bandwidth Error | 146 | 121 | 68 | 68 | 68 | 68 | 68 | 68 | 55 | 30 |
| High-order Measurement Error | 150 | 150 | 76 | 71 | 83 | 83 | 83 | 83 | 59 | 32 |
| LGS Focal Anisoplanatism Error | 208 | 208 | 208 | 208 | 208 | 208 | 167 | 167 | 167 | 76 |
| Asterism Deformation Error | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 29 |
| Multispectral Error | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 |
| Scintillation Error | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 | 14 |
| WFS Scintillation Error | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 |
| Uncorrectable Static Tel Abs | 66 | 66 | 66 | 66 | 55 | 55 | 55 | 55 | 55 | 43 |
| Uncorrectable Dynamic Tel Abs | 9 | 7 | 4 | 4 | 4 | 4 | 4 | 4 | 3 | 1 |
| Static WFS Zero-point Calibration Error | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 |
| Dynamic WFS Zero-point Calib Error | 80 | 80 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 |
| Go-to Control Errors | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 43 |
| Residual Na Layer Focus Change | 91 | 91 | 91 | 91 | 91 | 91 | 91 | 91 | 91 | 14 |
| DM Finite Stroke Error | 16 | 16 | 16 | 16 | 16 | 16 | 16 | 16 | 16 | 12 |
| DM Hysteresis | 13 | 13 | 13 | 13 | 13 | 13 | 13 | 13 | 13 | 13 |
| High-Order Aliasing Error | 37 | 37 | 37 | 37 | 21 | 21 | 21 | 21 | 21 | 21 |
| DM Drive Digitization | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |
| Uncorrectable AO System Abs | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 |
| Uncorrectable Instrument Abs | 110 | 110 | 110 | 110 | 110 | 30 | 30 | 30 | 30 | 30 |
| DM-to-lenslet Misregistration | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 |
| DM-to-lenslet Pupil Scale Error | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 |
| Angular Anisoplanatism Error | 24 | 24 | 24 | 24 | 24 | 24 | 24 | 24 | 24 | 24 |
| Tilt Measurement Error (one-axis) | 349 | 349 | 349 | 349 | 349 | 349 | 349 | 285 | 285 | 24 |
| Tilt Bandwidth Error (one-axis) | 45 | 44 | 44 | 44 | 44 | 44 | 44 | 67 | 67 | 7 |
| Tilt Anisoplanatism Error (one-axis) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Residual Centroid Anisoplanatism | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11 | 11 |
| Residual Atmospheric Dispersion | 26 | 26 | 26 | 26 | 26 | 26 | 26 | 26 | 26 | 1 |
| Science Instrument Mechanical Drift | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 |
| Long Exposure Field Rotation Errors | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 |
| Residual Tel Pointing Jitter (1-axis) | 239 | 239 | 239 | 239 | 239 | 239 | 239 | 179 | 179 | 18 |
| Total Wavefront Error (nm) = | 563 | 557 | 529 | 528 | 520 | 509 | 494 | 425 | 419 | 158 |
| Strehl at K-band = | 0.08 | 0.09 | 0.11 | 0.11 | 0.12 | 0.13 | 0.15 | 0.21 | 0.21 | 0.81 |
| Zenith angle (deg) | 10 |  |  |  |  |  |  |  |  |  |
| Guide star magnitude | 18.4 |  |  |  |  |  |  |  |  |  |
| HO WFS Rate (Hz) | 500 | 500 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1300 | 2000 |
| TT Rate (Hz) | 400 | 300 | 300 | 300 | 300 | 300 | 300 | 200 | 200 | 1000 |
| Laser | LP | LP | LP | LP | LP | LP | LP | LP | LP | CW |
| Laser power (W) | 16 | 16 | 40 | 40 | 40 | 40 | 40 | 40 | 100 | 150 |
| Off-axis projection distance (m) | 7 | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| WFS camera | CCD39 | CCD39 | CCD39 | CCID56 | CCID56 | CCID56 | CCID56 | CCID56 | CCID56 | CCID56 |
| HO decimation factor | 25 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 15 |
| TT decimation factor | 40 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 15 |
| Subaperture diameter (m) | 0.5625 | 0.5625 | 0.5625 | 0.5625 | 0.28125 | 0.28125 | 0.28125 | 0.28125 | 0.28125 | 0.28125 |
| Science Instrument | NIRC2 | NIRC2 | NIRC2 | NIRC2 | NIRC2 | TBD | TBD | TBD | TBD | TBD |
| Amplitude of vibrations (arcsec) | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.1 | 0.1 | 0.1 |
| Single laser backprojection benefit | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 |
| d0 (m) | 2.41 | 2.41 | 2.41 | 2.41 | 2.41 | 2.41 | 3.14 | 3.14 | 3.14 | 2.41 |
| TT sensor | STRAP | STRAP | STRAP | STRAP | STRAP | STRAP | STRAP | STRAP | STRAP | SNAP |