Palomar Adaptive Optics Test Plan

| Title | Low-order servo loop optimization |
| --- | --- |
| Version | 1.1 |
| Date released | 4/3/2007 |
| Lead | A. Bouchez |
| Time requested | 2 hours |
| Required conditions | Photometric, >90deg. from moon. |

Purpose - Determine optimal low-order servo loop parameters as a function of NGS magnitude. - Practice acquisition procedure with faint NGS.

Previous analysis Starting estimates (please note optimized values): NGS V magnitude; Acq integ. (s); Framerate (Hz); Integral gain. 10; 0.03; 700; 1.0. 12; 0.5; 400; 1.0. 14; 2.0; 250; 1.0. 15; 10.0; 200; 1.0. 16; 20.0; 150; 1.0. 17; 30.0; 100; 1.0. 18; 30.0; 50; 1.0.

Test procedure Targets for this test: isolated V=14, 15, 16, 17 stars near zenith. Use red stars (B-V = 1.0-2.5) to maximize PHARO sensitivity. - Follow LGS Target Acquisition testplan for star. - Only perform NGS tune-up on the first target. - Use LOWFS framerates from the table above. - DO perform "zero LOWFS centroids" step! - Record acquisition camera image before offsetting LOWFS. - Determine PHARO integration time for decent SNR images. Goal is 10-15k peak counts, but keep integration time <30s. - Take a PHARO sky - Open TT loop from TAO command line (TAO button? / PHARO button??) - Offset telescope from PHARO buttons (or TAO) - Take a PHARO sky. - Offset back to target. - Close TT loop (same way as opened). - Record average Strehl of 3 frames in table below. - Adjust framerate - Open TT loop. - Set framerate to one "level" higher - Take new LOWFS background (offset 60" to sky). - Close TT loop. - Record average Strehl of 3 frames. - Repeat for one level lower. - Set to optimal framerate and take a new sky. Record value in table above. - Adjust integral gain: - set ttm_integral_gain to 0.25, 0.5, 1.0, 1.5 - Record average Strehl of 3 frames. - If loop goes unstable, open TT loop and manually set TTM to (0,0). - Record optimal gain in table above. - Move on to next fainter star.

Results and conclusions V mag.; Star name; framerate; gain; PHARO frames; Average Strehl. 10.0. Note PHARO filter, integration, and sky frames: