## Sheet1
Requirement | Compliance |  | Verify | Notes
6 Overall Requirements
6.2 Performance Requirements
6.2.1 Parametric Performance Requirements
6.2.1.1.1 Transportation and Shipping Environment
Altitude |  | Design | I
Temperature |  | Design | I
Temperature shock |  | Design | I
Humidity |  | Design | I | IRMS will ship in vacuum bag with dessicant
Gravity orientation |  | Design | I | IRMS will not experience a variable gravity vector during shipping
Vibration |  | Design | I | IRMS is mounted on a shock isolating pallet
Shock |  | Design | I | IRMS is mounted on a shock isolating pallet
Acceleration
Due to transport |  | Design | I | IRMS is mounted on a shock isolating pallet
Due to seismic activity |  | Design | I | IRMS is mounted on a shock isolating pallet
6.2.1.1.2 Non-Operating Environment
Altitude |  | Design | I
Temperature
Range |  | Design | I
Rate of change |  | Design | I
Humidity |  | Design | I
Gravity orientation |  | Design | I
Vibration |  | Design | I
Shock |  | Design | I | Note that this parameter will not be tested
Acceleration
Due to handling |  | Design | I
Due to seismic activity |  | Design | I
6.2.1.1.3 Operating Environment
Altitude |  | Design | I
Temperature
Range |  | Design | T | CCR heads not compliant with the minimum operating temperature of -15 ° C based on manufacturer's documentation, however these are in use at WMKO on several instruments and function properly
Rate of change |  | Design | T | Applies to external components only
Humidity |  | Design | I
Gravity orientation |  | Design | I
Vibration |  | Design | I
Shock |  | Design | I
Acceleration |  | Design | I
6.2.2 Operational Performance Requirements
6.2.2.1 Air Borne Contaminants |  | Design | I
6.2.2.2 Audible Noise |  | Design | T
6.2.2.3 Power Failure Tolerance |  | Design | T
6.3 Implementation Requirements
6.3.2 Standards Implementation Requirements
6.3.2.1 Shipping Containers |  | Design, Fab | I
6.3.3 Regulatory Implementation Requirements
CFR 29 Part 1910
Subpart O §1910.212 |  | Design, Fab | I, D | IRMS is designed for operation and maintenance only by properly qualified and trained personnel
Subpart S §1910.302 through 1910.304 |  | Design, Fab | I, D
6.4 Design Requirements
6.4.1 Technological Design Requirements
6.4.1.1 Materials Suitability and Safety |  | Design | I
7 Optical Requirements
7.2 Performance Requirements
7.2.1 Parametric Performance Requirements
7.2.1.1 Typical Parameters
FOV
Imaging |  | Design, Config | I,T
Multi-object spectroscopy |  | Design, Config | I,T
Wavelength coverage
Y-band |  | Design | T
J‑band |  | Design | T
H‑band |  | Design | T
K‑band |  | Design | T
Imaging plate scale |  | Design | T
Spectral resolution |  | Design | T
Multi-object spectroscopy
0.7” slit width |  | Design | T
0.36” slit width |  | Design | T
Input focal ratio |  | Design | I
Science detector |  | Design | I
Guider
FOV |  | Design, Config | T | Requires on sky testing
Sensitivity |  | Design, Fab | T
7.2.1.2 Goal Parameters
Image quality
Imaging |  | Design, Fab | T
Spectroscopy |  | Design, Fab | T
Guider |  | Design, Fab | T | Requires on sky testing
Non-uniformity
Imaging |  | Design, Fab | T
Spectroscopy |  | Design, Fab | T
Guider |  | Design, Fab | T | Requires on sky testing
Distortion
Guider |  | Design, Fab | T | Requires on sky testing
Optical throughput
Imaging
Y‑band |  | Design, Fab | T
J‑band |  | Design, Fab | T
H‑band |  | Design, Fab | T
K‑band |  | Design, Fab | T
Spectroscopy
Y‑band |  | Design, Fab | T
J‑band |  | Design, Fab | T
H‑band |  | Design, Fab | T
K‑band |  | Design, Fab | T
Guider |  | Design, Fab | T
Instrument background
Y‑band |  | Design, Fab | T
J‑band |  | Design, Fab | T
H‑band |  | Design, Fab | T
K‑band |  | Design, Fab | T
Slit mask light blocking |  | Design, Fab | T
Ghosting
Imaging |  | Design, Fab | T
Spectroscopy |  | Design, Fab | T
7.2.2 Operational Performance Requirements
7.2.2.1 Observing modes
Direct Imaging |  | Design | D,T
Multi-object spectroscopy |  | Design | D,T
Multiplex up to 46 objects |  | Design | D,T
Slits deployable on nominal 2.7" pitch |  | Design | I
Minimum slit length 2.4" |  | Design | I,T
Adjacent slit bars form longer slits in increments of 2.7" |  | Design | D,T
Minimum slit width is 0.17", adjustable in increments of 0.03" |  | Design | D,T
7.2.2.2 Filters
10 filters |  | Design | I
7.2.2.3 Lyot stop
Deployable rotating hexagonal cold Lyot stop matched to the telescope pupil |  | Design | T
7.2.2.4 Spectrometer Grating
Reflection grating interchangeable with a plane mirror for imaging |  | Design | I
7.2.2.5 Spectral Coverage
Maximize spectral coverage in each wavelength band in a single exposure for slits located over at least the central 1' of the spectrometer FOV in the spectral dispersion direction |  | Design | T
7.3 Implementation Requirements
7.3.1 Feature Implementation Requirements
7.3.1.1 Dewar Window
Condensation control |  | Design | T
Temperature monitor |  | Design | T
No pressure deformations that affect the optical performance of the instrument |  | Design | I,T
7.3.1.2 Science Detector Focus Control
Remote fine focus |  | Design | T
7.3.1.3 Guider
Full supplied instrument FOV is accessible to guider pick-off |  | Design | I,T
Pick-off shadow is minimized |  | Design | I, D
7.4 Design Requirements
7.4.1 Technological Design Requirements
7.4.1.1 Optical Component Mountings
Alignment is maintained during cool down and warm up cycles |  | Design | T
Excessive stress is not placed on the optical components |  | Design | T
Alignment of optical components without excessive stress is maintained at all rotator angles and telescope elevations |  | Design | T
Materials compatible with the coatings |  | Design | T
Material used within the dewar compatible with vacuum and cryogenic environments |  | Design | T
7.4.1.2 Alignment Tolerancing
Documented optical alignment tolerance budget |  | Design, Fab | I
7.4.3 Standards Related Design Requirements
Drawings conform to ANSI /ASME standard Y14.18M-1986 |  | Design | I
7.4.4 Integration Related Design Requirements
7.4.4.1 Focal Position
Compatible with the f/15 NFIRAOS focus of the TMT |  | Design | I
8 Mechanical Requirements
8.2 Performance Requirements
8.2.1 Parametric Performance Requirements
8.2.1.1 General
Weight |  | Design | I
Overall dimensions |  | Design | I
Temperature
Operating Ambient |  | Design | T | Must be installed at observatory to test
Internal – opto‑mechanics
Operating Temperature |  | Design | T
Stability |  | Design | T
Maximum Gradient
Operating |  | Design | T
Transient (warm up or cool down) |  | Design | T
Internal – science detector
Operating Temperature |  | Design | T
Stability |  | Design | T
Science detector temperature variation (adequate zero point stability) |  | Design, Fab | T
Guider Enclosure Internal
Operating Temperature |  | Design | T
Stability |  | Design | T
Cool down time |  | Design | T
Warm up time |  | Design | T
Vacuum
Hold time |  | Design, Fab | T
Pressure |  | Design, Fab | T
Power Dissipation
To ambient |  | Design | T
To glycol supply |  | Design | T
Glycol cooling
Temperature rise |  | Design | T
Operating pressure |  | Design | T
Pressure drop |  | Design | T
Flow rate |  | Design | T
Flexure |  | Design | T
Alignment
X, Y and Z axis |  | Design, Fab | T
Rotation about X and Y axis |  | Design, Fab | T
8.2.1.2 IRMS Instrument Mechanisms
Dewar Internal Mechanisms
Configurable Slit Unit
Total slits (pairs of slit masking bars) |  | Design | I
Slit position accuracy |  | Design, Fab | T
Slit position repeatability |  | Design, Fab | T
Slit width error |  | Design, Fab | T
Slit width |  | Design, Fab | T
Slit width resolution |  | Design, Fab | T
Reconfiguration Time |  | Design | T
Filter wheels
In beam position repeatability |  | Design, Fab | T
In Beam position repeatability |  | Design, Fab | T
Cycle time |  | Design | T
Rotating Lyot stop
Rotation speed
Slew |  | Design | T
Tracking |  | Design | T
Rotation Range |  | Design | T
Tracking error |  | Design, Fab | T
Position repeatability |  | Design, Fab | T
Deploy/retract time |  | Design | T
Grating/Mirror exchange turret
Position accuracy |  | Design, Fab | T
In beam position repeatability |  | Design, Fab | T
Cycle time |  | Design | T
Flexure correction (if implemented)
Adjustment range |  | Design | T
Position accuracy |  | Design, Fab | T
Position repeatability |  | Design | T
Non-linearity |  | Design | T
Science detector focus
Adjustment range |  | Design | T
Position repeatability |  | Design, Fab | T
Position resolution |  | Design | T
Non-linearity |  | Design, Fab | T
Tip/tilt |  | Design, Fab | T
External Mechanisms
Guider Pick-off Positioning Mechanisms
Radial Range |  | Design | T
Angular Range |  | Design | T
Position accuracy |  | Design, Fab | T
Position repeatability |  | Design, Fab | T
Reconfiguration Time |  | Design | T
Guider Filter Wheel
In beam position accuracy |  | Design, Fab | T
In beam position repeatability |  | Design, Fab | T
Cycle time |  | Design | T
8.2.1.3 Rotator
Rotation speed
Slew |  | Design | T
Tracking |  | Design | T
Rotation Range |  | Design | T
Tracking error |  | Design, Fab | T
Skew |  | Design, Fab | T
8.2.1.4 Vacuum Integrity
>25 weeks under continuous cooling without pumping |  | Design, Fab | T
8.2.1.5 Power Dissipation
≤100 watts of heat into the telescope dome ambient environment |  | Design | I,T
8.2.1.6 Guider Thermal Interface and Enclosure
Thermal seal eliminates infiltration and thermal plumes into NFIRAOS |  | Design, Fab | T
Enclosure seals allow relative rotation beteen IRMS and NFIRAOS |  | Design, Fab | T
Enclosure maintains operating temperatures of contents |  | Design, Fab | T
8.2.2 Operational Performance Requirements
8.2.2.1 Operating Temperature Range |  | Design | I,T
8.2.2.2 Configurable Slit Unit
CSU meets performance requirements at all rotator positions |  | Design, Fab | T
8.2.2.3 Flexure Correction
Piezoelectric actuators operated by a closed loop servo |  | Design | I,D
8.2.2.4 Vibration (TBD)
Compatible with Generic Vibration Criteria Curve "C" |  | Design | T
Instrument does not produce vibrations in excess of those given in curve "C" |  | Design | T
8.3 Implementation Requirements
8.3.1 Feature Implementation Requirements
8.3.1.1 IRMS Instrument Dewar
Maximize the reliability of the mechanisms |  | Design | I,T
Provide status and performance information |  | Design | D
Provide a positive indication that the requested move(s) have been completed |  | Design | D
Operate properly with reduced speed over the ambient temperature range given in §6.2.1.1.3 |  | Design | D
8.3.1.2 Rotator
Manual mechanical lockout feature |  | Design | T
8.3.1.3 Instrument Support Assembly
Seismic standards for a zone 4 earthquake zone |  | Design | I
8.3.1.4 Access and Covers |  | Design | I
8.3.1.5 Glycol Cooling
Braided stainless steel hose and stainless steel fittings |  | Design | I
Custom manifolds |  | Design | I
Permanent connections made with JIC 37º flare compression fittings or SAE straight thread O-ring fittings |  | Design, Fab | I
No teflon tape |  | Fab | I
Removable connections ½” Parker Hannifin series FS |  | Design, Fab | I
Instrument supply coupler is male and the return coupler is female |  | Design | I
Flow meters |  | Design | D
Flow switch |  | Design | D
8.3.1.6 Vacuum Systems
No backstreaming or oil free pump |  | Design | I
Fittings |  | Design | I
8.3.1.6.1 Pressure Control
Low vacuum gauge |  | Design | I,T
High vacuum gauge |  | Design | I,T
Back-up high vacuum gauge |  | Design | I,T
Pressure relief valve(s) |  | Design | I
8.3.1.6.2 Gettering
Passive gettering |  | Design | I
Electro-polished stainless steel meshfor the sieve holder |  | Design | I
8.3.1.7 Cryogenic Systems
CCR heads vibration isolated from the instrument dewar |  | Design | T
Manifolds for the distribution of helium to the CCR heads |  | Design | I
8.3.2 Common Practices Implementation Requirements
8.3.2.1 Fit and Finish
Plated or painted |  | Fab | I
Welds ground to the surface/joint profile or dress quality |  | Fab | I
All welds and castings stress relieved |  | Fab | I
Machined componentsfree of tool marks, scratches and material flaws |  | Fab | I
All external enclosure and exposed structural elements finished in (TBD) epoxy paint |  | Fab | I
All burrs and sharp edges removed from all fabricated components unless the function of the component requires a sharp edge |  | Fab | I
Mild steel surfaces that cannot be painted for functional reasons protected by a non-tracking anti-corrosion dry film lubricant |  | Fab | I
8.3.2.2 Continuity of Shielding and Grounding
Dissimilar metals isolated where electrolytic corrosion may occur |  | Fab | I
All components of enclosures required to provide protective grounding or EMI shielding electrically bonded at multiple points |  | Design | I
Grounding straps tin plated copper braids ≥6 mm wide |  | Fab | I
Anodized aluminum parts free of anodizing where electrical contact is required |  | Fab | I
Painted metal partsfree of paint where electrical contact is required |  | Fab | I
8.3.2.3 Corrosion resistance
Metal components finished to prevent corrosion |  | Fab | I
All removable fastenersplated or treated to prevent corrosion |  | Fab | I
Painting, plating, or anodizing to MIL-STD-171E |  | Fab | I
8.3.2.4 Fasteners
Press fit studs or threaded inserts installed in the correct material |  | Fab | I
Self-tapping screws not used for removable covers or components |  | Fab | I
Phillips or hex socket heads |  | Fab | I
Kel-F or Vespel insert locknuts |  | Fab | I
8.3.2.5 Lubricants
Lubricants suited for low temperature environment |  | Design, Fab | I
8.3.2.6 Lubricated Components
Exposed lubricated components enclosed in a shroud or boot |  | Design | I
8.3.3 Standards Implementation Requirements
8.3.3.1 Structural
Structure meets zone 4 earthquake survival requirements, Telcordia Standard GR-63-CORE |  | Design | I
8.3.3.2 Vacuum Systems
ASME HPS‑2003 |  | Design | I
NFPA 99C, 2005 edition |  | Design | I
8.3.3.3 Cryogenic Systems
NFPA 55, 2005 edition |  | Design | I
8.4 Design Requirements
8.4.1 Technological Design Requirements
8.4.1.1 Vacuum and Cryogenic Components
Materials TML ≤ 1% |  | Design | I
Materials CVCM ≤ 0.1% |  | Design | I
Values for TML and CVCM determined in accord with ASTM E595-93 (2003)e1 |  | Design | I
LN2 fill overflow shield |  | Design | I
8.4.1.2 Opto-Mechanical Assemblies
Removable assemblies provided with locating pins or other features to permit repeatable removal and replacement |  | Design | I
Handling features provided on all components unless inherently easy to handle |  | Design | I
Handles for components with weights greater than 1 kg up to 25 kg |  | Design | I
Heavier components and subassemblies provided with lifting eyes or ‘A’ brackets |  | Design | I
8.4.1.3 Electrical/Electronic Assemblies and Enclosures
Service access |  | Design | D
Terminal access |  | Design | D
Meet wire bend radii and electrical spacing requirement |  | Design | I
Rack mounting conforms to 19 inch (482.6 mm)pattern of EIA 310-D |  | Design | I
Rack mounted modules installed using rack slides |  | Design, Fab | I
Front access only requires all rack slides extend far enough to permit disconnection of rear panel connections prior to removal of the rack module from the slides |  | Design, Fab | D
Systems predominantly of rack mounted modules:
All COTS modules, components and subsystems not available in rack mount configurations mounted in rack module chassis or on rack mount shelves |  | Design, Fab | I
All rack module chassis and shelves mounted on slides |  | Design, Fab | I
Components or modules mounted on shelves fully enclosed as required to meet all requirements for grounding, shielding and electrical safety |  | Design | I
Components or modules weighing < 0.5 kg may be mounted on hinged or screw mounted rack panels, provided that all other requirements for grounding, shielding and electrical safety are met |  | Design | I
Rails have tapped or equipped with captive tapped inserts |  | Design | I
No clip nuts |  | Design | I
Continuous shield |  | Design | I
No openings > 3 mm in diameter or 3 mm in width and 15 cm in length |  | Design | I
Removable covers that do not make continuous contact with the enclosure provided with a fastener every 15 cm or with conductive gaskets or finger stock |  | Design | I
8.4.1.4 Mechanisms
As few identical mechanical assemblies as possible |  | Design | I
Modular assemblies with a minimum of parts |  | Design | I
Provisions for simple installation and removal |  | Design | D
8.4.1.5 Drive Couplings
Shaft couplings pinned or locked |  | Design | I
Separable couplings to facilitate motor replacement |  | Design | I
8.4.1.6 Component Ratings
Structural elements and fasteners whose failure could cause injury to personnel or equipment must be selected for a safety factor of 10 over ultimate strength of the material. All other structures and fasteners should be designed with a safety factor of at least 5 |  | Design | I
10 year operating lifetime |  | Design | I
8.4.3 Standards Related Design Requirements
Enclosures conform to UL 508 |  | Design | I
NEMA type 4 or better enclosure to NEMA 250-1997 |  | Design | I
Mechanical drawings to ANSI Y14.5M-19994 (R1999) and ASME standard Y14.100-2000 |  | Design | I
8.4.4 Integration Related Design Requirements
8.4.4.1 Mounting Position |  | Design | I
8.4.4.2 Handling
All fixtures and equipment needed to disassemble the instrument dewar |  | Design | I
Service fixtures or stands designed with as low of a center of gravity as possible to resist tipping |  | Design | I
Seismic restraints as required |  | Design | I
9 Electronic/Electrical Requirements (TBD)
10 Safety Requirements
10.3 Performance Requirements
10.3.2 Operational Performance Requirements
No safety hazard to personnel or equipment |  | Design | D | IRMS is designed for operation by qualified personnel only
Interlocks, labeling and procedures provided |  | Design | D
Safety review completed? |  | Design | I
10.4 Implementation Requirements
10.4.1 Feature Implementation Requirements
10.4.1.1 Local Control
IRMS mechanism remote control lockout |  | Design | D
Rotator local control switch |  | Design | D | Applies to rotator only
Rotator motion stop switch |  | Design | D | Applies to rotator only
Telescope emergency stop is implemented |  | Design | D | Applies to rotator only
10.4.1.2 Mechanical
Pinch hazards clearly marked with a hazard warning label or equipped with shrouds |  | Design | I
Rotator:
Mechanical lockout feature |  | Design | D
Electrical lockout |  | Design | D
10.4.4 Regulatory Implementation Requirements
See §6.3.3
10.5 Design Requirements
10.5.1 Technological Design Requirements
10.5.1.1 IRMS Instrument Dewar
No motion when power applied |  | Design | T
Servo position “wind up” prevention |  | Design | I | Not applicable to instrument mechanism stepper motors which operate open loop. Not applicable to FCS piezo
Inhibit motion on position error |  | Design | T | Not applicable to instrument mechanism stepper motors which operate open loop. Not applicable to FCS piezo
N.C. contact limit switches |  | Design | D
Disconnected limit appears active |  | Design | T
Motion limit error sensing |  | Design | T | Not applicable to instrument mechanism stepper motors which operate open loop. Not applicable to FCS piezo
Encoder status loops |  | Design | I | Not applicable to instrument mechanism stepper motors which operate open loop. Not applicable to FCS piezo
10.5.1.2 Rotator
No motion when power applied |  | Design | T
Servo position “wind up” prevention |  | Design | I
Inhibit motion on position error |  | Design | T
N.C. contact limit switches |  | Design | D
Disconnected limit appears active |  | Design | T
Motion limit error sensing |  | Design | T
10.5.2 Regulatory Design Requirements
Conform to all applicable regulatory requirements |  | Design | I
11 Software Requirements (TBD)
12 Interface Requirements
12.2 Performance Requirements
12.2.2 Operational Performance Requirements
12.2.2.1 Handling
12.3 Implementation Requirements
12.3.1 Feature Implementation Requirements
12.3.1.1 Optical Requirements
12.3.1.2 Mechanical
12.3.2 Common Practice Implementation Requirements
12.3.2.1 Glycol Cooling
12.3.2.2 Vacuum and Cryogenics
12.3.2.3 Stray Light
12.4 Design Requirements
12.4.4 Integration Related Design Requirements
12.4.4.1 Optical Interface
12.4.4.2 Mechanical Interface
12.4.4.3 Electrical/Electronic Interface
14 Spares Requirements
Spares for all "consumable" or limited lifetime components |  | Design, Fab | I
20% of the in service quantity (at least 1 unit) |  | Design, Fab | I
15 Service and Maintenance Requirements
Provisions for disassembly |  | Design | I
Handling fixtures and any specialized tools provided with IRMS |  | Design | I
Written procedures provided |  | Doc | I
16 Documentation Requirements
16.1 Documentation Package
System overview and design description |  | Doc | I
User’s manual |  | Doc | I
Revised fabrication/procurement drawings, specifications, and schematics |  | Doc | I
Bills of material including supplier information for all components |  | Doc | I
Maintenance manual:
Procedures for handling, assembly and disassembly accurately reflecting the as-built instrument |  | Doc | I
Routine maintenance and inspection procedures |  | Doc | I
Alignment procedures |  | Doc | I
Troubleshooting guide |  | Doc | I
Repair procedures |  | Doc | I
Acceptance Test Plan documents and all performance data and results of acceptance testing |  | Doc | I
Descriptions of all recommend spare parts |  | Doc | I
All manufacturer’s manuals and documentation for COTS components |  | Doc | I
All software design documents and related documents |  | Doc | I
Safety plan and procedures |  | Doc | I
16.2 Drawings
16.2.1 Drawing Standards
SI units, dimensioned in millimeters |  | Doc | I
Optical components conform to ANSI/ASME standard Y14.18M-1986 |  | Doc | I
Mechanical drawings shall conform to ANSI Y14.5M-1994 (R1999) and ASME standard Y14.100-2000 |  | Doc | I
Each sheet conforms to ANSI Y14.1-1995 (R2002), |  | Doc | I
Each drawing shall have a title block with required information |  | Doc | I
All drawings include parts and materials lists in accordance with ANSI Y14.34‑2003 |  | Doc | I
Assembly drawings shall include all relevant views, including isometric and exploded |  | Doc | I
All detail drawings include all views, geometry, dimensions and feature controls required to duplicate the part in accordance with ANSI Y14.5M-1994 (R1999) |  | Doc | I
Multi and sectional view drawings developed in accordance with ANSI Y14.3M-1994 |  | Doc | I
Fluid power system schematics drawn in accordance with ASME Y32.10-1967 (R1994) |  | Doc | I
Dimensions and tolerances indicated in accordance with ANSI 14.5M-1994 (R1999) |  | Doc | I
Surface finishes described in accordance with ANSI 14.5M-1994 (R1999) |  | Doc | I
Electronic drawing format is at least SolidWorks 2012 |  | Doc | I
CAD software for 3D drawings is SolidWorks |  | Doc | I
16.2.2 Required Drawings
As‑built detailed mechanical drawings |  | Doc | I
As‑built detailed drawings for all optical components |  | Doc | I
As‑built assembly drawings for all assemblies |  | Doc | I
16.3 Electrical/Electronic Documentation (TBD)
16.4 Software Documentation (TBD)