Keck Adaptive Optics Note 510

NGAO System Design Phase

Technical Risk Evaluation

Peter Wizinowich, Rich Dekany, Don Gavel, Claire Max, Christopher Neyman

March 3, 2008

Abstract

This note is intended to summarize the significant technical risks associated with the NGAO system as identified during the system design phase.  The current version represents the fourth iteration on the risks (the first iteration was included in the proposal (KAON 400), and iterations were performed during and after the system architecture phase of the system design).  The risk evaluation methodology is described in section 1.  Section 2 identifies and ranks the technical risks, including risk mitigation options, and presents the resultant risk matrix.  The final section provides more specific risk mitigation plans for the remaining design phases.

1. Methodology

The JPL risk evaluation matrix approach used for the Keck Interferometer was selected to track the significant technical risks.  This matrix ranks each risk by the consequences and likelihood of the risk occurring.  A scale of 1 to 5 is used with higher numbers representing higher risk.

Likelihood of Occurrence:

| Level | Definition |
| --- | --- |
| 5 | Very High > 70%, almost certain |
| 4 | High >50%, more likely than not |
| 3 | Moderate >30%, significant likelihood |
| 2 | Low > 1%, unlikely |
| 1 | Very Low <1%, very unlikely |

Consequence of Occurrence

(replaced JPL’s usage of “mission return” with “science return”):

| Level | Performance Risk Definition |
| --- | --- |
| 5 | Project Failure |
| 4 | Significant reduction in science return |
| 3 | Moderate reduction in science return |
| 2 | Small reduction in science return |
| 1 | Minimal or no impact to science return |

(replaced JPL’s usage of “launch” with “schedule”)

| Level | Implementation Risk Definition |
| --- | --- |
| 5 | Overrun budget & contingency. Cannot deliver. |
| 4 | Consume all contingency, budget or schedule |
| 3 | Significant reduction in contingency or schedule slack |
| 2 | Small reduction in budget or schedule slack |
| 1 | Minimal reduction in budget or schedule slack |

A JPL-format technical risk matrix using these definitions is shown in the next section.  In this risk matrix red represents high risks that require implementation of new processes or a change in the baseline plan, yellow represents medium risks that need to be aggressively managed including considering alternative approaches, and green represents relatively low risks that should at least be monitored.

2. Technical Risks Identification and Ranking

2.1 Technical Risk Matrix

The current risk matrix is shown in the following Figure.

2.2 Significant Technical Risks

The following table lists the significant technical risks that have been identified.  The risks have been sorted in descending order by highest combined consequences and likelihood scores, followed by highest likelihood and highest consequence.  Each risk has a unique number, a trend column which will be used for tracking which way the risks are moving, a consequence ranking, a likelihood ranking, a description, the status of the risk and plans for mitigation.

| # | Trend | Conse-quence | Like-lihood | Description | Status | Mitigation |
| --- | --- | --- | --- | --- | --- | --- |
| 1 |  | 3 | 4 | Inadequate PSF calibration to support precision astrometry, photometry and companion sensitivity science. | The importance of PSF calibration and approaches to this calibration are documented in KAONs 474, 480 and 497. Based on thes KAONs, NGAO team members wrote a 2-year PSF reconstruction proposal to the CfAO that was funded; work began in Nov/07. Science cases for which high accuracy PSF calibration has the most impact are: Galactic Center General Relativistic effects astrometry, narrow-field proper motion astrometry, and two target sets for detection of planets around low mass stars. | 1) Participate in and monitor the CfAO funded effort, including support for PSF reconstruction tests & demonstrations with the existing Keck AO systems. 2) PSF calibration system-level preliminary design and design of the PSF calibration sequences and pipeline. 3) Collaborate with others to implement an atmospheric profiler for Mauna Kea, in support of PSF calibration. |
| 2 |  | 4 | 3 | Required lasers unavailable | Performance budget assumes SOR-type laser return. We would need more laser power from an LMCT-type laser. | 1) See discussion in programmatic risk KAON 566. |
| 3 |  | 2 | 4 | Inadequate sky coverage to support the wavefront error budget and hence science cases. | All of our evaluations have assumed AO corrected low order wavefront sensors using low noise near-IR detectors. These evaluations are documented in KAONs 470, 487, 492 and 504. | 1) Prototype a near-IR tip/tilt sensor to demonstrate that adequate detectors are available and that the AO correction is adequate. 2) Demonstrate the technique in the lab and/or on-sky. |
| 4 |  | 2 | 4 | Wavefront error budget not achieved due to inadequate assumptions and calculations | The largest issues are addressed in this table as separate risks. In addition, we anchored our assumptions through site monitoring (KAONs 303, 415, 496, 503), sodium return measurements (KAONs 416, 417, 419), comparison to the Keck II AO performance (KAON 461), and understanding of the telescope wavefront error contributions (KAONs 468, 469, 482) | 1) Anchor to Keck II NGWFC results. 2) Revisit all the assumptions and calculations during the preliminary design phase. |
| 5 |  | 3 | 3 | Inadequate tomographic reconstruction accuracy to support the wavefront error budget and hence specific science cases. | Laser tomography of the atmosphere has not yet been demonstrated or used for AO science. Tomography experiments with natural guide stars (ESO's MAD and Palomar's PT) have begun to set limits on tomographic measurement error. We have compared multiple tomography codes as part of the system design phase (KAON 475); our NGAO tomography predictions have been made using one of these codes (KAON 429). | 1) Closely monitor the results of tomography experiments being performed by other groups (Gemini, MAD, MMT). 2) Continue to perform lab experiments at the LAO directly in support of NGAO tomography issues. |
| 6 |  | 3 | 3 | Astrometry performance requirement not achieved | Astrometry error budget not yet adequately understood. Current understanding and recommendations are summarized in KAON 480. | 1) We will continue to work with the UCLA Galactic Center team and with the CIT proper motions team to understand the limitations imposed by the existing Keck AO system and science instrument. 2) A full error budget will be developed during the preliminary design phase. |
| 7 |  | 3 | 3 | Tomographic reconstruction computer architecture not yet tested in hardware | The proposed architecture is is a new, untried concept for astronomy. Massively parallelized architectures are assumed for cost and performance reasons. | 1) Monitor the progress of other projects (i.e., P3K). 2) Benchmark on a software simulator. 3) Benchmark on a scaled subset of hardware. |
| 8 |  | 4 | 2 | Complexity and instability of interactions in the overall software control system | NGAO will be significantly more complex than the existing Keck AO system with many more potentially interacting control loops and significantly more motion control. We are addressing this issue with significant attention to the final science operations product and utilization of significant operational "lessons learned" experience from the current LGS AO system. | 1) Good system-level design with attention to science observing sequences and operations. |
| 9 |  | 4 | 2 | Adequate wavefront sensor CCDs not available | Fast low-noise high pixel count detectors required. CCID-56 is a prime candidate but is not yet available. | 1) Monitor the progress of the AODP-funded CCID-56 project. 2) Evaluate alternative options. |
| 10 |  | 2 | 3 | Photometry performance requirement not achieved | Error budget not adequately understood. Current understanding and recommendations summarized in KAON 474. | 1) Develop a more complete understanding of this performance budget, which will be closely tied to the quality of PSF determination (listed as a separate risk item). |
| 11 |  | 2 | 3 | Inadequate tip/tilt performance for 1st relay DM mounted on a tip/tilt stage | 1st stage DM is ~100 mm in diameter. We do not yet know whether the required tip/tilt bandwidth can be achieved. | 1) Evaluate BW requirement for the 1st stage DM given that tweeter MEMs may be on their own tip/tilt stages. 2) Determine achievable tip/tilt bandwidth performance during preliminary design. 3) Develop options for a separate tip/tilt mirror (either use existing fold or redesign). |
| 12 |  | 2 | 3 | Keck Interferometer needs not met. | Needs & options documented in KAONs 428 and 483. Conceptual design for implementation partially addressed in the optical relay system design report (KAON 549). A complete layout of the interferometer feed and an analysis of the polarization impact needs to be developed next. | 1) Complete the preliminary design with the interferometer requirements in mind. |
| 13 |  | 3 | 2 | Complexity and cost for an adequate multiple LGS projection and sensing system. | An ~ 9 LGS system is needed to achieve the requirements. Design and packaging of both the laser projection and LGS wavefront sensor package to meet the performance requirements will be challenging. | 1) Complete the preliminary design. |
| 14 |  | 3 | 2 | No workable design for deployable near-IR integral field spectrograph (d-IFS). | This instrument will be key to the future of extragalactic science with NGAO. A conceptual design has been developed for the object selection mechanism for NGAO's low order wavefront sensors, which could also apply to the d-IFS heads. | 1) Complete the conceptual design for the d-IFS instrument. 2) Complete the preliminary design for the IFU and low order wavefront sensor heads, during the NGAO preliminary design phase. |
| 15 |  | 3 | 2 | Rayleigh-scattered background on LGS WFS cannot be calibrated out. | A first-order Rayleigh rejection trade study has been performed (KAON 490) | 1) Learn from Gemini MCAO experience. 2) Perform additional modeling during preliminary design phase. |
| 16 |  | 3 | 2 | K-band background requirement not achieved. | The background performance budget was evaluated in KAON 501 with the conclusion that the AO optics need to be cooled to 260 K. | 1) Re-evaluate expected K-band background based on conceptual & preliminary optical designs. 2) Develop enclosure design. 3) Revisit science requirement, especially for narrow field NIR instruments. |
| 17 |  | 3 | 2 | Inadequate MEMS performance | Good lab experience, but little on-sky performance data and no on-sky lifetime data yet. Impact of windows on throughput could be an issue. | 1) Monitor VILLAGES on-sky experiments and LAO lab experience. 2) Evaluate pros/cons of various window options (including no window) on MEMS. |
| 18 |  | 3 | 2 | Woofer-tweeter performance inadequate. | Woofer-tweeters are used for the science instruments & the low order wavefront sensors. The tweeters are used open loop. Such systems have not yet been demonstrated on-sky. | 1) Need to design servo control. 2) LAO & others (i.e., U. Vic) are doing woofer-tweeter demonstrations. Need to quantify the performance results. |
| 19 |  | 3 | 2 | Space command shutdowns | Recently US Space Command has been requiring frequent blackout periods. Keck is already actively pursuing this issue in collaboration with other Observatories. | 1) Monitor progress of current discussions with Space Command. |
| 20 |  | 1 | 3 | Low mass star companions sensitivity science requirement not achieved | Current requirements & understanding of companion sensitivity documented in KAON 497. Current low mass stars companion sensitivity requirements are only partially met by a simple coronagraph (unlikely to have an advanced coronagraph in 1st generation instruments). | 1) Revisit science requirements in light of KAON 497. In particular, re-evaluate the three target sets discussed in the SCRD. |
| 21 |  | 2 | 2 | Science requirements inadequately understood and/or defined. | Science cases requirements are defined in KAONs 455 and 456. All Key Science Drivers are complete. Case development for four of the Science Drivers is not yet complete; work is ongoing. | 1) Complete the science case requirements for the last four Science Drivers. 2) Revisit Key Science Drivers and Science Drivers as needed during preliminary design. |
| 22 |  | 2 | 2 | Required dichroic performance not achieved | Multiple large dichroics with excellent performance are required to meet the throughput, background and wavefront performance budgets. Dichroic changers must be repeatable. | 1) Discuss with vendors. 2) Evaluate whether coating tests are valuable. 3) Design dichroic changer. |
| 23 |  | 3 | 1 | Impact of telescope vibrations on wavefront error budget higher than predicted. | Measurements and impact documented in KAON 482. Tip/tilt vibrations from Keck II experience added to performance budget tool. | 1) Improve tip/tilt vibration model in the wavefront budget during the preliminary design. 2) Reduce tip/tilt vibrations. |
| 24 |  | 2 | 1 | LOWFS-based tip/tilt correction for narrow field science instruments inadequate | Tip/tilt errors from the 2nd AO relay or opto-mechanical drifts will not be sensed by the low order wavefront sensors. | 1) Evaluate the error budget impact. 2) Design a stable system. 3) Design a metrology system if needed |

4.  Technical Risk Mitigation Plans

4.1 Preliminary Design Phase

Many of the above technical risk items will be addressed during the preliminary design phase.  Specific actions planned for the preliminary design phase, on the medium to high risk items, include the following (the numbers below correspond to the first column in the above risk table):

- PSF calibration.

- Collaboration and support for the CfAO-funded PSF reconstruction proposal.  Support the effort for 2nd year funding for this project.

- Produce a system level design for PSF calibration.  This includes defining the requirements for the hardware and software that will be needed to support PSF calibration, including such things as implications for observing and calibration sequences, the real-time control system design and data products, atmospheric profilometry and data post-processing.

- Investigate options for collaborative implementation of an appropriate atmospheric profilometer on Mauna Kea including the Chun et al. proposal to the Mauna Kea Directors and reuse of the TMT monitoring equipment.

- Lasers.

- See the Programmatic Risks KAON 566.

- Sky coverage.

- Initial plans included prototyping of a tip/tilt sensor and object selection mechanism (as proposed in KAON 565) during the preliminary design.  For budget reasons we have reduced the available preliminary design phase budget for this work and hence the scope.  However, Caltech and HIA are collaborating on a TMT IRIS LOWFS study and we have recently started investigating whether a TMT/NGAO collaboration is possible on this design/implementation.  Some cost sharing, if available might allow us to get further in this prototyping.

- If the above collaborative prototyping is not feasible we will determine whether we can move the fabrication of the first LOWFS assembly into the detailed design phase in order to identify problems early.

- We will continue with the preliminary design of the LOWFS assembly including investigating the availability of the required sensors.

- Wavefront error budget.

- Revisit the assumptions and calculations.

- Tomographic reconstruction.

- Monitor and summarize the results of tomographic experiments, and compare with our simulation codes as appropriate.

- Define and implement appropriate experiments at the LAO.

- Astrometry performance.

- Develop a full astrometry performance budget.

- Work with the UCLA Galactic Center team and CIT proper motions team to develop and anchor this budget.

- Tomographic reconstruction computer architecture.

- Monitor the progress of similar designs and implementations.

- Proceed with the preliminary design and analysis.

- Design a benchmark experiment to be performed during the detailed design phase.

- Software control complexity.

- Complete the preliminary design of the overall software architecture.

- Wavefront sensor CCDs.

- Monitor the progress of suitable CCDs developments.

- Photometry performance.

- Develop a better understanding of the photometry performance budget.

- Inadequate woofer tip/tilt performance.

- Develop a preliminary design that includes discussions with vendors.

- Use experiments at LAO to anchor modeling.

- Interferometer needs.

- Complete the optical relay design for the interferometer feed.

- Evaluate the polarization and field/pupil rotation implications of this design, and redesign as appropriate, including consideration of the option of modifying the existing Keck AO or interferometer optics to work with NGAO.

Appendix I.  Preliminary Technical Risk Evaluation Results dated August 15, 2007

As of August 15, 2007 the NGAO system architecture down select process was still on going.  Five possible architectures were identified during the team work shop at UC Santa Cruz in early July (see KAON 499).  The risk table shown in this Appendix was compiled to assist in the down-selection process for NGAO system architecture and as a repository for risk evaluation for the final system design report.

This appendix explains our initial methodology for risk evaluation.  It also included a risk table with about 70 items. These are organized by system level risks, component risks, and architecture specific risks.

# Methodology

Risk areas in the project were identified.  Each risk was evaluated for its impact and likelihood.  Impact level was assigned in one of the following 4 categories:

- Major - Project objectives at risk (mandatory change to one or more of project scope, schedule, or resources)

- Moderate - Project objectives can still be met, but would require significant changes to plan

- Low - No major plan changes required; the risk is an inconvenience or it will be addressed through minor allocation of contingency resources

- Unknown - Impact is not quantifiable at this time

These categories are broad and include cost and schedule risk as inclusively.  Each risk was also assigned a likelihood of occurrence based on the following 3 categories.

- Likely - 50% or higher

- Unlikely - 10% to 50%

- Very unlikely - 10% or less

Using these criteria the following risk table was developed, see Table 1.

Table 1: Risk evaluation for NGAO as of August 15, 2007.

| Ref. # | Description | Impact | Likelihood | Mitigation Plans | System Design Phase Mitigation |
| --- | --- | --- | --- | --- | --- |
|  | System Level Risks |  |  |  |  |
| 1 | Achieving science requirements |  |  |  |  |
| a | Long exposure time performance | Moderate | Likely | On instrument metrology |  |
| b | (add other parameters?) |  |  |  |  |
| 2 | Science requirements inadequately understood & changing | Major | Unlikely | Talk to the astronomers a lot | Science Core Requirements document and flowdown |
| 3 | Delivered PSF too variable (spatially and temporally) to satisfy astrometry and photometry requirements | Moderate | Likely |  | CfAO funded PSF study (2008-2009) |
| 4 | Adequately meeting interferometer needs | Unknown | Likely | Review proposed performance with KI team | KI support trade study and KAON |
| 5 | Achieving contrast performance budget | Unknown | Unlikely | (Need to verify science requirements) | Contrast trade study & KAON |
| 6 | Achieving defined photometry budget | Unknown | Unlikely | (Need to verify science requirements) | Photometry trade study & KAON |
| 7 | Achieving defined astrometry budget | Unknown | Unlikely | (Need to verify science requirements) | Astrometry trade study & KAON |
| 8 | Achieving desired SNRs | Unknown | Unlikely | Managing throughput in optical design, making provisions for long exposure stability | Throughput and background trade study & KAON |
| 9 | Achieving polarimetry requirement | Unknown | Unlikely | Control effects that rotate or scramble polarization | Polarimetry trade study |
| 10 | Wavefront error budget assumptions & accuracy |  |  |  |  |
| a | Bandwidth error assumptions. Assumption that closed loop bandwidth is 1/15 of sample rate. The rate of ~1/20 has been demonstrated, but would significantly impact error budget. | Moderate | Unlikely | Investigate and simulate control loop impact. |  |
| b | Sodium return expectations not met | Major | Likely | Refine and adjust assumptions based on data from current systems | See LAO web page for current info |
| c | low noise CCDs for WFS. | Major | Unlikely | Another design turn for CCID-56, more laser power | AODP CCD project |
| d | Impact of telescope vibration | Moderate | Likely | Reduce telescope vibrations | Vibration trade study & KAON |
| e | Tomography. No sky demonstration. |  |  |  | MAD, MMT, LAO bench experiment |
| i | Codes contain assumptions that are untested in actual operating conditions | Major | Likely | Refine and adjust assumptions based on testing |  |
| ii | Alignment and registration - beacons and WFS | Moderate | Unlikely | Design opto-mechanics for closed loop beacon positioning and stability. Implement test procedures during I&T to ensure proper alignment and registration. |  |
| f | Tip/tilt tomography. No sky demonstration of benefits of multiple TT stars | Moderate | Unlikely | LAO laboratory experiments |  |
| g | Rotating LGS constellation limits performance for long exposures | Moderate | Likely | De-rotate, configurable add beacons? |  |
| h | MCAO mirrors are not at proper conjugates or correct "statistical position" for the actual Cn2 profile | Moderate | Unlikely | Get MASS/DIMM data for Mauna Kea before detailed design phase | Collate TMT and other Mauna Kea seeing measurements |
| 11 | Risk of not being able to find adequate tip/tilt stars for certain science cases | Low | Likely | System provides gradual degradation, TT stars AO corrected |  |
| 12 | Rayleigh background on LGS WFS cannot calibrated out | Major | Likely | Issue for GS MCAO, will be tested by them. Use long period pulsed laser and electronic shutter on HOWFS CCD to gate out Rayleigh | Rayleigh Rejection trade study |
|  | Component Risks |  |  |  |  |
| 13 | Availability of required lasers | Major | Likely | Continue to pursue laser development |  |
| 14 | Fiber transport. Mitigation is conventional beam transport | Moderate | Likely | Testing programs underway for fibers | Keck I LGS experiment with Gemini Laser, Subaru Fiber; ESO & Subaru experience |
| 15 | Availability DM with small pitch and adequate stroke | Major | Unlikely | Use 48 x 48, 5 mm, add a second DM |  |
| 16 | MEMs mirror window/no window | Moderate | Likely | Account for windows in budget or develop MEMs without windows |  |
| 17 | MEMs mirror lifetime | Major | Likely | Work with MEMs vendors, other AO project |  |
| 18 | DM on a tip/tilt stage | Major | Likely |  |  |
| a | DM incompatible with operation on TT stage | Major | Unlikely | Use a separate TT mirror |  |
| b | Problems with DM interface cabling on TT stage | Major | Likely | Address in DM design |  |
| c | Insufficient TT rejection | Moderate | Unlikely | Add a second TT mirror |  |
| 19 | Switchyard approach: |  |  |  |  |
| a | Dichroics. Size and performance. | Major | Likely | (at report we will not have this level of risk). Test coating samples to confirm performance before completing design |  |
| b | Performance and reliability of dichroic changers. | Moderate | Unlikely |  |  |
| 20 | K-mirror. Size, performance. | Moderate | Unlikely | Other architectures for derotation, better coatings |  |
| 21 | Achieving real-time control performance requirements | Major | Unlikely | Benchmark tests, simulations anchored to RTC hardware performance, prototype testing |  |
| 22 | Fitting system on telescope | Major | Unlikely | Design process will ensure compatible system |  |
| 23 | Thermal/mechanical performance of AO system environmental enclosure | Moderate | Unlikely | Careful design, thermal performance modeling including FEA |  |
| 24 | Design & cost of interfacing with existing instruments exceeds value of doing so | Unknown | Unlikely | Replace those instruments |  |
| 25 | MOAO not demonstrated. | Moderate | Likely | MCAO gives reasonable sky coverage, VILLAGES testing planned. Other testing programs, perhaps on existing Keck AO system. |  |
| 26 | Fast LOWFS IR based camera |  |  |  |  |
| a | Detector performance | Moderate | Unlikely | Some performance data on hand. Testing continues. |  |
| b | Detector availability | Major | Unlikely | Two sources of supply |  |
| 27 | Calibration unit with LGS simulators |  |  |  |  |
| a | Finding space for it | Major | Unlikely | Will be designed-in from the beginning as an essential capability |  |
| b | Achieving required level of performance | Moderate | Unlikely | On-sky calibration can substitute at greater expense |  |
| 28 | Rayleigh Laser |  |  |  |  |
| a | Complexity of pulse tracker or other range gate method | Unknown | Unlikely |  |  |
| b | Additional background light | Major | Unlikely |  |  |

Table 2: Architecture specific risk for NGAO, as of August 15, 2007.

| Ref. # | Description | Impact | Likelihood | Mitigation Plans | System Design Phase Mitigation |
| --- | --- | --- | --- | --- | --- |
|  | candidate 1 split relay |  |  |  |  |
|  | Unclear that dNIRI can fit close enough to elevation journal | Major | Likely |  |  |
|  | Unclear if there is enough space for dNIRI & narrow field at same time | Major | Likely |  |  |
|  | Calibration: the non-common path accuracy between the TT location and narrow field science instruments (This is particularly true due to the adoption of rotators over a single k-mirror field de-rotator) | Major | Likely |  |  |
|  | candidate 2 AO secondary |  |  |  |  |
|  | Development of AM2 is costly and uncertain | Major | Likely |  |  |
|  | Actual tip/tilt performance of the AM2 is Unknown | Major | Likely |  |  |
|  | Fitting error for AM2 worse than expected | Major | Likely |  |  |
|  | candidate 3 large relay |  |  |  |  |
|  | Large instrument that needs to be cooled | Major | Likely |  |  |
|  | Unclear if the instrument will fit on the platform | Moderate | Unlikely |  |  |
|  | MCAO option only provides 60” field fully corrected (50% EE) | Moderate | Unlikely |  |  |
|  | MCAO requires 2 DMs, one at ground and one at 5km | Moderate | Unlikely |  |  |
|  | candidate 4 Keck I upgrade |  |  |  |  |
|  | Higher background | Moderate |  |  |  |
|  | Parts not designed for Low Temp operation | Major | Likely |  |  |
|  | dNIRI feed hard to fit in front of AO | Moderate | Likely |  |  |
|  | Some of the hardware will be obsolete by the time of NGAO | Moderate | Likely |  |  |
|  | candidate 5 cascade relay |  |  |  |  |
|  | Cannot be packaged | Major | Likely |  |  |
|  | Cannot support interferometer | Moderate | Unlikely |  |  |
|  | High emissivity | Major | Unlikely |  |  |
|  | Complication of woofer-tweeter control | Major | Likely |  |  |
|  | Lower transmission for both the LGS path (loss of laser return) and instruments path (reduced sensitivity, but potentially offset by higher Strehl with less risky architectural approach) | Major | Likely |  |  |
|  | LOWFS away from science instruments, though all are not rotating | Moderate | Likely |  |  |
|  | Potentially more scintillation, static aberrations due to large number of surfaces that need to be controlled | Moderate | Likely |  |  |