DRAFT

Next Generation Adaptive Optics Preliminary Design

Compliance Matrix

Keck Adaptive Optics Note 772

Christopher Neyman

May 24, 2010

Abstract

A discussion of the compliance verification process for the Next Generation Adaptive Optics (NGAO)

This work was part of the preliminary design phase of the NGAO project, a standard phase of new instrument development at W. M. Keck Observatory.

Revision History

| Revision | Date | Author (s) | Reason for revision / remarks |
| --- | --- | --- | --- |
| 1.0 | May 24, 2010 | C. Neyman | Initial release |

# Introduction

Before adaptive optics observations can begin, the various NGAO wavefront sensors, deformable mirrors, and

# Background

# Summary Statistics

System Requirements All Subsystems:

The one requirement not in compliance (denoted “No” in table below) is related to the mass limited requirement for systems that are mounted inside the secondary mirror support structure. The LGS Facility design team commented, “The mass of the components will exceed that in the spec.  However, there is an existing 700 pound counterweight that will be removed to offset any weight increase by the components in the secondary.”  This issue will be addressed during the detailed design.  The two requirements listed as partially compliant involve documentation requirements.  It is not expected at preliminary design review to be able to satisfy the requirement for all final documentation.  The other systems all listed this as being verified only at detailed design or during delivery of the subsystem.  Two of the three requirements listed at “Goal” are all related to failure testing of the software.  Software methodology cannot guarantee compliance during the design phases, the possibility exists of unknown bugs that may not be found until the system is built and tested.  Software unlike hardware cannot test many identical units to determine a statistical mean time between failure.  NGAO control software will be verified to the extent possible during full scale development of software.  The last requirement that is listed as “Goal” involves the requirement to simulate the telescope functions with the AO control software.  The software design team commented that they are not currently tracking this as a design feature of controls system, but it could be added during the detailed design phase.

|  | N/A | DD | Yes | No | Partial | Goal |
| --- | --- | --- | --- | --- | --- | --- |
| Compliance of 11 subsystems for all 102 System Requirements | 568 | 378 | 170 | 1 | 2 | 3 |
| % excluding N/A |  | 68.23% | 30.69% | 0.18% | 0.36% | 0.54% |
| % including N/A | 50.62% | 33.69% | 15.15% | 0.09% | 0.18% | 0.27% |

Overall AO:

The one requirement listed at partial involves not blocking the Keck telescope pupil.  The design team of the NGS WFS notes that: “The sensors low order 5x5 mode will sometimes loose light (during extreme nutation).”  The one requirement listed as a “Goal Only” by the NGS WFS design team,  involves the non-operational conditions (temperature, humidity, etc.) in the observatory dome environment; no reason was given for the noncompliance.    Both issues will be addressed during the detailed design phase.

|  | N/A | DD | Yes | No | Partial | Goal |
| --- | --- | --- | --- | --- | --- | --- |
| Compliance of 8 major subsystem for all 30 Overall AO functional Requirements | 102 | 80 | 56 | 0 | 1 | 1 |
| % excluding N/A |  | 57.97% | 40.58% | 0.00% | 0.72% | 0.72% |

| Subsystem Specific Functional Requirements | DD | Yes | No | Partial | Goal |
| --- | --- | --- | --- | --- | --- |
| AO Enclosure | 12 | 18 | 0 | 0 | 0 |
| AO Bench | 21 | 47 | 0 | 2 | 0 |
| LGS WFS | 11 | 24 | 1 | 0 | 0 |
| NGS WFS | 10 | 12 | 1 | 0 | 0 |
| LOWFS | 5 | 11 | 0 | 0 | 0 |
| Acquisition Camera | 5 | 11 | 0 | 7 | 0 |
| Alignment, Calibration, & Diagnostics | 2 | 39 | 0 | 0 | 0 |
| LGS Facility | 26 | 103 | 1 | 1 | 0 |
| Controls | 8 | 48 | 1 | 0 | 0 |
| Science Operations Tools | 8 | 38 | 0 | 3 | 0 |

AO Enclosures:

Design is compliant to the level that it can be verified at preliminary design.

AO Bench:

The two requirements that are listed as partially compliant are related to Lateral Color of the wide field and narrow field relays respectively.  The specifications are not met for all bands.  This issue will be followed up on during the detailed design phase.

LGS WFS:

The one requirement listed as not compliant (No) is related to the Patrolling Asterism LGS WFS.  The requirement specifies an operating dynamical range of at least plus/minus 1.4 arc seconds (on-sky).  The design has plus/minus 0.86 arc seconds.  This issue will be followed up on during the detailed design phase.

NGS WFS:

The one requirement listed as not compliant (No) appears to be an error in the original draft of the requirement or the interpretation of the requirement by the NGS WFS design team.  This item has been noted in the Contour requirements database and will be addressed during the final parts of the preliminary design phase (June 2010).

LOWFS:

Design is compliant to the level that it can be verified at preliminary design.

Acquisition Camera:

Design is compliant to the level that it can be verified at preliminary design.

Alignment, Calibration, and Diagnostics:

Requirements that are listed as partially compliant can be organized into two groups.  These are related to four requirements for an atmospheric simulation capability that are only partially supported by the calibration unit design.  At present this deficiency is acceptable, but will be revisited during the detailed design phase.  The other partially compliant requirements (three in total) are related to the formal statement of the requirements on calibration source brightness in terms of the final SNR at the science instrument.  The calibration unit design team believes that the specifications can be met, but no formal verification has been documented, therefore they are listed as partially compliant.  The issues of required calibration source brightness will be revisited during the detailed design phase.

RTC:

Design is compliant to the level that it can be verified at preliminary design.

LGS Facility:

The one requirement listed as No, is related to the strength of material for items mounted on or above the telescope.  This issue will be addressed during the detailed design phase.   The requirement denoted as partial involves the power required in the enclosure housing the laser units mounted to the elevation ring.  The design was modified to require less power in this location, so the current enclosure design provides for less power than stated in the formal requirement.  This issue is will be addressed during the detailed design phase.

Controls:

The one requirement denoted in the table as not compliant (i.e. No) is a result of the formal statement of the requirement.  The software team will support all functionality listed in the requirement but it will not be done by a distributed control systems on a single “command processor” as stated in the requirement.   If the KCSF/EPICS framework is considered the “command processor” then the software team feels their design is compliant.  This issue will be addressed during the detailed design phase.

Science Operations Tools:

The requirements that are listed in partial compliance are related to the need to more completely specify the design, functionality, and operations which either part of the controls design and the science tools design.  Some confusion on the exact scope of each resulted in controls requirements being assigned to the Sci. Ops. Tools.  This issue will be addressed during the detailed design phase.

# Discussion

# Conclusions

This note has outlined the methods for calibrating the AO system and the science instrument before an observing run.  In some cases, these methods are unique to NGAO and should be tested during the detailed design phase on an actual ‘go-to’ adaptive optics system.  Possible locations include the UCSC Laboratory for Adaptive Optics and the VILLAGES system at Lick Observatory.  This note also outlines features for the AO system calibration unit.  Any modification will need to be included in the revised NGAO functional requirements for the calibration system and the optical design of the AO calibration source.

## References

1.