Palomar Adaptive Optics Test Plan

| Title | BTO setup |
| --- | --- |
| Version | 2.3 |
| Date released | 7/25/2007 |
| Lead | A. Bouchez |
| Time requested | 1 hr during the first afternoon of each observing run. Parts C and D must be performed after safety interlock checkout. |
| Required conditions | Dome closed |

Purpose - Power up BTO, align mirrors using 660nm stimulus beam, save default positions. - Verify BTO alignment to LLT. - Align LLT to 660nm stimulus beam.

Test procedure Notes: - Requires 1 AO/BTO operator + 1 personnel in dome. - Request that Coude block supervisor lock be unlocked before beginning procedure. - Wear protective goggles while in Coude room and dome (will need to be removed for parts of LLT alignment procedure). A. Power up Coude hardware: - Uncover BTO Coude optics. - Remove blue FSM mirror cover. - Open ports in Coude bench enclosure. - Dust optics in main beam with Aero-Duster. - Power up ESP300 and FSM controllers. - Power up power strip behind Coude bench (laser & diagnostics electronics). - Power up 660nm laser using the following procedure: - Unplug black wire from power supply. - Turn on power supply. Check that power supply is set to 3.75V, 0.75A. - Turn off power supply and reattach wire. - Power on laser. - Power up Coude lab video camera, intercom, and both microphones. - Test intercom volume. Coude controls adjusts volume as heard from data room... B. Power up Dome hardware: - Verify that telescope focus is nominal for AO (~61 mm) - Power up all BTO electronics in dome. - telnet viswa_power - > /on C2 - > /on C3 - > /on trolley - > /on bto - > /x - Wait ~3 minutes for BTO computer to boot. - Start up BTO GUI. - ssh -Y aousr@lgs.palomar.caltech.edu (198.202.125.152) - bto_ctrl - > reset all (Note: this moves the laser focus stage on laser bench!) - > set trigger=internal - > set hadec=TCS - Prepare trolley optics for operation - Move trolley to +5200000. - Remove optics covers. - Dust optics with Aero-Duster - Set Track ON to move BTO mirrors to last zenith position. C. Align BTO (after clearing personnel from dome) - Verify that Coude light path is clear of obstructions (eg. Coude M5). - Open Coude block. - Enable perimeter alarm (all other zones disabled). - Verify that track is ON. - Erect the BTO by closing the servo loop to Q1, then Q2, then Q3. - If no light is visible on Q1... - Visually check Coude beam path. - Verify that stimulus beam is precisely centered on both Coude irises. - Verify that telescope pointing data is OK (see step 9.5) - Manually offset M1x,y until light falls on Q1, the close servo loop. - Set zenith and save default (saves stimulus zenith position & BTO default) D. Align LLT to laser (requires 1 personnel in prime focus, 1 in data room) - Remove LLT primary cover, FSM cover, any other optics covers. - Verify that UTT mirror controller is powered on, and positioned at [0,0]. - Lock BTO servo loop with high speed on Q3. - Adjust the Q3 beamsplitter to center the laser on the UTT mirror. - Adjust the top fold mirror to center the laser on it. - Install reticles in the LLT primary mirror hole and on the LLT secondary. - Iterate between UTT mirror adjustments (performed by the AO operator) and adjustments of the LLT final fold mirror to center the laser on both retticles. - Remove the reticles and verify that the beam pattern on the dome ceiling is evenly illuminated and centered on the LLT secondary shadow. - Open the BTO servo loop. - Verify that the laser beam path in prime focus is clear of obstructions and everything appears ready for the high-power 589nm laser.

Results and conclusions