Palomar Adaptive Optics Test Plan

| Title | NGS AO Checkout |
| --- | --- |
| Version | 3.1 |
| Date released | 4/2/2007 |
| Lead | A. Bouchez |
| Time requested | 20 minutes every LGS night. Start after 6º twilight. |
| Required conditions | N/A |

Purpose Verify acceptable performance of the NGS AO system for subsequent engineering tests, and measuring seeing, on-axis Strehl and isoplanatic angle.

Test procedure - Slew to a V=7 A star near zenith. - Verify that AO system is set up for NGS. - Acquire the primary with the AO system with an initial frame rate of 1000Hz. Adjust the frame rate to get 100-200 cts/subap on the WFS. - Offset to sky and record a wavefront sensor sky. - Perform both coarse and fine WFS-DM registration. - Verify that co_default2 is loaded and that focus offloading is on. - Close the TT and DM loops. - Create a flat map on the star. - IDL> ao_make_dm_flat_map, filename - Load the flat map when complete. - While waiting for the flat map, check signal levels on PHARO: - Initial setup: 25mas FOV, filter BrG, ND 1%, standard cross, 2s integration. - Adjust the exposure time and filter to get 10-15k peak counts. - Open TT and DM loops, offset to sky, and record a new WFS background. - While on sky, record 3 PHARO sky frames. - Offset back to star and close TT and DM loops. - Record 3 images with PHARO. Note the directory, frame numbers, and Strehl. - Open the DM loop only. - Record 1 image with PHARO, 30s integration. Note frame number and FWHM. - Record the telescope focus.

Results (use one column per night) UT date. Weather conditions. PHARO directory. Target name. 13. Sky frames. 15. Closed-loop frames. 15. Closed-loop Strehl. 17. Open-loop frame. 17. Open-loop FWHM. 18. Telescope focus.