## Slide 1: Flat Field Forensics
Timothee Greffe – COO – 2017/07/06

## Slide 2
Intro | Downing Dip | Bilgi Bump | Next
CCD 231-C6 from e2v
Used in ZTF and WaSP cameras
All tests were performed using the WaSP camera
6k by 6k sensor
15um pitch
350 ke- Full Well Capacity

## Slide 3
Intro | Downing Dip | Bilgi Bump | Next
CCD231 Topology:
Image Area
“Up” Amplifiers
“Down” Amplifiers

## Slide 4
Intro | Downing Dip | Bilgi Bump | Next
Serial
Register
Rφ
Transfer
Gate
Iφ4
Iφ3
Iφ2
Iφ1
Potential
10
0
Reminder : Charge Transfer

## Slide 5
Intro | Downing Dip | Bilgi Bump | Next
Initial goal and field of study : Steep the PTC by fine tuning

## Slide 6
Intro | Downing Dip | Bilgi Bump | Next
Rφ
TG
Iφ4
Iφ3
Iφ2
Iφ1
Potential
10
-4
2
One parameter available : Voltage of the Image phase clocks

## Slide 7
Strange Dip appears for low common-mode voltage
Intro | Downing Dip | Bilgi Bump | Next

## Slide 8
Extract from “Bulk Silicon CCDs, Point Spread Function and Photon Transfer Curves: CCD testing activities at ESO”
Mark Downing et al.
Intro | Downing Dip | Bilgi Bump | Next

## Slide 9
Intro | Downing Dip | Bilgi Bump | Next
One might think : “ Who cares since we are not seeing anything wrong when driven with recommended voltage?”
All is not for the best, we must cultivate our garden.

## Slide 10
Intro | Downing Dip | Bilgi Bump | Next
Focus on -4V to 2V voltage swing.
Downing Dip
Blooming/FWC
How can the noise decreases for a given luminosity level and recover good statistic later?

## Slide 11
Intro | Downing Dip | Bilgi Bump | Next
Hypothesis: Drift of the whole image area after a given threshold

## Slide 12
Intro | Downing Dip | Bilgi Bump | Next
A,σA
B,σB
1
1
0.5
0.5
If uncorrelated noise
C,σC
σC ?
=> But the dip in the PTC is not ½ depth…
If correlated noise between A and B…
1
1
0.5
Additive term null in 0 and 1
0.5
What if during a pixel shift, we mix charges from a pixel to the next one?

## Slide 13
A drift mechanism can explain the “Downing Dip”
=> We have to figure out if it occurs, where, when and how….
Intro | Downing Dip | Bilgi Bump | Next

## Slide 14
Intro | Downing Dip | Bilgi Bump | Next
Hypothesis: Drift of the whole image area after a given threshold
Reference level for Cross-Correlation

## Slide 15
Intro | Downing Dip | Bilgi Bump | Next
There is indeed a drift of the whole image area during the Downing Dip!

## Slide 16
Intro | Downing Dip | Bilgi Bump | Next
TGD | IPHI4 | IPHI3 | IPHI2 | IPHI1 | IPHI4 | IPHI3 | IPHI2 | IPHI1
Time
Charge Transfer : How it should work
Serial Register

## Slide 17
Intro | Downing Dip | Bilgi Bump | Next
TGD | IPHI4 | IPHI3 | IPHI2 | IPHI1 | IPHI4 | IPHI3 | IPHI2 | IPHI1
Main Idea: What if the last IPHI4 node is missing?
Time
Serial Register

## Slide 18
Intro | Downing Dip | Bilgi Bump | Next
TGD | IPHI3 | IPHI2 | IPHI1 | IPHI4 | IPHI3 | IPHI2 | IPHI1
Recommended clocking with IPHI4 missing.
Time
Serial Register
?
?
?
?
?

## Slide 19
Intro | Downing Dip | Bilgi Bump | Next
TGD | IPHI3 | IPHI2 | IPHI1 | IPHI4 | IPHI3 | IPHI2 | IPHI1
Time
Serial Register
X
STATE I
Because of structural traps, charges can’t flow immediately to the serial register:
Everything seems to be ok.

## Slide 20
Intro | Downing Dip | Bilgi Bump | Next
TGD | IPHI3 | IPHI2 | IPHI1 | IPHI4 | IPHI3 | IPHI2 | IPHI1
Time
Serial Register
STATE II
All the charges are flowing immediately to the serial register:
=> The whole image is shifted

## Slide 21
Intro | Downing Dip | Bilgi Bump | Next
TGD | IPHI3 | IPHI2 | IPHI1 | IPHI4 | IPHI3 | IPHI2 | IPHI1
New timing to get rid of this effect:
Time
Serial Register

## Slide 22
Intro | Downing Dip | Bilgi Bump | Next
No Dip, No shift anymore!
“Downing Dip” defeated

## Slide 23
Intro | Downing Dip | Bilgi Bump | Next
Back to the recommended voltage: 0V to 10V

## Slide 24
Intro | Downing Dip | Bilgi Bump | Next
When looking closely :

## Slide 25
Intro | Downing Dip | Bilgi Bump | Next
Named after Pavan Bilgi, famous student at COO.

## Slide 26
Intro | Downing Dip | Bilgi Bump | Next
No Pixel Shifting
=> Different mechanism leads to this behavior
Phenomenon already visible when using a different voltage on the image phases.
=> Not introduced during charge transfer over the image area.
Again, an effect with threshold implied!

## Slide 27
Intro | Downing Dip | Bilgi Bump | Next
TGD | IPHI3 | IPHI2 | IPHI1 | IPHI4 | IPHI3 | IPHI2 | IPHI1
Time
Serial Register
Voltage of the serial register is increased to “help” the charges flowing through it.

## Slide 28
Intro | Downing Dip | Bilgi Bump | Next
The “Bilgi Dump” disappeared!
“Bilgi Bump” defeated

## Slide 29: Future investigations
We’d like to see the last phase node missing or tied together with the previous one.
  => But yet no certitude : Our model rely on badly understood structural trap phenomenon
Intro | Downing Dip | Bilgi Bump | Next
Call for ideas and comments.
A SWIR Camera might be used to inspect the sensor itself.

## Slide 30: Thank you for your attention
Intro | Downing Dip | Bilgi Bump | Next