## Slide 1: Performance analysis of NGAO NGSWFS with and without the IF dichroic
Caltech Optical Observatories
6th April, 2010

## Slide 2: Design without IF dichroic

## Slide 3: Input to the NGS sensor
Design characteristics:
NGS light is picked off in collimated space and focused using a (BASF2-N15-BASF2) triplet
F/# = 20.012
Plate scale = 1.063 mm/”
Without IF dichroic

## Slide 4: Input to the NGS sensor – spot diagram at the NGS sensor pick-off focal plane
Without IF dichroic

## Slide 5: OPD at the NGS sensor pick-off focal plane
Without IF dichroic

## Slide 6: Design with IF dichroic – if one does nothing (w/ the telescope) but re-focus the NGSWFS
I am using a N-BK7 dichroic that is 12.5 mm thick (and 135 mm diameter). The fold angle is 45 deg. As used by Reni in her Version 10 optical design.
The NGS WFS needs to move back by ~6 mm to accommodate the introduction of the IF dichroic.

## Slide 7: Input to the NGS sensor
Design characteristics:
NGS light is picked off in collimated space and focused using a (BASF2-N15-BASF2) triplet
F/# = 20.18495 (0.8% change)
Plate scale = 1.064 mm/”
With IF dichroic – refocus NGSWFS (instead of the telescope)

## Slide 8: Input to the NGS sensor – spot diagram at the NGS sensor pick-off focal plane
Worst case spot is almost 40% bigger, though this may not mean much for the 63x63 sampling mode. A lot of color (in the smaller spot).
One could envision re-optimizing the triplet if both modes were equally used.
With IF dichroic – refocus NGSWFS (instead of the telescope)

## Slide 9: OPD at the NGS sensor pick-off focal plane
7.7 waves as opposed to <6 waves of P-P WFE. (RMS is 1.51 waves instead of 1.15 waves.
With IF dichroic – refocus NGSWFS (instead of the telescope)

## Slide 10: With IF dichroic with telescope re-focus
In this optical design, the science instrument (and hence the WFS is static), the telescope is refocused to keep the science instrument in focus following the introduction of the IF dichroic in the optical beam.

## Slide 11: Input to the NGS sensor
Design characteristics:
NGS light is picked off in collimated space and focused using a (BASF2-N15-BASF2) triplet
F/# = 19.9856 (-0.13% change)
Plate scale = 1.053 mm/” (0.9% change)
With IF dichroic – refocus telescope (and not move NGSWFS)

## Slide 12: Input to the NGS sensor – spot diagram at the NGS sensor pick-off focal plane
Worst case spot is almost 37% bigger, though this may not mean much for the 63x63 sampling mode. A lot of color (in the smaller spot).
One could envision re-optimizing the triplet if both modes were equally used.
With IF dichroic – refocus telescope (and not move NGSWFS)

## Slide 13: OPD at the NGS sensor pick-off focal plane
8.5 waves as opposed to <6 waves of P-P WFE. (RMS is 1.6 waves instead of 1.15 waves.
With IF dichroic – refocus telescope (and not move NGSWFS)

## Slide 14: Implications
The NGSWFS sub-apertures (in 63x63 mode) will see a spot blur of ~ 3 um NGSWFS pick-off plane [the plate scale in this mode is 1.053 mm/”] – so, the effect is negligible with the introduction of the IF dichroic whether the telescope is refocused or if the WFS is moved to account for the path-length change due to the introduction of the dichoric.