## Template
data base field names | ID # | Who nominated or originated requirement | Section | Category | Sub Category | WBS | Description | Rationale | Traceability | Status | Verification method | Version | Created | Modified | Comments
data type |  | String | Selection (string) | Selection (string) | Selection (string) | Selection (string) | text | text | text | selection | selection | string | date | date | text string
comments | Just a unique number in the database (leave blank) |  | Assign to Sean's requirements document sections: should make it easier to draft requirements documents from database (See Adkins, "Draft Engineering Guideline For the Preparation of Requirements Documents" version 1.1, September 2006) |  |  | What part of the system is the requirement assigned to (by WBS) | Text of the requirement | Why requirement is needed, plus technical reference, KAONS, other requirements documents, etc. | What requirements this one is derived from: should be a reference to other data bases or design document |  | How the requirement is tested | What version | Date Created | history of modification | anything else
|  |  | pick one from list: | pick one from list: | pick one from list: | pick one from list: |  |  |  | pick one from list: | pick one from list:
|  | Chris Neyman | Overall | Performance Requirements | Parametric | 3.2.2 AO Enclosure | Cooling: The optical components are enclosed and cooled to -20 C. If needed, parts of the system, such as electronics that do not require a direct interface to the optical path, might be located in nearby housing and kept at a warmer temperature. | Sensitivity in K band requires cooling AO system to -20 C to achieve emissivity, Reference: KAON 501 | Science K band sensitivity SRD sections 2.4, 14.1,6.7, see also KAON 501 for temperature derivation | draft | test | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Overall | Implementation Requirements | Feature | 3.2.2 AO Enclosure | Input window: There shall be an input window to the cooled enclosure | An input window or windows are required to allow light into but environmentally seal off the cooled interior of the optics enclosure. The purpose is to prevent condensation from forming on optical surfaces. | KAON 499 "NGAO System Architecture Definition" | draft | inspection | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Overall | Implementation Requirements | Feature | 3.2.2 AO Enclosure | Output window: There shall be an output window to the cooled enclosure. Need for this requirement is TBD | An output window or windows are required to allow light out of but environmentally seal off the cooled interior of the optics enclosure. The purpose is to prevent condensation from forming on optical surfaces. | KAON 499 "NGAO System Architecture Definition" | draft |  | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Mechanical | Implementation Requirements | Feature | 3.2.2 AO Enclosure | Size: AO enclosure will be located within the boundary of the already existing Nasmyth platform and will not extend about the roof line of the current AO enclosure. | This is a design decision based on consideration of available space at Keck observatory | KAON 499 "NGAO System Architecture Definition" | draft | inspection | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Mechanical | Performance Requirements | Parametric | 3.2.2 AO Enclosure | Weight: AO enclosure, AO facility, instruments, and support equipment will not exceed the weight limit of the current Keck I and II left Nasmyth platforms. | This is the current observatory requirement for Nasmyth deck. The following is a point of information not a requirement: The weight limit is currently estimated at 22,000 lbs and 40 pounds per square foot (CN: must verify this and convert to SI units) | see NGAO Twiki http://.../NGAO/OptoMechanicalDesign see link to document, "Investigation of the Right Nasmyth Platform of the Keck II Telescope" by Medwadowski, 3/26/2001 | draft | analysis | 0.3 | 2007-09-11 | 2007-10-10 | This weight limit should go into the Instrument Baseline Requirements Document
|  | Chris Neyman | Mechanical | Implementation Requirements | Feature | 3.2.2 AO Enclosure | Access requirements: The enclosure and its associated mechanism and optics shall be situated on the Nasmyth platform in such a way so as to not preclude access to the telescope elevation bearing and the Nasmyth “basement” for normal telescope maintenance functions. | Need to access telescope bearing area for periodic maintenance | Observatory requirement, based on discussion with Jason Chin staff Keck Observatory (see also SRD v1.15 section 8.4 number 9) | draft | inspection | 0.3 | 2007-09-11 | 2007-10-10 | The need for access should go into the Instrument Baseline Requirements Document
|  | Chris Neyman | Mechanical | Performance Requirements | Operational | 3.2.2 AO Enclosure | Cooling and waste heat removal: The Instrument must not radiate more than 50 watts of heat into the telescope dome ambient environment. All heat generated by the Instrument in excess of this amount must be carried away by a glycol based cooling system. | In order to maintain telescope image quality dome temperature must be maintained close to predicted nighttime temperature | see section: 8.2.1.3 Power Dissipation, Instrument Baseline Requirements, Sean Adkins, Version 0.1, January 15, 200 | draft | analysis | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Mechanical | Performance Requirements | Parametric | 3.2.2 AO Enclosure | Vibration: AO enclosure, AO relay, and associated system must be consistent with observatory vibration standards. Additional standards for NGAO are TBD | Observatory requirement for new instrumentation, NGAO may impose more stringent standards | see section: 8.2.2.2 Vibration, Instrument Baseline Requirements, Sean Adkins, Version 0.1, January 15, 201 | draft | test | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Mechanical | Implementation Requirements | Feature | 3.2.2 AO Enclosure | Instrument handling: The AO enclosure will be equipped with a suitable mechanism for the installation and replacement of TBD instruments. This mechanism could be a crane, cart, or other mechanism. | It is envisioned that instruments will be installed and replaced throughout the lifetime of NGAO | Chris Neyman engineering decision, see additional requirement in section, 8.4.4.2 Handling, Instrument Baseline Requirements, Sean Adkins, Version 0.1, January 15, 2007 | draft | inspection | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Electronic/Electrical | Implementation Requirements | Feature | 3.2.2 AO Enclosure | Diagnostics and monitors: The AO enclosure will provide diagnostics and sensors for vibration. (Accelerometers) | This capability will be need to ensure that telescope and AO vibration environment is with specification for AO operations. | Engineering decision by Chris Neyman in consultation with Jason Chin | draft | inspection | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Electronic/Electrical | Implementation Requirements | Feature | 3.2.2 AO Enclosure | Diagnostics and monitors: The AO enclosure will provide diagnostics for cooling system flow rate. (Glycol flow) | This capability is needed to ensure that AO glycol cooling system is functioning correctly | Engineering decision by Chris Neyman in consultation with Jason Chin
|  | Chris Neyman | Electronic/Electrical | Implementation Requirements | Feature | 3.2.2 AO Enclosure | Diagnostics and monitors: The AO enclosure will provide temperature sensors. | This capability is needed to ensure that AO cooling system is functioning correctly | Engineering decision by Chris Neyman in consultation with Jason Chin
|  | Chris Neyman | Electronic/Electrical | Implementation Requirements | Feature | 3.2.2 AO Enclosure | Diagnostics and monitors: The AO enclosure will provide humidity sensors. | This capability is needed to ensure that AO cooling system is functioning correctly, and warm of condensation potential inside AO enclosure | Engineering decision by Chris Neyman in consultation with Jason Chin
|  | Chris Neyman | Electronic/Electrical | Implementation Requirements | Feature | 3.2.2 AO Enclosure | Diagnostics and monitors: The AO enclosure will provide particle count monitors and control the cleanliness of the AO environment. | This capability is needed to ensure that AO system will meet its emissivity specifications | Engineering decision by Chris Neyman in consultation with Jason Chin
|  | Chris Neyman | Electronic/Electrical | Implementation Requirements | Feature | 3.2.2 AO Enclosure | Diagnostics and monitors: The AO enclosure will provide video cameras for remote monitoring of the AO enclosure from the summit control rooms and by remote observers (Hawaii and mainland United States). | This capability is standard in the current AO systems | Engineering decision by Chris Neyman in consultation with Jason Chin
|  | Chris Neyman | Electronic/Electrical | Implementation Requirements | Feature | 3.2.2 AO Enclosure | Diagnostics and monitors: The AO enclosure will be able to monitor the background light levels in and around the AO bench. | This capability is standard in the current AO systems | Engineering decision by Chris Neyman in consultation with Jason Chin
|  | Chris Neyman | Electronic/Electrical | Implementation Requirements | Common Practices | 3.2.2 AO Enclosure | Background light and stray light: The encoders and other electronics will have no visible or IR light sources such as LEDs or optical encoders, etc. inside the optical path of the AO system and instruments. | common practice at observatory | see section 9.3.2.1 Stray Light, Instrument Baseline Requirements, Sean Adkins, Version 0.1, January 15, 2007 | draft | test | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Overall | Implementation Requirements | Feature | 3.2.3.1 Field Rotation | Input K-mirror: The AO system will have a derotator at the front end of the optical train. It must be located optically upstream of the science instruments, wavefront sensors, and tip/tilt sensors. | The purpose of this element is to remove the effects of field rotation caused by the AO system being located on a Nasmyth platform of an altitude over azimuth mounted telescope. It is envisioned that this can be accomplished by a K-mirror arrangement. | KAON 499 "NGAO System Architecture Definition" | draft | inspection | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Overall | Implementation Requirements | Feature | 3.2.3.2 Optical Relay | First stage relay: A main relay "first stage" that contains a deformable mirror at an optical conjugate to the telescope pupil | The telescope pupil (limiting aperture) is the primary mirror. We define this as the "ground layer" or 0km conjugate. The first stage DM is intended to correct to moderate order over a wide field of view to feed multiple deployable near-IR IFU and the LOWFS. | KAON 499 "NGAO System Architecture Definition" | draft | inspection | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Overall | Implementation Requirements | Feature | 3.2.3.2 Optical Relay | Second Stage Relay: A second stage relay with a second deformable mirror located at an optical conjugate to the ground layer. The position of the second stage in the optical train is in series, after the first stage | The second DM is intended to correct the incoming wavefront to an accuracy consistent with the requirements for high Strehl narrow field instruments. | KAON 499 "NGAO System Architecture Definition" | draft | inspection | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Overall | Implementation Requirements | Feature | 3.2.3.2 Optical Relay | LOWFS optical interface: A wide field output from the first relay provides an interface for up to 3 NGS low order wavefront sensors (LOWFS) | A wide field interface is needed so that a large natural guide stars can be selected resulting in high sky coverage | KAON 499 "NGAO System Architecture Definition" | draft | inspection | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Overall | Implementation Requirements | Feature | 3.2.3.2 Optical Relay | Deployable Integral field spectrograph interface: A wide field output from the main relay also directs light into the optical inputs of multiple deployable integral field spectrographs | A wide field interface is needed for this required science instrument | KAON 499 "NGAO System Architecture Definition", Science Instrument requirements | draft | inspection | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Overall | Implementation Requirements | Feature | 3.2.3.2 Optical Relay | High Strehl science instrument interface: A narrow field output from the second relay directs light into one of five instrument stations. The direction of the light into any one of these instrument stations shall not require physically moving any of the instruments. | It is envisioned that there will be an optical “switchyard” with a movable flat mirror that can redirect light into several statically located instruments. Three instruments are show in the plane of figure 5 in KAON 499; the other two locations are symmetrically located above and below the plane of the diagram | KAON 499 "NGAO System Architecture Definition" | draft | inspection | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Overall | Implementation Requirements | Feature | 3.2.3.2 Optical Relay | Acquisition Camera Interface: The need for and location of an acquisition camera is TBD | This sensor is needed to aid in the acquisition of (1) the 3 NGS for tip tilt tracking, (2) science objects and (3) LGS | KAON 499 "NGAO System Architecture Definition" | draft | inspection | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Overall | Implementation Requirements | Feature | 3.2.3.2 Optical Relay | Calibration (Truth) wavefront sensor interface: Location of this sensor is TBD | This sensor is needed to track variation in LGS spot elongation resulting from Na density changes | KAON 499 "NGAO System Architecture Definition" | draft | inspection | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Overall | Implementation Requirements | Feature | 3.2.3.2 Optical Relay | NGS wavefront sensor interface: Location of this sensor is behind both optical relays | This sensor is used for high order NGS adaptive optics observation for testing the system and for science observations | KAON 499 "NGAO System Architecture Definition", Science requirement NGS mode, see: Architecture_reqments_summary_v7.xls | draft | inspection | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | First relay field of view: The first stage of the optical relay will pass a circular unvignetted field of view of 150 arc seconds diameter, without vignetting beams coming from point sources at 80 km altitude. (These fields are referenced to the sky). | This field of view is established by considering the system requirements for wavefront error and sky coverage, while folding in the availability of natural stars for low-order wavefront sensing and assuming the architectural design decision to place the LOWFS downstream of the first relay. Relevant requirements from the SRD are XXX.X, YYY.Y, and ZZZ.Z. The calculations are done using the error budget spreadsheet [reference]. This requirement is stressed by the Extragalactic science cases [reference?]. Wavefront sensors downstream of the fist relay will be using sodium laser guide stars, where light originates from a range of finite altitudes. This requirement is intended to include all anticipated altitudes without cutting off rays from any part of the telescope pupil | See KAON 504 "NGAO Performance vs. Technical Field of View for LOWFS Guide Stars", wavefront error budget spread sheet v1.26 Science Case: Architecture_reqments_summary_v7.xls (row sky coverage) | draft | test | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | The second stage of the relay will pass a circular unvignetted field of view of 20 arc seconds diameter with a goal of 40 arc seconds. (These fields are referenced to the sky) | This requirement is driven by need to image the science target itself along with surrounding astrometric reference stars | This requirement is derived from SRD requirements XXXX. The stressing science case is the imaging of the Galactic center see also: Architecture_reqments_summary_v7.xls | draft | test | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman, Don Gavel | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Static optical quality of first relay: The static optical quality of the first AO relay should be as high as possible over the entire field of view of the AO system. Uncorrectable static aberrations of the AO relay should be no more than 30nm, including chromatic focal shift. | TBD | This requirement is derived from the Performance Budget, Wavefront Error and Encirled Energy spreadsheet. It meets the static aberration requirements of all science cases. | draft | test | 0.3 | 2007-09-11 | 2008-02-01
|  | Chris Neyman, Don Gavel | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Static optical quality of second relay: The static optical quality of the second AO relay should be as high as possible over the entire field of view of the AO system. Uncorrectable static aberrations of the AO first and second relays shall be no more than 30 nm, including chromatic focal shift. | TBD | This requirement is derived from the Performance Budget, Wavefront Error and Encirled Energy spreadsheet. It meets the static aberration requirements of all science cases | draft | test | 0.3 | 2007-09-11 | 2008-02-01
|  | Chris Neyman, Don Gavel | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Lateral color shall be no more than TBD for the first relay. | TBD | TBD | draft | test | 0.3 | 2007-09-11 | 2008-02-01
|  | Renate Kupke | Optical | Performance Requirements | Parametric | 3.2.3.2. Optical Relay | Lateral color shall be no more than TBD for the second relay. | [5 empty] | 2008-02-01 | 2008-02-01
|  | Chris Neyman, Don Gavel | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Chromatic focal shift: Allowable chromatic focal shift is an uncorrectable static aberration. It should be minimized so that in combination with the other static aberrations it does not exceed 30nm. | TBD | TBD | draft | test | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman, Don Gavel | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Curvature of output focal plane, first relay: Sufficient to implement deployable integral field spectrograph and low order wave front sensor object selection mechanisms. | TBD | TBD | draft | test | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman, Don Gavel | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Curvature of output focal plane, second relay: Known, so that instruments fed by the second relay will compensate, if necessary. | TBD | TBD | draft | test | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman, Don Gavel | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Field distortion, first relay: The optical distortion shall be TBD percent across the field of the first relay | TBD | TBD | draft | test | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman, Don Gavel | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Field distortion, second relay: The optical distortion shall be TBD percent across the field of the second relay when used in combination with the first relay. | TBD | TBD | draft | test | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman, Don Gavel | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Transmission, first relay: Transmission in the wavelength range XXXnm to YYYnm [IR Science wavebands] will be TBD %. Transmission in the wavelength range XXXnm to YYYnm [Visible Science wavebands] will be TBD % Transmission at the sodium D2 line (~589nm) will be TBD% Transmission in the wavelength range XXXnm to YYYnm [ NGS wfs sensing bands ] will be TBD % |  |  | draft | test | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman, Don Gavel | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Transmission, second relay: Transmission in the wavelength range XXXnm to YYYnm [IR Science wavebands] will be TBD %. Transmission in the wavelength range XXXnm to YYYnm [Visible Science wavebands] will be TBD % |  |  | draft | test | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman, Don Gavel | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Output focal ratio, first relay: The output focal ratio of the first relay shall be made compatible with the input to the second relay, the inputs to the LOWFS, and the inputs to the units | The two relays have to work together | KAON 499 "NGAO System Architecture Definition" | draft | test | 0.3 | 2007-09-11 | 2007-10-10
|  | Don Gavel | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Output focal ratio, second relay: The output focal ratio of the second relay will be 45.6 | The output focal ratio was chosen as an architectural design decision considering back focal distance, room for the instrument switchyard, reasonable output pupil size and location, and reasonable plate scale for infrared science instruments. This requirement may be revisited as science instrument concepts mature. | Engineering decision Don Gavel (Brian Bauman) | draft | demonstration | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman, Don Gavel | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Output pupil location, entire relay: The optical relays will provide pupil locations that match the Keck telescope when used with its f/15 secondary. The output pupil as seen from the converging beam at the output of the second relay shall be located 19.948 m upstream of the focal plane of the second relay. | The pupil location was chosen as an architectural design decision considering back focal distance, room for the instrument switchyard, reasonable output pupil size and location, and reasonable plate scale for infrared science instruments. This requirement may be revisited as science instrument concepts mature. | Science requirement see (row 9 AO back ground) Architecture_reqments_summary_v7.xls http://www.oir.caltech.edu/twiki_oir/bin/view.cgi/Keck/NGAO/OpticalDesign for exact specification of telescope | draft | demonstration | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman, Don Gavel | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Output pupil location, first relay: The output pupil location of the second relay will be infiinity for the first relay. | The pupil location was chosen to simplify the design of the LGS WFS and LOWFS pick-off mechanisms and wavefront sensors. A telecentric system does not require tilt with field to keep the chief ray constant on entering the wavefront sensor. There are also no pupil scale changes with conjugate distance for the LGS WFS. | Science requirement see (row 9 AO back ground) Architecture_reqments_summary_v7.xls http://www.oir.caltech.edu/twiki_oir/bin/view.cgi/Keck/NGAO/OpticalDesign for exact specification of telescope | draft | demonstration | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman, Don Gavel | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Pupil image size internal to the first relay: The pupil image size within the first relay will be 100 mm | This pupil size is driven by the choice of deformable mirror for the first relay. The requirement is to accommodate a DM having 21 actuators across the diameter of the pupil and we are assuming that the actuator pitch is 5mm. | KAON 499 "NGAO System Architecture Definition" | draft | demonstration | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman, Don Gavel | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Pupil image size internal to the second relay: The pupil image size of the second relay will be 25 mm. | This pupil size is driven by the choice of deformable mirror for the second relay. The requirement is to accommodate a DM having 64 actuators across the diameter of the pupil and we are assuming that the actuator pitch is 400 microns | KAON 499 "NGAO System Architecture Definition" | draft | demonstration | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman, Don Gavel | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Pupil grid distortion on woofer first relay, TBD | Requirement is based on tenth of a subaperture registration error and 64 subapertures across the pupil. | Standard rule of thumb for AO systems (Chris Neyman engineering decision) | draft | demonstration | 0.3 | 2007-09-11 | 2007-10-10
|  | Renate Kupke | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Pupil grid distortion on tweeter, second relay TBD | The telescope primary will not map perfectly to the DM due to optical distortions. |  |  |  |  | 2008-02-01
|  | Renate Kupke | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Pupil aberration, woofer mirror, first relay TBD | The telescope will map to the woofer mirror with a point spread function whose width depends on the field being corrected, in this case 150 arcseconds. This point spread function blurs the correction applied to the DM. |  |  |  |  | 2008-02-01
|  | Renate Kupke | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Pupil aberration, tweeter mirror, second relay TBD | The telescope will map to the tweeter mirror with a point spread function whose width depends on the field being corrected, in this case 30 arcseconds. |  |  |  |  | 2008-02-01
|  | Chris Neyman, Don Gavel | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Output pupil distortion, entire relay: The combined first and second relay shall provide less than TBD% pupil distortion in the output beam. | This requirement will be driven by requirements or good practice for the instruments. For example it may take into account the feasible design of cold stops and other optics in the infrared instruments. | Wavefront error budget spread sheet version 1.26, dm-lenslet registration | draft | demonstration | 0.3 | 2007-09-11 | 2007-10-10
|  | Renate Kupke | Optics | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Pupil tilt, woofer mirror, first relay | There will be a large tilt of the pupil at the plane of the woofer, due to the nature of producing the pupil with an off-axis parabola. This tilt will act like an uncertainty in the conjugate height being corrected. |  |  |  |  | 2008-02-01
|  | Renate Kupke | Optics | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Pupil tilt, tweeter mirror, second relay | There will be a large tilt of the pupil at the plane of the woofer, due to the nature of producing the pupil with an off-axis parabola. This tilt will act like an uncertainty in the conjugate height being corrected. |  |  |  |  | 2008-02-01
|  | Chris Neyman, Don Gavel | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | Output pupil tilt, entire relay: | The combined first and second relay shall provide an output pupil that is tilted less than TBD degrees off the chief ray normal. | Wavefront error budget spread sheet version 1.26, dm-lenslet registration | draft | demonstration | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Optical | Performance Requirements | Operational | 3.2.3.2 Optical Relay | Athermalization, performance at operating temperature: The optical relay shall be able to satisfy its optical requirements at the operating temperature specified for the cooled enclosure, allowing for a variation of plus or minus xxx C. See requirement Cooling( row 9 this spread sheet) for the operating temperature | Need to be able to operate the AO system when it is not at its operational temperature. | Engineering decision by Chris Neyman | draft | test | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Optical | Performance Requirements | Operational | 3.2.3.2 Optical Relay | Athermalization, performance at dome ambient temperature: both cooled (-20 C) and ambient temperatures (0 C). | Need to be able to operate the AO system when it is not at its operational temperature. Specifically when the enclosure is at the ambient temperature of the Keck domes (0 C) | Engineering decision by Chris Neyman | draft | test | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Optical | Implementation Requirements | Feature | 3.2.3.2 Optical Relay | Alignment: As a goal the alignment of the optical relay should be maintained whenever possible by accurate machining of static fixtures. | At the AO system should have few adjustments as necessary, the bulk of the optical alignment tolerances should be maintained by static fixtures. | Engineering decision by Chris Neyman | draft | demonstration | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Optical | Performance Requirements | Parametric | 3.2.3.2 Optical Relay | LGS background light in other sensors: Filters, baffles, and selection of dichroics shall be made so as to limit the background light from laser guide stars from entering into other detectors. The background shall be less than TBD mag/arcsec2 at the LOWFS and less than TBD mag/arcsec2 at the visible science instruments. | Increasing in the background above the natural limit set by the night sky, will degrade tracking and instrument performance. It is possible that the LOWFS detector will be insensitive to LGS light to such an extinct that these measures are not needed, this analysis is TBD. | KAON 499 "NGAO System Architecture Definition" | draft | test | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Mechanical | Performance Requirements | Operational | 3.2.3.2 Optical Relay | No mechanism vignetting of optical beam: Mechanical systems inside and around the AO enclosure will not obscure the optical beam from a 202 arc second transferred field of view from the first relay and a 40 arc second transferred field from the second relay. | Needed to meet field of view optical requirement | See rows 21 and 22 this spreadsheet | draft | test | 0.3 | 2007-09-11 | 2007-10-10 | Circular reface in requirement
|  | Chris Neyman | Optical | Implementation Requirements | Operational | 3.2.3.3 Optical Switchyard | Optical interface to multi-object deployable near-IR IFU and LOWFS: There shall be an optically flat dichroic splitter that feeds appropriate bands of infrared light to multi-object deployable near-IR IFU and the LOWFS. Dichroic choices are: • Mirror (sends all light to multi-object deployable near-IR IFU and LOWFS) • J band only • H band only • Both J and H bands The dichroic shall maintain optical quality in order to meet wavefront specifications (this spread sheet) along both paths, transmitted and reflected. | The splitters are need for operation of the LOWFS at the same time as near IR high Strehl instruments that are located behind the second relay. | KAON 499 "NGAO System Architecture Definition" | draft | test | 0.3 | 2007-09-11 | 2007-10-10 | Circular reference in requirement
|  | Chris Neyman | Optical | Implementation Requirements | Operational | 3.2.3.3 Optical Switchyard | Optical interface to LGS wavefront sensors: There shall be a sodium D2a line reflecting dichroic that sends LGS light to the wavefront sensors. The dichroic shall maintain optical quality in order to meet wavefront specifications 7.2.2.3 along both paths, transmitted and reflected. The dichroic shall pass wavelength range 0.6 to 2.2 microns. | Need to use LGS in combination with other NGAO optical subsystems | KAON 499 "NGAO System Architecture Definition" | draft | test | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Optical | Implementation Requirements | Operational | 3.2.3.3 Optical Switchyard | Optical interface to acquisition camera is TBD | Need to use acquisition camera in combination with other NGAO optical subsystems | KAON 499 "NGAO System Architecture Definition" | draft | inspection | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Optical | Implementation Requirements | Operational | 3.2.3.3 Optical Switchyard | Optical interface to truth wavefront sensor is TBD | Need to use truth wavefront sensor in combination with other NGAO optical subsystems | KAON 499 "NGAO System Architecture Definition" | draft | inspection | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Optical | Implementation Requirements | Operational | 3.2.3.3 Optical Switchyard | Optical interface to narrow field science instruments. A optically flat mirror shall be used to selected the instrument into which the light from the second optical relay shall be directed. The flat shall maintain optical quality in order to meet wavefront specifications 7.2.2.4. | Need to use narrow field instruments in combination with other NGAO optical subsystems | KAON 499 "NGAO System Architecture Definition" | draft | inspection | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Mechanical | Implementation Requirements | Feature | 3.2.3.3 Optical Switchyard | 7.2.3.7 Motion control: 1.The input K mirror must provide image or pupil rotation compensation, which is selectable by the AO control software. 2. A moving optical element will select between the five instruments located at the output of the second stage relay. 3. The atmospheric dispersion compensator will be removable from the second relay. 4. The atmospheric dispersion compensator will be driven based on the telescope elevation and the direction of the Keck zenith on the AO bench. 5. The LGS dichroic will be removable. 6. The LOWFS and multi-object deployable near-IR IFU pickoff will be selectable between one of several optical elements in order to optimize LOWFS performance and scientific instrument throughput. | These motion control options were specified in KAON 499, others are TBD. | KAON 499 "NGAO System Architecture Definition" | draft | inspection | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman | Mechanical | Performance Requirements | Parametric | 3.2.3.3 Optical Switchyard | Mechanism motions: Optical relay Speed of mechanism motions is TBD Accuracy of the mechanism motions is TBD | These motion control options were specified in KAON 499, others are TBD. | KAON 499 "NGAO System Architecture Definition" | draft | test | 0.3 | 2007-09-11 | 2007-10-10
|  | Chris Neyman, Don Gavel | Optical | Performance Requirements | Parametric | 3.2.3.8 Atmospheric Dispersion Correction | Atmospheric dispersion compensation: Optical relay will include atmospheric dispersion compensation for the output of the 2nd optical relay. Atmospheric dispersion shall be controlled to 0.177 milli arc seconds across J band, 0.144 milli arc seconds across H band, and 0.077 milli arc seconds across K band (or specify particular bands in nm). | This is driven by science requirements for imaging quality over given optical bandwidths for instruments in the narrow field science output path. | Reference to system requirement high Strehl and tip tilt residual error. Parameters are from wavefront error budget spread sheet v.126 | draft | demonstration | 0.3 | 2007-09-11 | 2007-10-10

## AdkinsTaxonomy
See Adkins, "Draft Engineering Guideline For the Preparation of Requirements Documents" version 1.1, September 2006
Adkins requirements taxonomy
Document Sections (roughly by engineering discipline) | Categories | Sub Categories
Performance Requirements
Overall |  | Parametric Performance Requirements
Optical |  | Operational Performance Requirements
Mechanical | Implementation Requirements
Electronic/Electrical |  | Feature Implementation Requirements
Safety |  | Common Practices Implementation Requirements
Software |  | Standards Implementation Requirements
Interface |  | Regulatory Implementation Requirements
Reliability | Design requirements
Spares |  | Technological Design Requirements
Service and Maintenance |  | Regulatory Design Requirements
Documentation |  | Standards Related Design Requirements
Integration Related Design Requirements

## RequirementsOutline
| Outline of requirements document | See Adkins, "Draft Engineering Guideline For the Preparation of Requirements Documents" version 1.1, September 2006
1 | Introduction
2 | Scope and Applicability
3 | Related Documents
4 | Revision History
5 | Background
6 | Overall System Requirements
Performance Requirements
Implementation Requirements
Design Requirements
7 | Optical Requirements
Performance Requirements
Implementation Requirements
Design Requirements
8 | Mechanical Requirements
Performance Requirements
Implementation Requirements
Design Requirements
9 | Electronic/Electrical Requirements
Performance Requirements
Implementation Requirements
Design Requirements
10 | Safety Requirements
Performance Requirements
Implementation Requirements
Design Requirements
11 | Software Requirements
12 | Interface Requirements
Performance Requirements
Implementation Requirements
Design Requirements
13 | Reliability Requirements
14 | Spares Requirements
15 | Service and Maintenance Requirements
16 | Documentation Requirements
Glossary

## DataBaseFields (hidden)
[6 empty] | WBS | WBS (Sys. Design Plan)
data base field names | ID # | Who nominated or originated requirement | Section | Category | Sub Category |  | L.1 | L.2 | L.3 | L.4 | Description | Rationale | Traceability | Status | Verification method | Version | Created | Modified
data type |  | String | Selection (string) | Selection (string) | Selection (string) | Selection (string) | int | int | int | int | text | text | text | selection | selection | string | date | date
comments | Just a unique number in the database |  | Assign to Sean's requirements document sections: should make it easier to draft requirements documents from database (see taxonomy, outline) |  |  | What part of the system is the requirement assigned to (by WBS) | What part of the system is the requirement assigned to (by WBS) |  |  |  | Text of the requirement | Why requirement is needed, plus technical reference, KAONS, other requirements documents, etc. | What requirements this one is derived from: should be a reference to other data bases or design document |  | How the requirement is tested | What version | Date Created | history of modification
Selections |  |  | pick one: | pick one: | pick one: | pick one: | [7 empty] | pick one: | pick one:
|  |  | Overall | Performance Requirements | if performance: | [8 empty] | draft | inspection
|  |  | Optical |  | Parametric | [8 empty] | final | analysis
|  |  | Mechanical |  | Operational | [8 empty] | approved | demonstration
|  |  | Electronic/Electrical | Implementation Requirements | if implementation: | [8 empty] | pending approval | test
|  |  | Safety |  | Feature | [8 empty] | disapprove
|  |  | Software |  | Common Practices
|  |  | Interface |  | Standards
|  |  | Reliability |  | Regulatory
|  |  | Spares | Design requirements | if design:
|  |  | Service and Maintenance |  | Technological
|  |  | Documentation |  | Regulatory
Standards
Integration

## pickfields (hidden)
[6 empty] | WBS | WBS (Sys. Design Plan)
data base field names | ID # | Who nominated or originated requirement | Section | Category | Sub Category |  | L.1 | L.2 | L.3 | L.4 | Status | Verification method
data type |  | String | Selection (string) | Selection (string) | Selection (string) | Selection | int | int | int | int | selection | selection
comments | Just a unique number in the database |  | Assign to Sean's requirements document sections: should make it easier to draft requirements documents from database (see taxonomy, outline) |  |  |  | What part of the system is the requirement assigned to (by WBS) |  |  |  |  | How the requirement is tested
Selections |  |  | pick one: | pick one: | pick one: | pick one |  |  |  |  | pick one: | pick one:
|  |  | Overall | Performance Requirements | if performance: | 3.2 AO System |  |  |  |  | draft | inspection
|  |  | Optical | Implementation Requirements | Parametric | 3.2.1 AO System Architecture |  |  |  |  | final | analysis
|  |  | Mechanical | Design requirements | Operational | 3.2.2 AO Enclosure |  |  |  |  | approved | demonstration
|  |  | Electronic/Electrical |  | if implementation: | 3.2.3 Opto-mechanical |  |  |  |  | pending approval | test
|  |  | Safety |  | Feature | 3.2.3.1 Field Rotation |  |  |  |  | disapprove
|  |  | Software |  | Common Practices | 3.2.3.2 Optical Relay
|  |  | Interface |  | Standards | 3.2.3.3 Optical Switchyard
|  |  | Reliability |  | Regulatory | 3.2.3.4 Optical Support Structure
|  |  | Spares |  | if design: | 3.2.3.5 Wavefront Sensors
|  |  | Service and Maintenance |  | Technological | 3.2.3.5.1 High Order LGS Wavefront Sensors
|  |  | Documentation |  | Regulatory | 3.2.3.5.2 High Order NGS Wavefront Sensor
[5 empty] | Standards | 3.2.3.5.3 Low Order NGS Wavefront Sensors
[5 empty] | Integration | 3.2.3.5.4 Calibration Wavefront Sensor
3.2.3.6 Wavefront Correctors
3.2.3.6.1 Tip/Tilt Corrector
3.2.3.6.2 Deformable Mirror
3.2.3.6.3 Tip/tilt Vibration Mitigation
3.2.3.7 Acquisition Cameras
3.2.3.7.1 NGS Acquisition Camera
3.2.3.7.2 LGS Acquisition Camera
3.2.3.8 Atmospheric Dispersion Correction
3.2.3.9 Alignment, Calibration, Diagnostics, Metrology and Monitoring
3.2.3.10 Atmospheric Profiler
3.2.3.11 d-IFU & LOWFS AO & Object Selection
3.2.3.12 PSF Monitoring Camera
3.2.4 Non-real-time Control
3.2.4.1 Non-RTC Software
3.2.4.2 Non-RTC Electronics
3.2.5 Real-time Control
3.2.5.1 RTC Architecture Analysis and Design Study
3.2.5.2 RTC Software Module Definition
3.2.5.3 RTC Hardware Module Definition
3.3 Laser Facility
3.3.1 Laser System Architecture
3.3.2 Laser Enclosure
3.3.3 Laser
3.3.4 Laser Launch Facility
3.3.4.1 Laser Beam Transport
3.3.4.2 Laser Pointing & Diagnostics
3.3.4.3 Laser Launch Telescope
3.3.5 Laser Safety Systems
3.3.5.1 Personnel and Equipment Safety Systems
3.3.5.2 Aircraft, Satellite & Laser Traffic Control Safety Systems
3.3.6 Laser System Control
3.3.6.1 Laser System Software
3.3.6.2 Laser System Electronics