Next Generation Adaptive Optics System

Laser Launch Facility Beam Transport Optics

Preliminary Design

(Draft)

Sept 3, 2009

VersionV1.1

Prepared By Author

# Revision History

| Revision | Date | Author (s) | Reason for revision / remarks |
| --- | --- | --- | --- |
| 1.0 | Aug 24, 2009 |  | Initial release |
| 1.1 | Sep 03, 2009 |  | Mods to include BTO figures. JC |

# TABLE OF CONTENTS

Revision History	2

TABLE OF CONTENTS	3

1	Introduction	5

2	References	7

2.1	Referenced Documents	7

2.2	Acronyms and Abbreviations	7

3	Overview	8

4	Requirements	8

5	Design	9

5.1	Opto-Mechanical Design	9

5.1.1	Optical Mechanical Layout	9

5.1.2	Flexure	10

5.1.3	Keck II Laser Pointing Model and Data	12

5.1.4	Keck I Flexure Testing	12

5.2	Mechanical Design	12

5.2.1	Tube Design	12

5.3	Scalability	12

5.4	Zemax Optical Model (Thomas)	13

5.5	Error Budget and Tolerances (Thomas)	13

5.6	Electrical Design (Ed)	13

5.6.1	Motion Control System	13

5.7	Diagnostics (Ed)	13

5.8	Safety (Ed)	13

5.9	Interfaces	14

5.9.1	External Interfaces	14

5.9.2	Internal Interfaces within the LLF	14

6	System Performance	14

6.1	Optical (Thomas)	14

6.1.1	Transmission	14

6.1.2	Wavefront Error	14

6.1.3	Pointing Errors	14

6.2	Mechanical (Jim)	14

6.2.1	Mass on Telescope and Impacts	14

6.2.2	Heat Dissipation and Glycol requirements	14

6.3	Electrical (Ed)	14

7	Operations (Thomas)	14

7.1	Modes	14

7.1.1	Operational Mode	14

7.1.2	Initialization Mode	15

7.1.3	Alignment Mode	15

7.2	Procedures	15

7.2.1	Alignment	15

7.2.2	Cleaning (Maintenance)	15

7.2.3	Motion Stage Lubrication (Preventative Maintenance)	15

7.3	Operational Resources and Preventative Maintenance	15

7.4	Configuration Management	15

8	Development and Testing (Thomas)	16

9	Requirements Compliance Verification	17

10	Risk and Risk Reduction Plan (Thomas)	22

11	Deliverables	22

12	Management (Jason)	23

12.1	Budget	23

12.2	Budget Reduction Possibilities	23

12.3	Schedule	23

13	Plans for the next Phase (Jason)	23

14	Appendix A.  Requirements	25

# Introduction

As part of the Next Generation Adaptive Optics System (NGAO), a Laser Launch Facility (LLF) System is needed to propagate the laser beam.  One component of the LLF System is the Beam Generation System (BGS).  The BGS is located within the secondary f/15 module on the telescope.  It receives the laser beam(s) from the Beam Transport Optics (BTO), formats them into the required asterism, and provides the beam pointing on the sky.

This document provides the design for the Laser Launch System BGS in support of the NGAO Preliminary Design Phase.

Things to keep in mind during this phase.

Taken from Meeting 7 PD phase:

- Objectives.

  - Deliver documented designs for each system, sub-system & component, hardware or software, of sufficient detail to establish through inspection & analysis the feasibility of the proposed design, & the likelihood that the design will meet the requirements.

  - Present the project plan to completion, including a detailed schedule & budget.

- Principal activities

  - Design, prototyping, simulation and analysis.

- Key deliverables

  - Preliminary technical specifications, requirements for subsystems, a preliminary Operations Concept Document, Interface Design document(s), & a Preliminary Design report.

- Observing Operations Concept Document

- System Requirements Document

- Functional Requirements & Interface Control Document(s)

  - AO system, laser system, science operations tools, science instruments

  - Managed within Contour database

- Preliminary Design Manual (the document to read to understand the design & performance of the NGAO facility – draft PDM posted)

  - Flowdown of requirements to design

  - Solidworks & Zemax model(s)

  - Software design (RTC, non-RTC & science ops tools)

  - Performance Budget Reports (wavefront, EE, astrometry, contrast, …)

  - Science Performance Analysis Report

- Science Instrument Design Manual

- Risk Assessment & Management Report

- Systems Engineering Management Plan

  - Project plan to completion, including a detailed schedule & budget

  - Phased implementation option(s)

  - Cost estimation

  - Justification for any procurements during DD

From the WBS Definition

| Phase | WBS Element | Deliverable |
| --- | --- | --- |
| PD | Develop a preliminary design for the systems required for delivering the laser power from the laser to the sky. This includes: 1) Laser Beam Transport: Develop preliminary design for delivering the laser power from the laser to the launch telescope. 2) Laser Pointing and Diagnostics: Develop preliminary design for determining and controlling the alignment and pointing of the laser beams. Develop preliminary design for regularly monitoring the beam quality, laser power, and health of the laser launch system. 3) Laser Launch Telescope: Develop the preliminary design for the telescope needed to launch multiple laser beacons Does not include: Software control of these systems is part of laser system control (WBS 5.5) Software control of safety shutters and interlocks is part of laser safety system (WBS 5.4). | 1. Preliminary optical design for optics located in laser enclosure (beam transport, laser pointing, diagnostics) including ray trace design, preliminary tolerances, and preliminary alignment plan 2. Preliminary optical design for launch telescope 3. Preliminary optical design for optics located along telescope tube and behind secondary (beam transport, laser pointing, and diagnostics) 4. Mechanical design for mechanical elements located in laser enclosure (beam transport, laser pointing, and diagnostics) including mechanical drawings, mechanism for motion control, optic mounts 5. Mechanical design for laser launch telescope and mechanical supports 6. Mechanical design for mechanical elements located along telescope tube and behind secondary 7. Electrical design for beam transport, laser pointing, diagnostics and launch telescope including electrical system for motion control, monitoring 8. Interfaces (internal to NGAO) 9. Assembly, alignment, and test plans 10. Verify compliance, update requirements matrices and updated requirements (preliminary specifications) 11. Document design. |

# References

## Referenced Documents

Documents referenced are listed in Table 1.  Copies of these documents may be obtained from the source listed in the table.

| Ref. # | Document # | Revision or Effective Date | Source | Title |
| --- | --- | --- | --- | --- |
| 1 | KAON 511 | 0.3 | WMKO | NGAO System Design Manual |
| 2 | KAON 525 | 1.0 | WMKO | K1 LGSAO Safety System Preliminary Design Report |
| 3 | KAON 562 | 1.0 | WMKO | NGAO Design Changes |
| 4 | KAON 574 | 1.0 | WMKO | NGAO Systems Engineering Management Plan |
| 5 | KAON 642 | April 10, 2009 | WMKO | NGAO Design Changes in Support of Build-to-Cost Guidelines |
| 6 | KAON 572 | 0.1 | WMKO | Instrument Baseline Requirements Document |

Table 1:  Reference Document

## Acronyms and Abbreviations

Table 2 defines the acronyms and abbreviations used in this document.

| Acronym/Abbreviation | Definition |
| --- | --- |
| K1 | Keck 1 |
| K2 | Keck 2 |
| KAON | Keck Adaptive Optics Note |
| LGS | Laser Guide Star |
| LLF | Laser Launch Facility |
| NGAO | Next Generation Adaptive Optics System |
| NGS | Natural Guide Star |
| WMKO | W.M.K. Observatory |

Table 2:  Acronyms and Abbreviations

# Overview

The LLF layout is shown in Figure 1 which was presented in the NGAO System Design Review in KAON 511. The Beam Transport Optics (BTO) includes the functions represented by the components that are not in the dotted rectangles.  The main function of the BTO is to transport the beam from the laser enclosure, along the telescope structure, to the Beam Generation System (BGS).

Figure 1:  Laser Launch Facility Layout

The location of where the BTO fits into the overall NGAO System is shown in Figure 2.  The BTO will have a mechanical interface to the mainly to the telescope and partially in the f/15 module.

Figure 2:  Laser Launch Facility BTO (shaded in red) within the NGAO System

# Requirements

The requirements for the Beam Transport Optical System are presented in Appendix A.  The BTO is part of the BTOS and will apply the requirements as outlined by the BTOS.

In addition to those in the BTOS requirement, the following requirements are being added to BTO.

| Short Name | ID | Section | Category | Priority | Description |
| --- | --- | --- | --- | --- | --- |
| BTO Mounting to telescope | FR- | Overall | Implementation | Important | The mounting of the BTO to the telescope structure will be compliant as not to impact the telescope performance. |
| BTO vignetting | FR- | Overall | Implementation | Important | The BTO will not add any additional vignetting to the telescope field of view. The BTO will travel along existing telescope structures and will not protrude visually into the telescopes field of view. |
| BTO exterior surface | FR- | Overall | Implementation | Important | The BTO’s exterior surface shall be properly coated to minimize emissivity and reflections. |
| BTO Scalability | FR- | Overall | Optical | Important | The BTO optical design shall be able to be used with the existing Keck II dye laser output. |

Table 3:  Additional Requirements for the BTO

# Design

## Opto-Mechanical Design

The main criteria for the BTO design are to maximize throughput and minimize wavefront errors by minimizing reflections, no additional vignetting of the primary telescope, while maintaining polarization of the beams.  The design will also be scalable to add additional beams if necessary.  The BTO receives the beams from the Switchyard located in the laser enclosure on the elevation ring and outputs the beams to the Beam Generation System (BGS) in the secondary module.

### Optical Mechanical Layout

Figure 3 shows the path of the BTO from a side view of the primary telescope.  The right side of figure (-Y) is where the Laser Service Enclosure (LSE) and Switchyard will be located.  The beams will travel in a sealed tube from the LSE at a diagonal toward the top end of the telescope in the -Y direction.  This tube will run outside of the pupil without vignetting the primary as shown in Figure 4.  One reason for implementation of the tubes in the +Y direction of the telescope is for the BTO to travel on the outside (+Z) of the existing telescope spiders that emanate from the top socket.  This is crucial for the beam to go directly into the BGS without additional fold mirrors.  An additional reason for the BTO to travel in the +Y half of the telescope is that tubes located on the outside of the spiders (+Z) will not collide with the Nasmyth Deck.  Tubes extending beyond the current telescope structure in the +Z and –Y direction can collide with the Nasmyth Deck as the telescope comes down to 70 degrees zenith angle.

This design has a total of six mirrors, two at the laser enclosure, two on the telescope structure, and two at the secondary socket to fold the beam into the BGS.  Motion control

+Z

+Y

Figure 3:  BTO on telescope (side view)

+Y

+X

Figure 4:  BTO on telescope (front view)

### Flexure

Understanding the telescope flexure is crucial to determine the needed motion control to assure the laser beams will enter the BGS properly at the top end.  Although the telescope was designed with trusses that function as a Serrurier type telescope, the flexure impacting the laser beams will be different. The lasers will be located on the elevation ring while its initial target will be located at a corner of the top ring.  The Serrurier design assumes the primary and secondary mirrors move together due to gravity.

#### Blue Book Design

Based on the blue book design of the telescope, the top ring having a 2 ton instrument package will deflect 1.7mm with gravity, resulting in a 19 arcsec tilt about the elevation axis or 1.2mm of motion at the top end.  No data was provided for the corner of the top ring where the laser beams will land as they transverse toward the secondary from the elevation ring.  The total displacement is 2.9mm?? or 19 arcsec.

Using this blue data and a Zemax model of the beam path from the laser enclosure to the secondary socket, the resultant deflection will results in 3mm of motion rotated 19 arcsec.  The position error would mainly in the Y-axis at 2.17 mm and the resulting 0.67mm in the x axis.

To compensate for this motion, a single steering mirror can be placed at the laser enclosure location.  The maximum tilt required would be the following:

| Axis | Tilt Required (arcsec) | Resulting Displacement (mm) |
| --- | --- | --- |
| X | 18.4 | 1.5 |
| Y | 0.026 | 1.2 |

The resultant change in angle at the secondary socket is 39.5 arcsec.  The resultant angular motion is shown in Figure 5.  This figure shows the angular motion at both the top ring and the spider.

Beam Displacement at top ring

Beam Displacement at spider

No Flexure    Uncorrected    Corrected

Figure 5:  Angular Motion

This 40 arcsec angular change will result in a 0.24 arcsec change on sky with the LTA/BTOB magnification factor.  Adding a 0.145 mm centroid offset can easily account for this error.

Assuming a motion device located at the LSE to provide 1 urad of resolution on a gimbal mount with a lever arm of 35mm.  The resultant 100urad range on this surface would provide the needed correction for the BTO into the BGS without additional motion devices.  The 35nm of incremental motion from this actuator would be sufficient to provide 25 um of motion at the secondary socket.  For such fine motion, a piezoelectric device is the probably required.

Until the blue book flexure model is verified, a conservative approach is to provide contingency for an additional motion stage at top ring location that steers the beam toward the secondary socket.  Using a similar device type as that suggested at the LSE, the required resolution will be larger by a factor of 5 to 5urad or 175nm of travel.  This larger resolution is the resultant of a shorter lever arm to the secondary.  By adding a mirror at this location, the resolution of the device at the LSE is also relaxed to 1.5urad or 50nm of motion.  Although this resolution is relaxed, it is still marginal for using standard actuators instead of piezoelectric type.

### Keck II Laser Pointing Model and Data

Using Keck II laser pointing model, the laser pointing has a worst case offset of 12mm of motion.  This takes into account of all factors of which telescope flexure is one component.    Other factors include laser internal flexure, residual sodium distance, secondary tilt, telescope pointing models, and acquisition camera offsets of which are all lumped together into the model.

A test was conducted in Keck II to understand this same flexure prior to installation of the laser on the elevation ring.  Prior to laser installation on the Keck II elevation ring, the internal structure of the elevation ring was stiffened near the laser attachment points.  KAON xxx describes the result of this test which concluded a motion of 12mm between the elevation ring and the L4 laser exit lens.  The L4 mechanical support, located on a corner of the top ring at a location 116 degrees about the + X axis, is a cantilever design.  The large 20 cm L4 optic created a significant tilt which may have skewed the data.

### Keck I Flexure Testing

Due to the inconsistency of the data, a new flexure test was completed with the following results:

These results conclude

## Mechanical Design

A mechanical unit sized xxx will be located at the start of the BTO at the laser enclosure.  This unit will be attached to the LSE by means of …

### Tube Design

The BTO tubes will be of different sizes depending on its location.  Those residing in front (+Z) of the spiders will be of n” diameter and cannot extend beyond the thickness of the spiders.  Tubes not in the telescope aperture may be larger for structural support.  The tube will be sealed from the LSE to the BGS with a modest n CFM of purge air starting at the LSE to maintain cleanliness within the tube structure.  The exterior of the tube shall be painted with low emissivity paint.

## Scalability

The current plan is to install three 25W lasers in the existing Keck II LSE.  If this is unattainable, an auxiliary location may be needed to host one of the three lasers.  This laser can be placed at the Right Bent Cassegrain location.  The output beam from this third laser can go directly in the +Z direction of the telescope to join the existing BTO path.  This design will allow the additional beams if necessary without significant changes to the design.

## Zemax Optical Model (Thomas)

## Error Budget and Tolerances (Thomas)

## Electrical Design (Ed)

### Motion Control System

Describe what motion controls are necessary and to what level of accuracy and repeatability.  Assume the controls team to provide the needed motion control.  Specify motion parameters including motor requirements.

Describe the control system that will be implemented.  This can be included in the controls section; but should be mentioned here.

## Diagnostics (Ed)

What diagnostics will be available?  Will there be a camera at the entrance of the BGS or output of the BTO to examine beam quality?  Power?

## Safety (Ed)

What are the safety concerns and how are they mitigated?  Shutter control shall be tied to?  Any emergency stops?  Will the laser path be completely sealed?  Will the PSD be used for safety interlocking?

## Interfaces

### External Interfaces

#### Mechanical Interface to the f/15 module (Jim/Ed)

#### Mechanical Interfaces to the telescope structure including the elevation ring (Jim)

#### Infrastructure Interfaces such as Power, Pneumatic and Glycol (Ed)

### Internal Interfaces within the LLF

#### Mechanical and Optical Interface to the Beam Generation System (Mechanical) (Jim) (Optical) (Thomas)

#### Mechanical and Optical Interface to the Switchyard (Mechanical) (Jim) (Optical) (Thomas)

#### Electrical Interface to the LGS Control System  (Ed)

#### Electrical Interface to the Safety System (Ed)

# System Performance

## Optical (Thomas)

### Transmission

### Wavefront Error

### Pointing Errors

## Mechanical (Jim)

### Mass on Telescope and Impacts

### Heat Dissipation and Glycol requirements

## Electrical (Ed)

Power needed

# Operations (Thomas)

## Modes

### Operational Mode

During normal operations, the BTO shall operate at full laser power with three beams operating at 25 watts of CW 589nm light.  The BTO, along with the entire LLF, will maintain the beam to within 0.3 arc sec tolerance.  The BTO will use its position sensing diodes, as well as inputs from the Adaptive Optics System to maintain the laser returns properly on sky as well as the wavefront sensors.

### Initialization Mode

During the initialization mode, the operator will initialize all motion control devices in the BTO to a position referenced by the fiducial location.  This location may be nominal with respect to the zenith angle of the telescope.  The laser beams will be blocked and prevented from entering the BTO during the initialization mode.

### Alignment Mode

During alignment mode, the laser beams will be operated at low power (TBD) initially.  The power shall be sufficient to operate the PSD to properly center the laser beams at the BGS entrance pupil within the tolerances.  Remote adjustments shall be made to position the BTO mirrors based on PSD inputs.  Once alignment is verified, the laser power will be ramped back to operational powers while maintaining PSD positioning.

## Procedures

The following procedures shall be provided as part of future phases for the BTO.

### Alignment

A rough alignment procedure shall be provided to set up the BTO on the telescope.  This rough alignment shall not include the use of a 589nm laser.  This procedure shall be provided during DDR.

A fine alignment procedure shall be provided to operate the BTO in alignment mode as specified in 7.1.3.  This procedure shall take into consideration any safety issues due to the high power nature of the laser.  A final procedure shall be turned over to the operations team during BTO handover.

### Cleaning (Maintenance)

The BTO shall be designed for a 10 year life time.  Due to difficult access to the BTO devices on the telescope, the BTO shall be designed as a sealed unit to minimize particulates contamination.  However, considerations shall be made in the design to clean the optics in the BTO and possibly recoat as necessary.  The frequency of cleaning shall depend on how well the BTO is sealed to prevent particulates contamination.  This procedure shall be turned over to the operations team during BTO handover.  BTO cleaning will require both personnel and telescope resources to access BTO optics.

### Motion Stage Lubrication (Preventative Maintenance)

Since the BTO shall include a number of motion stages, these stages shall be lubricated per the manufacturer.  The frequency shall be based on manufacturer recommendation as well as how well the BTO is sealed.  BTO stage lubrication will require both personnel and telescope resources to access BTO stages.

## Operational Resources and Preventative Maintenance

Additional information shall be provided during DDR on required operational resources.  Outside of troubleshooting of the BTO, Operations team will be required to support maintenance procedures in 7.2.

## Configuration Management

Any Configurations management issues (usually s/w)

# Development and Testing (Thomas)

The following provides a methodology on how the BTO system will be developed and tested.  More detail plans shall be provided in the DDR

# Requirements Compliance Verification

The following table shows the requirements compliance.

| Short Name | ID | Compliance |
| --- | --- | --- |
| Beam Transport System - definition (linked) | FR-1969 | By Design |
| Standards - new instruments and facilities (linked) | FR-1970 | By Design |
| Central projection of Laser Guide Stars (linked) | FR-1971 | By Design |
| Reuse Keck I or Keck II Laser Launch Telescope (linked) | FR-1972 | N/A |
| Elevation range (linked) | FR-1973 | By Design and Test |
| Interface to LGS Control System - software | FR-1974 | Will complied with once LGS Control System is designated and Test |
| Interface to LGS Control System - electrical | FR-1975 | By Design to be flexible to interface with LGS Control System and Test |
| Reimage Laser Unit pupil to Laser Launch Telescope | FR-1976 | N/A |
| Input beam format | FR-1977 | By Design and Test |
| Output beam format | FR-1978 | By Design and Test |
| Laser Launch Telescope - functional quality | FR-1979 | N/A |
| LGS Focus control | FR-1980 | N/A |
| Transmission (linked) | FR-1982 | By Design and Test |
| Damage threshold | FR-1983 | By Design |
| Automated alignment | FR-1984 | By Design |
| Beam splitting | FR-1985 | N/A |
| Asterism generation | FR-1986 | N/A |
| Deployable LGS control | FR-1987 | N/A |
| Offload mirrors - function | FR-1988 | N/A |
| Offload mirrors - update rate | FR-1989 | By Design |
| Offload mirrors - range and precision | FR-1990 | N/A |
| Image natural stars with Laser Launch telescope | FR-1991 | N/A |
| Quarter wave plate | FR-1992 | N/A |
| Quarter wave plate control | FR-1993 | N/A |
| Laser safety - general requirement | FR-1994 | By Design |
| Laser safety - Interior finish | FR-1995 | By Design |
| Laser safety - E-Stop | FR-1996 | N/A |
| Laser safety - Status Indicator | FR-1997 | N/A |
| Laser safety: laser radiation exposure | FR-1998 | By Design |
| Laser safety: hazard labelling and warning signs | FR-1999 | By Design |
| Pointing reference | FR-2000 | By Design |
| Shutter | FR-2001 | N/A |
| Environmental monitoring | FR-2002 | By Design |
| Yield strength | FR-2003 | By Design |
| Positive pressure | FR-2004 | By Design |
| Installation and removal process | FR-2005 | N/A |
| Installation and removal repeatability | FR-2006 | By Design |
| Installation and removal handling | FR-2007 | By Design |
| Electrical power capacity | FR-2008 | By Design |
| Allowable volume | FR-2019 | By Design |
| Maintenance accessibility | FR-2020 | By Design |
| Mechanical Interface | FR-2021 | By Design |
| Wavefront error | FR-1981 | By Design and Test |

Table 4:  Beam Transport Optical System Compliance Status

| Short Name | ID | Description |
| --- | --- | --- |
| LGS production capability | FR-1932 | By Design |
| LGS Facility subsystems | FR-1933 | By Design |
| Laser System - definition | FR-1934 | By Design |
| Laser Enclosure - definition | FR-1935 | By Design |
| Beam Transport System - definition | FR-1936 | By Design |
| LGS Control System - definition | FR-1937 | By Design |
| LGS Safety System - definition | FR-1938 | N/A |
| LGS Traffic Control System - definition | FR-1939 | N/A |
| Operational lifetime | FR-1940 | By Design |
| Downtime | FR-1941 | By Design |
| LGS asterism configuration | FR-1942 | By Design |
| LGS asterism power levels | FR-1943 | By Design |
| LGS asterism orientation | FR-1944 | By Design |
| LGS photon return | FR-1945 | By Design |
| Transmission | FR-1946 | By Design and Test |
| Polarization | FR-1947 | By Design |
| Laser Guide Star projected size | FR-1948 | By Design |
| Uplink tip tilt offload | FR-1949 | N/A |
| LGS position - stability | FR-1950 | By Design |
| LGS position - tip tilt residual | FR-1951 | By Design |
| LGS position - blind pointing | FR-1952 | By Design |
| Flexure compensation | FR-1953 | By Design |
| Central projection of Laser Guide Stars | FR-1954 | By Design |
| Reuse Keck I launch telescope | FR-1955 | N/A |
| Location of Laser Units | FR-1956 | By Design |
| Cannot vignette Keck Telescope | FR-1957 | By Design |
| Elevation range | FR-1958 | By Design |
| Standards - new instruments and facilities | FR-1959 | By Design |
| Laser safety standards | FR-1960 | By Design |
| Laser light leaks | FR-1961 | By Design |
| Interface to Keck Telescope structure | FR-1962 | By Design |
| Interface to Keck Observatory facilities | FR-1963 | By Design |
| Interface to NGAO software - Multi-system Command Sequencer | FR-1964 | By Design |
| Interface to NGAO software - Data Server | FR-1965 | By Design |
| Interface to AO system - Real Time Control | FR-1966 | By Design. |
| Interface to NGAO software - acquisition system | FR-1967 | By Design |
| Allowable mass | FR-2018 | By Design |

Table 5: LGSF Facility Compliance Status

The following table shows the area of non-compliance.

| Short Name | ID | Non Compliance |
| --- | --- | --- |

Table 6:  Non-Compliance Status

# Risk and Risk Reduction Plan (Thomas)

Identify any procurement that is needed to purchase in the DDR phase.

Based on the risk guidelines of KAON xxx, list the risks associated with the BTO.  The following table shows individual risks within BGS, their ranking and mitigation plans if necessary.

|  | Risk | Ranking | Mitigation |
| --- | --- | --- | --- |
| 1 | Risk 1 | Low, Medium, High |  |
| 2 | Risk 2 | Low, Med, High |  |

# Deliverables

Specify what the deliverables for this system are.  Hardware, software, documentation, procedures.

Figure 6 shows the deliverables for the BGS.

Figure 6: BTO Deliverables

# Management (Jason)

## Budget

Based on the deliverables, what is the expected cost, include procurement, labor.  Include a section on contingency.  Based on the risk, a % of contingency should be available.

The effort estimate for the BGS is x hours of labor and $x for procurements.

| Item | DDR | FSD | I&T | Handover | Total |
| --- | --- | --- | --- | --- | --- |
| Systems Engineer | 40 | 40 | 80 | 40 |  |
| Optical Engineer | 80 | 80 | 80 | 40 |  |
| Electrical Engineer | 80 | 80 | 80 | 40 |  |
| Electrical Technician |  | 80 | 80 | 40 |  |
| Software Engineer | 80 | 120 | 120 | 40 |  |
| Mechanical Engineer | 80 | 80 | 80 | 40 |  |
| Mechanical Technician |  | 80 | 160 |  |  |
| Laser Safety Officer | 10 | 10 | 10 | 10 |  |
| Total |  |  |  |  |  |

Table 7:  BTO Effort Estimates (Hours)

Insert table on Procurements.

| Subsystem Procurement | Cost |
| --- | --- |
| Optical Components and Hardware | 20,000 |
| Mechanical Components and Hardware | 20,000 |
| Motion Devices | 10,000 |
| Diagnostics Devices | 5,000 |
| Safety Devices | 2,000 |
| Controller | 10,000 |
| Software | 0 |
| Total |  |

Table 8:  Procurements

## Budget Reduction Possibilities

## Schedule

# Plans for the next Phase (Jason)

The following effort is planned for the Detailed Design Phase:

- Completed designs and fabrication drawings

- Software keywords definition

- Completed Interface Control Document

- Integration and Test Plans

- Operational and Maintenance Plans

- Handover Plans

- Spares Recommendation

- Retirement of risks

- Budget and Schedule

# Appendix A.  Requirements

| Short Name | ID | Section | Category | Priority | Description |
| --- | --- | --- | --- | --- | --- |
| Beam Transport System - definition (linked) | FR-1969 | Overall | Implementation | Important | The Beam Transport System shall be composed of the optical mechanical system that optically transmits or relays the output beams of the Laser Units along the telescope structure and projects them onto the sky. The Beam Transport System includes opto-mechanical systems located inside the Laser Enclosure, along the telescope tube structure, and across to the telescope secondary mirror support structure. This system also includes opto-mechanical systems inside the secondary structure including the Laser Launch Telescope itself. This system also consists of mechanical and electrical components needed for beam steering and centering. It is also responsible for the generation of the required number and orientation of individual Laser Guider Stars. Monitoring functions such as beam quality, laser power, and polarization not included as part of the supplied Laser Units are also included in this system. |
| Standards - new instruments and facilities (linked) | FR-1970 | Overall | Implementation | Important | The Beam Transport Facility shall comply with the full set of Keck instrument baseline requirements (TBC). This includes baseline conditions for the operational, non-operational, and shipping environments. It includes the standard for both the vibration environment at the observatory and the allowable amount of vibration a system is permitted to produce. It includes interfaces to the Observatory glycol cooling system. It includes general standards for Optical, Mechanical, Electrical, and Software best practices. It includes implementation requirements that mandate certain technical solutions particularly well adapted to use at the Observatory. It also includes baseline documentation standards. |
| Central projection of Laser Guide Stars (linked) | FR-1971 | Mechanical | Implementation | Important | The Beam Transport System shall project each Laser Guide Star beam from an area located behind the Keck Telescope secondary mirror. |
| Reuse Keck I or Keck II Laser Launch Telescope (linked) | FR-1972 | Optical | Implementation | Important | The Beam Transport System shall use a launch telescope identical in design to the Keck I Laser Launch Telescope manufactured by Galileo Avionica. |
| Elevation range (linked) | FR-1973 | Mechanical | Performance | Important | The Beam Transport System shall meet all requirements over its operational elevation range of 20 degrees to 90.5 degrees. The LGS facility shall be able to function in a stand-by mode (TBD) over the elevation range (-5 degrees to 90.5 degrees). |
| Interface to LGS Control System - software | FR-1974 | Software | Interface | Important | The Beam Transport System shall interface with the LGS Control System software via TBD software interfaces. |
| Interface to LGS Control System - electrical | FR-1975 | Electrical | Interface | Important | The Beam Transport System shall interface with the LGS Control System hardware via a TBD format. |
| Reimage Laser Unit pupil to Laser Launch Telescope | FR-1976 | Optical | Interface | Important | The Beam Transport System shall reimage the exit pupil of the Laser Units onto the entrance pupil of the Laser Launch Telescope. |
| Input beam format | FR-1977 | Optical | Interface | Important | The Beam Transport System shall accept three input beams of diameter 3 mm from each Laser Unit. The Laser Unit will deliver a collimated beam with an output beam waist diameter at 1/e2 equal to 3.0 mm plus/minus 0.1 mm. Output beam waist location: 0.0 m plus/minus 0.5 m with respect to the output aperture of each Laser Unit. |
| Output beam format | FR-1978 | Optical | Interface | Important | The Laser Launch Telescope shall have an output Gaussian intensity profile with a 1/e2 diameter of 0.3 m (TBC). |
| Laser Launch Telescope - functional quality | FR-1979 | Optical | Interface | Important | The Laser Launch Telescope shall have as-built optical quality of 60 nm rms (TBC). The requirement is applicable at all operating elevations between 20 and 90 degrees and at all operating temperatures. |
| LGS Focus control | FR-1980 | Optical | Functional | Important | The Beam Transport System shall provide an mechanism so that the LGS beams can be focus at the sodium layer for ranges between 80 km and 270 km. |
| Transmission (linked) | FR-1982 | Optical | Performance | Important | The LGS facility optical transmission shall be equal or greater than 75% at a wavelength of 589 nm. The requirement is for all optics from the output of each Laser Unit to the sky. It includes all transmission losses in the Beam Transport System and the Laser Launch Telescope. It also includes losses from the Laser Launch Telescope secondary obscuration. |
| Damage threshold | FR-1983 | Optical | Performance | Important | All optical components and coatings used in the Beam Transport System optics shall withstand 100 W CW laser power if they transmit or reflect one laser beam. All optical components and coating used in the Beam Transport System optics shall withstand 300 W CW laser power if they transmit or reflect 7 or more laser beams. The energy density will vary as the beam size increases or decreases in the Beam Transport Optical system. In general laser beam energy density will vary depending on the instantaneous size and f/number of individual laser beams, however typical maximum energy density for optical coating would be to withstand greater than or equal to 4.5 kW (TBC) per centimeter squared CW laser power for wavelengths between 580-600 nm (TBC). |
| Automated alignment | FR-1984 | Optical | Functional | Important | The Beam Transport System shall include automated optical alignment mechanism for the laser optical path from the Laser Units output to the input of the Laser Launch Telescope. |
| Beam splitting | FR-1985 | Optical | Functional | Important | The Beam Transport System shall include automated beam splitters and steering mirror to produce the power fraction required and number of LGS on the sky from the outputs of the Laser Units. |
| Asterism generation | FR-1986 | Optical | Functional | Important | The Beam Transport System shall include optics to orient and position the laser beams for input to the Laser Launch Telescope so that the required asterism is generated on the sky. |
| Deployable LGS control | FR-1987 | Optical | Functional | Important | The Beam Transport System shall include optics to orient and position the laser beams for input to the Laser Launch Telescope so that 3 LGS are created which can be positioned at locations coincident with natural guide stars used for tip tilt correction by other NGAO systems. |
| Offload mirrors - function | FR-1988 | N/A | Functional | Unassigned | The Beam Transport System shall include a single mirror for repointing the on axis LGS asterism. Each deployable LGS (Point and Shoot asterism) beam shall have its own independent mirror for repointing. These mirrors shall be driven by pointing error offload from the up-link tip tilt steering mirrors inside the LGS wavefront sensors of the AO system. |
| Offload mirrors - update rate | FR-1989 | N/A | Performance | Unassigned | The Beam Transport System mirrors used to repoint the LGS beams shall be driven by offload error from the up-link tip tilt mirrors at a rate of 10 Hz (TBC). |
| Offload mirrors - range and precision | FR-1990 | N/A | Performance | Unassigned | The Beam Transport System mirrors used for off loading up-link tip tilt correction shall have a range of 30 arc seconds (TBC) and a positioning tolerance of 0.3 arc seconds (TBC). |
| Image natural stars with Laser Launch telescope | FR-1991 | N/A | Functional | Important | The Beam Transport System shall be able to make pointing and optical alignment checks of the Laser Launch Telescope using natural stars. This will require a sensor such as a CCD (TBC) and other optics to be mounted near the Laser Launch Telescope looking upward at natural stars. |
| Quarter wave plate | FR-1992 | N/A | Functional | Important | The Beam Transport System shall have several quarter wave plates needed to convert the linearly polarized output of the Laser Units to circularly polarized output. |
| Quarter wave plate control | FR-1993 | Optical | Functional | Important | The quarter wave plates shall be automated to compensate for depolarizing effects of the other optics in the Beam Transport System. |
| Laser safety - general requirement | FR-1994 | Overall | N/A | Important | The Beam Transport System shall be delivered in compliance of ANSI Standards for Laser Safety. Areas on non-compliance must be agreed upon by NGAO System Team. Additional highlighted safety requirements are meant to bring additional attention to the design, but are in compliance with the standards. |
| Laser safety - Interior finish | FR-1995 | Mechanical | Safety | Important | The Beam Transport System shall be fabricated with properly finished surfaces that come in contact with the laser and viewed by personnel. The surfaces shall comply with nominal hazard zones for the laser. |
| Laser safety - E-Stop | FR-1996 | Electrical | Safety | Important | The Beam Transport System shall provide an emergency stop button to the laser safety system to terminate the laser beams from entering the Launch Facility. |
| Laser safety - Status Indicator | FR-1997 | Electrical | Safety | Important | The Beam Transport System shall provide status indicators on the outside of the enclosure. These indicators are for personnel to determine the laser status prior to entry. The indicators shall be momentary if any light source is used so as to not contaminate the dome environment. |
| Laser safety: laser radiation exposure | FR-1998 | Overall | Safety | Important | The Beam Transport System shall insure personnel must not be exposed to the maximum exposure levels (MPE) as defined by the ANSI standards (Z136.1) and beam hazard analysis. Interlocks shall be installed to meet these requirements. |
| Laser safety: hazard labeling and warning signs | FR-1999 | Mechanical | Safety | Important | The Beam Transport System shall include hazard labels for a Class-4 lasers in accordance with ANSI Standards. |
| Pointing reference | FR-2000 | Mechanical | Functional | Important | The Beam Transport System shall have a mean to initialize its motion control devices to known fiducial. |
| Shutter | FR-2001 | Mechanical | Safety | Important | The Beam Transport System shall accept a command from the Laser Safety System to close a final shutter if necessary. The location of the final shutter will depend on the Beam Transport System design. |
| Environmental monitoring | FR-2002 | Mechanical | Interface | Important | The Beam Transport System shall provide inputs to the environmental control and monitoring system for temperature and humidity. The subsystem of the LGS facility that performs environmental control is LGS Control System (TBC). |
| Yield strength | FR-2003 | Mechanical | Safety | Important | The Beam Transport System shall be designed and constructed to a minimum safety factor of TBD on yield strength for all structural elements. |
| Positive pressure | FR-2004 | Mechanical | Functional | Important | The Beam Transport System shall be positive pressure with dried filtered air to minimize particulates inside of the Launch Facility. The flow rate will be TBD. The output flow shall be distributed for laminar flow and not a point source. |
| Installation and removal process | FR-2005 | Mechanical | Performance | Important | The Beam Transport System component systems shall each be able to be removed and installed within 4 hours or 1/2 day of daycrew operations. The installation shall have a repeatable factor of TBD. |
| Installation and removal repeatability | FR-2006 | Mechanical | Functional | Important | The components of the Beam Transport System shall be installed with kinematic fixture for easy of removal and reinstallation. The repeatability factor is TBD. |
| Installation and removal handling | FR-2007 | Mechanical | Implementation | Important | Mechanical supports shall be implemented to simplify the installation and removal of the parts of the Beam Transport System that are located inside the telescope secondary and along the telescope tube structure. If possible, the interface to the Laser Launch Telescope shall be installed as a single unit. The components of the Beam Transport System shall provide supports such as lifting bolts to assist in installation and removal. |
| Electrical power capacity | FR-2008 | Electrical | Performance | Important | The Beam Transport System shall be provided with 1500 Watts of 120 VAC power. |
| Allowable volume | FR-2019 | Mechanical | Functional | Essential | The Beam Transport System volume shall conform to volume limits of the current f/15 secondary and its mounting structure. Those parts of the Beam Transport System mounted behind the f/15 secondary mirror must allow for the removal, storage and installation of the f/15 secondary module. Furthermore, these parts shall 1) not extend beyond the module in the x,y-directions and 2) must not extend more than 1 m beyond the top of the telescope structure. |
| Maintenance accessibility | FR-2020 | N/A | N/A | Important | The diagnostics of the Beam Transport System shall not interfere with maintenance of existing equipment in the f/15 Secondary module. When possible 50 cm (20 inches) of access shall be provided around equipment that requires maintenance. |
| Mechanical Interface | FR-2021 | Mechanical | Interface | Important | The Beam Transport System shall interface with the telescope, secondary socket, and f/15 structure (TBD). |
| Wavefront error | FR-1981 | Optical | Performance | Important | The Beam Transport Facility optics excluding the Laser Launch telescope shall have an RMS wavefront error of 60 nm rms (TBC). |

Table 9: Beam Transport Optical System Requirements

| Short Name | ID | Section | Category | Priority | Description |
| --- | --- | --- | --- | --- | --- |
| LGS production capability | FR-1932 | Overall | Functional | Essential | The Laser Guide Star Facility shall be responsible for the production of laser guide stars used by the NGAO Adaptive Optics system for sensing optical distortions caused by atmospheric turbulence. This includes all systems needed for production, setup, calibration, pointing, focusing, and control used to produce laser guide stars in the atmospheric sodium layer. |
| LGS Facility subsystems | FR-1933 | Overall | Implementation | Essential | The Laser Guide Star Facility shall be composed of the following systems: Laser System, Laser Enclosure, Beam Transport System, LGS Control System, LGS Safety Systems, and Laser Traffic Control System. (note Laser Guide Star abbreviated LGS) |
| Laser System - definition | FR-1934 | Overall | Implementation | Essential | The Laser System shall be composed of two or three (TBD) individual Laser Units. In addition, the system consists of other associated mechanisms, electronics, control software, and other closely associated hardware provided by the laser supplier/vendor. |
| Laser Enclosure - definition | FR-1935 | Overall | Implementation | Essential | The Laser Enclosure shall be composed of a fully enclosed room mounted on the Keck Telescope structure (tube or elevation ring TBD) that houses the Laser System. This system is composed of the enclosing room along with associated mechanisms and electronics. |
| Beam Transport System - definition | FR-1936 | Overall | Implementation | Essential | The Beam Transport System shall be composed of the optical mechanical system that optically transmits or relays the output beams of the Laser Units along the telescope structure and projects them onto the sky. The Beam Transport System includes opto-mechanical systems located inside the Laser Enclosure, along the telescope tube structure, and across to the telescope secondary mirror support structure. This system also includes opto-mechanical systems inside the secondary structure including the Laser Launch Telescope itself. This system also consists of mechanical and electrical components needed for beam steering and centering. It is also responsible for the generation of the required number and orientation of individual Laser Guider Stars. Monitoring functions such as beam quality, laser power, and polarization not included as part of the supplied Laser Units are also included in this system. |
| LGS Control System - definition | FR-1937 | Overall | Implementation | Essential | The Laser Guide Star Control System shall be responsible for controlling the Beam Transport System for purposes of laser beam steering/pointing and laser beam diagnostics. The LGS Control System shall control the overall state of the Laser System by an appropriate interface to the Laser Unit control software. The LGS system shall have interfaces with the Laser Traffic Control System and the LGS Safety System. The Laser Guide Star Control System shall coordinate/interface to the overall NGAO Multi-system Command Sequencer, the AO Control System and the NGAO Data Server. |
| LGS Safety System - definition | FR-1938 | Overall | Implementation | Essential | The LGS Safety Systems shall include three sub-systems: Laser Safety System, Aircraft Safety System, and the Satellite Safety System. The Laser Safety System shall be responsible for ensuring safe interaction between observatory personnel and the Laser Units. The Laser Safety System shall provide environmental monitors and controls for both the Laser Enclosure and Beam Transport system, ensuring that the Laser Units are not damaged (i.e. "safe"). The Aircraft Safety System shall be responsible for ensuring that aircraft are not illuminated by lasers propagated from the Keck Telescope dome. The Satellite Safety system shall be responsible for ensuring that satellites are not illuminated by lasers propagated from the Keck Telescope dome, in coordination with United States Space Command. |
| LGS Traffic Control System - definition | FR-1939 | Overall | Implementation | Essential | The LGS Traffic Control System is defined as the computer system that monitors the status and coordinates the activities of all telescopes on Mauna Kea which participate in Mauna Kea Laser Traffic control, to prevent accidental illumination of any participating telescopes. This includes all telescopes which are either propagating lasers or may be impacted by laser light from other telescopes. The LGS Facility shall have an interface to this Mauna Kea wide service. |
| Operational lifetime | FR-1940 | Overall | Performance | Essential | The LGS Facility system and its components shall be designed to have a 10 year operational lifetime. |
| Downtime | FR-1941 | Overall | Performance | Essential | The LGS Facility shall have less than 5% (TBC) downtime lost to faults of all systems inclusive. |
| LGS asterism configuration | FR-1942 | Optical | Functional | Essential | The laser guide stars shall be arranged in the following pattern or asterism on the sky. The on-axis pattern is composed of a fixed LGS asterism consisting of one on-axis LGS and three LGSs symmetrically located on a radius of 10 arc seconds (TBC). In addition, the LGS asterism shall have three movable or point-and-shoot (PNS) LGSs to be used to AO correct or "sharpen" three natural guide stars located outside of the fixed asterism. The three LGSs will be positioned at the locations of these randomly located NGSs. As such, the movable LGSs will be located in an annulus that has an outer radius of 60 arc seconds and an inner radius of 15 (TBC) arc seconds. |
| LGS asterism power levels | FR-1943 | Optical | Implementation | Essential | The LGSs shall be allocated different power levels depending on function. A total of 50W (TBC) of laser power will be distributed uniformly between the four fixed LGSs. A total of 25W (TBC) of laser power will be distributed uniformly between the three movable or point-and -shoot LGSs. |
| LGS asterism orientation | FR-1944 | Mechanical | Functional | Essential | The Laser Guide Star facility shall provide a mechanism either optical or mechanical (TBD) for keeping the individual laser guide stars fixed with respect to field stars. The Laser Guide Star asterism shall maintain a fixed orientation with respect to background field stars when the Adaptive Optics system is operated in its "fixed field" mode. When the AO system is operated in its "fixed pupil" mode the LGS asterism shall remain fixed with respect to the telescope pupil and rotate with respect to the background stars. In this mode, the movable or Point-and-Shoot asterism will NOT maintain orientation with the three NGS used for tip-tilt compensation. |
| LGS photon return | FR-1945 | Optical | Performance | Essential | The LGS Facility shall produce a return flux of 740 ph/cm2/sec (TBC) for each of the 4 LGS in the on axis fixed asterism (science asterism) and 495 photons/cm2/sec (TBC) for each of the 3 LGS in the movable asterism. These flux values are measured at the Keck Telescope primary mirror (i.e. the pupil) when the laser projection system is pointed at zenith. This requirement assumes a sodium column density of 3e9 atoms/cm2, mean distance to the sodium layer of 90 km, a one-way atmospheric transmission of 89.6%, and laser beam powers of 9.4 W and 6.25 W respectively. Return flux variations will result from changes in the sodium column density and the direction of the earth's magnetic field. |
| Transmission | FR-1946 | Optical | Performance | Essential | The LGS facility optical transmission shall be greater than or equal to 75% at a wavelength of 589 nm. The requirement is for all optics from the output of each Laser Unit to the sky. It includes all transmission losses in the Beam Transport System and the Laser Launch Telescope. It also includes losses from the Laser Launch Telescope secondary obscuration. |
| Polarization | FR-1947 | Optical | Performance | Essential | The laser beams leaving the Laser Launch Telescope shall be circularly polarized to greater than or equal to 98% (TBC). The handedness, right or left, is arbitrary (TBC) but the beams should be fully polarized to the value given in one or the other of these states. |
| Laser Guide Star projected size | FR-1948 | Optical | Performance | Essential | The LGS spots at the sodium layer shall be less than or equal to 0.9 arc seconds (TBC) without considering the effects of atmospheric turbulence. |
| Uplink tip tilt offload | FR-1949 | Optical | Interface | Essential | The LGS Facility shall be able to adjust pointing of the lasers in order to offload the correction of up-link tip tilt from the up-tip-tilt mirrors. |
| LGS position - stability | FR-1950 | Optical | Performance | Essential | The LGS Facility shall maintain the asterism orientation and shape to 1 arc second (TBC) rms. The asterism orientation must be maintained as the telescope tracks astronomical targets for periods up to one hour (TBC). |
| LGS position - tip tilt residual | FR-1951 | Optical | Performance | Essential | The projected (up-link) tip tilt residual for each LGS shall be less than or equal to 0.5 arc seconds rms (one sigma) on-axis (TBC). This includes the residual from atmospheric turbulence, wind shake, and vibrations. The correction of this residual motion shall be performed on the return trip (down-link) by the AO System Real Time Controller and fast steering mirrors location inside the laser guide star wavefront sensor optics. |
| LGS position - blind pointing | FR-1952 | Optical | Performance | Essential | The blind pointing of the LGS asterism shall be less than 10 arc seconds rms (TBC) with a goal of 1 arc second rms (TBC). |
| Flexure compensation | FR-1953 | Mechanical | Performance | Essential | The LGS Facility shall be capable of correcting for the effects of flexure between the telescope top end structure and the elevation ring structure by up to 20 mm (TBC) in translation (perpendicular to the primary mirror optical axis) and 2 milliradians (TBC) in tilt (relative to the primary mirror optical axis) over elevation angles between 0 and 90 degrees. |
| Central projection of Laser Guide Stars | FR-1954 | Mechanical | Implementation | Essential | The Laser Guide Star Facility shall project each Laser Guide Star beam from an area located behind the Keck Telescope secondary mirror. |
| Reuse Keck I launch telescope | FR-1955 | Optical | Implementation | Essential | The Laser Guide Star Facility shall use a launch telescope identical in design to the Keck I Laser Launch Telescope manufactured by Galileo Avionica. |
| Location of Laser Units | FR-1956 | Mechanical | Implementation | Essential | The laser units shall be mounted on the elevation ring of the telescope. This location varies in orientation with respect to the gravity vector as the telescope tracks in elevation. Additional equipment required by the laser units, such as support electronics, control computers, or cooling systems may be located in the laser enclosures or on the telescope Nasmyth platform. If mounting of the laser units on the elevation ring is not feasible, an alternate location must be agreed upon with the NGAO design team. |
| Cannot vignette Keck Telescope | FR-1957 | Optical | Functional | Essential | No part of the Laser Guide Star Facility shall vignette the main telescope beam. In particular, any parts of the Beam Transport System that cross the secondary mirror support struts (spiders) shall not enlarge these obstructions to incoming starlight. |
| Elevation range | FR-1958 | Mechanical | Performance | Essential | The LGS facility shall meet all requirements over its operational elevation range of 20 degrees to 90.5 degrees. The LGS facility shall be able to function in a non-propagating standby mode (TBD) over the full elevation range of the telescope (-5 degrees to 90.5 degrees). |
| Standards - new instruments and facilities | FR-1959 | Overall | Implementation | Essential | The Laser Guide Star Facility shall comply with the full set of Keck instrument baseline requirements (TBC). This includes baseline conditions for the operational, non-operational and shipping environments. It includes the standard for both the vibration environment at the observatory and the allowable amount of vibration a system is permitted to produce. It includes interfaces to the Observatory glycol cooling system. It includes general standards for Optical, Mechanical, Electrical, and Software best practices. It includes implementation requirements that mandate certain technical solutions particularly well adapted to use at the Observatory. It also includes baseline documentation standards. |
| Laser safety standards | FR-1960 | Overall | Safety | Essential | The LGS Facility shall conform to ANSI standard Z136.1-2000 American National Standard for the Safe Use of Lasers. The LGS Facility shall conform to ANSI Z136.6 -2005 "Safe Use of Lasers Outdoors", Approved December 22, 2005. |
| Laser light leaks | FR-1961 | Overall | Performance | Essential | The Laser Guide Star Facility shall not emit laser radiation or other light into the Keck Observatory dome during normal operations. One exception is made for laser radiation at 589 nm that leaves the output aperture of the Laser Launch Telescope. Laser radiation and other stray light shall not leak from other parts of the Beam Transport System or the Laser Enclosure into the dome. |
| Interface to Keck Telescope structure | FR-1962 | Overall | Interface | Essential | The LGS facility shall have interfaces with the telescope structure at several locations (TBD). These interfaces must not affect the performance of the Keck Telescope when used with NGAO or other instruments. Details will be provided in the subsystem requirements and an interface control document. |
| Interface to Keck Observatory facilities | FR-1963 | Overall | Interface | Essential | The LGS facility shall have an interface with the Keck Observatory facilities such as electrical power and glycol cooling. Details will be provided in the subsystem requirements and an interface control document. |
| Interface to NGAO software - Multi-system Command Sequencer | FR-1964 | Overall | Interface | Essential | The LGS Control system shall have an interface to the NGAO Multi-system Command Sequencer. The overall coordination of the LGS Facility with the NGAO System will be done through this interface. Details will be provided in the subsystem requirements and an interface control document. |
| Interface to NGAO software - Data Server | FR-1965 | Overall | Interface | Essential | The LGS Control system shall have an interface to the NGAO Data Sever for the purpose of transmitting LGS Facility telemetry, diagnostics, command and status information to the data server. Details will be provided in the subsystem requirements and an interface control document. |
| Interface to AO system - Real Time Control | FR-1966 | Overall | Interface | Essential | The LGS Facility shall have an interface with the AO Control System to communicate offloads between the AO system LGS wavefront sensor and the LGS Facility Beam Transport system and to communicate laser shutter status to the AO Control System. Details will be provided in the subsystem requirements and an interface control document. |
| Interface to NGAO software - acquisition system | FR-1967 | Overall | Interface | Essential | The LGS Facility shall have an interface with the AO Control System for communication between the acquisition and guiding control tasks and the LGS Facility Beam Transport System. Details will be provided in the subsystem requirements and an interface control document. |
| Allowable mass | FR-2018 | Mechanical | Functional | Essential | The LGS Facility mass shall conform to following weight limits. The maximum weight of the LGS Facility on the azimuth rotating part of the telescope shall not exceed 10,000kg. The maximum weight of the LGS Facility on the elevation moving portion of the telescope shall not exceed 1700 kg (elevation ring mounting). The maximum weight of the LGS Facility in the top-end module (secondary) shall not exceed 150kg. All mass values are (TBC). |

Table 10:  Overall LGS Facility Requirements