- Notes on reviewing the Motion Control Review Material

Don Gavel; EW responses in BLUE

There is little information on stages needed in the calibration / simulator unit (design is not ready, as noted in the report)

Agreed

Image derotator: where is the calculation of the required accuracy? Is the entry in the “Accuracy/Repeatability” column relative or absolute? (this seems to have a consequence on whether the listed stage meets spec).

This reference is missing. The reference was actually stating that the number was a place holder and came from the existing AO system. Analysis beyond what is presented in http://www.oir.caltech.edu/twiki_oir/pub/Keck/NGAO/NewKAONs/KAON708.pdf is needed.

In general, I think that all the resolution/accuracy specifications need to have backup calculations (based on path length, lever arm, etc.). Shouldn’t these be in the motion control design document or have cross-references to the appropriate design document?

Agreed, probably part of the detailed design phase.

The PI M- designators for the LOWFS and LGSWFS pickoff motors (theta-phi degrees of freedom) don’t seem to be the same as the URSxxx type motors called out by Alex Delacroix in his latest opto-mechanical design. Are these up to date?

I believe so. The PI stages are suggested in http://www.oir.caltech.edu/twiki_oir/pub/Keck/NGAO/WFS/NGAO_OSM_Design_Update_V19.pdf and in e-mail from Alex on 11Feb10. If the design has gone back to the Newport stages, I was not aware of it.

Tracking of the LGSWFS focus needs to be specified. This will need to be continuous as the sodium layer changes distance with sec(zenith).

Noted

ADC motion needs to be specified (talk to Reni Kupke)

Will contact Reni

(Perhaps this is new) There is a 2-position flat to switch the output of the 2nd AO relay between two instruments that needs to be done to a fairly high, repeatable, position accuracy.

The existence of this is known, but perhaps not obvious in the documents. It is the ‘instrument select’ item just after the HODM T/T stage in the master list (682). I don’t have any information on the required performance, nothing special had been assumed.