Keck Adaptive Optics Note 456

Next Generation Adaptive Optics:

System Requirements Document

Version 1.15

Oct. 3, 2007

1	Introduction	1

2	Scope and Applicability	1

3	References	2

3.1	Related Documents	2

3.2	Referenced Drawings	2

4	Revision History	2

5	Background	4

5.1	Purpose	4

5.2	Motivation for the Development of NGAO	4

5.3	Overview	5

6	Overall Requirements	7

6.1	Science Requirements	7

6.1.1	Purpose and Objectives	7

6.1.2	Science Performance Requirements	7

Table 1. Asteroid Companions Survey driven requirements	8

Table 2. Asteroid Companions Orbit Determination driven requirements	11

Table 3. Asteroid size, shape, and composition derived requirements	12

Table 4. Moons of giant planets derived requirements	14

Table 5.  Uranus and Neptune derived requirements	16

Table 6. Planets Around Low Mass Stars derived requirements	16

Table 7.  Debris Disks derived requirements	22

Table 8.  Young Stellar Objects derived requirements	22

Table 9.  Astrometry Science in Sparse Fields derived requirements	22

Table 10a. General relativity effects in the Galactic Center derived requirements	22

Table 10b. Radial velocity measurements derived requirements	23

Table 11.  Resolved Stellar Populations in Crowded Fields derived requirements	24

Table 12. Nearby AGNs derived requirements	25

Table 13. QSO Host galaxies derived requirements	27

Table 14. High-Redshift Galaxies derived requirements	28

Table 15a. Spectroscopic studies of distant galaxies lensed by galaxies	32

Table 15b. Imaging studies of distant galaxies lensed by galaxies.	33

Table 16.  Other: Backup Science derived requirements	35

6.1.3	Science Instrument Requirements	36

Table 18. Science Instrument Requirements	37

6.1.4	Science Operations Requirements	38

Table 19. Science Operations Requirements, Raw Data Quality	38

Table 20. Science Operations Requirements, Data Products Quality	40

Table 21. Science Operations Requirements, Archiving and Retrieval	41

6.2	Observatory Overall Requirements	42

6.2.1	Purpose and Objectives	42

6.2.2	Facility Requirements	42

Table 21. Facility Requirements	42

6.2.3	Telescope and Dome Environment Requirements:	45

Table 22. Telescope and Done Environment Requirements	45

6.2.4	Observatory Science Instrument Requirements	46

Table 23. Observatory Science Instrument Requirements	46

6.2.5	Observatory Operational Requirements	47

Table 24. Observatory Operational Requirements	47

6.2.6	Observatory Implementation Requirements	49

Table 25. Observatory Implementation Requirements	49

7	Optical Requirements	51

7.1	Purpose and Objectives	51

7.2	Performance Requirements	51

7.3	Implementation Requirements	51

7.4	Design Requirements	51

8	Mechanical Requirements	58

8.1	Purpose and Objectives	58

8.2	Performance Requirements	58

8.3	Implementation Requirements	58

8.4	Design Requirements	59

9	Electronic/Electrical Requirements	61

9.1	Purpose and Objectives	61

9.2	Performance Requirements	61

9.3	Implementation Requirements	61

9.4	Design Requirements	61

10	Safety Requirements	62

10.1	Purpose and Objectives	62

10.2	Scope	62

10.3	Laser Safety Requirements	62

10.4	Laser Projection Safety Requirements	62

10.4.1	Aircraft Safety	62

10.4.2	Space Command	62

11	Software Requirements	63

11.1	Purpose and Objectives	63

11.2	Scope	63

11.3	Performance Requirements	63

11.4	Implementation Requirements	63

11.5	Design Requirements	63

12	Interface Requirements	64

12.1	Purpose and Objectives	64

12.2	Performance Requirements	64

12.3	Implementation Requirements	64

12.4	Design Requirements	64

12.4.1	Optical Interface	64

12.4.2	Mechanical Interface	64

12.4.3	Electrical/Electronic Interface	65

12.4.4	Software Interface	65

13	Reliability Requirements	67

13.1	Purpose	67

13.2	Scope	67

13.3	Performance	67

14	Spares Requirements	67

15	Service and Maintenance Requirements	67

16	Documentation Requirements	68

16.1	Documentation Package	68

16.2	Drawings	68

16.3	Electrical/Electronic Documentation	68

16.4	Software	68

17	Glossary	69

Figure 1 Keck telescope structure	6

Table 1 NGAO baseline Mauna Kea Cn2 Profile	36

Table 2 Glossary of Terms	69

Requirements Tables

Table 1. Asteroid Companions Survey driven requirements	8

Table 2. Asteroid Companions Orbit Determination driven requirements	11

Table 3. Asteroid size, shape, and composition derived requirements	12

Table 4. Moons of giant planets derived requirements	14

Table 5.  Uranus and Neptune derived requirements	16

Table 6. Planets Around Low Mass Stars derived requirements	16

Table 7.  Debris Disks derived requirements	22

Table 8.  Young Stellar Objects derived requirements	22

Table 9.  Astrometry Science in Sparse Fields derived requirements	22

Table 10a. General relativity effects in the Galactic Center derived requirements	22

Table 10b. Radial velocity measurements derived requirements	23

Table 11.  Resolved Stellar Populations in Crowded Fields derived requirements	24

Table 12. Nearby AGNs derived requirements	25

Table 13. QSO Host galaxies derived requirements	27

Table 14. High-Redshift Galaxies derived requirements	28

Table 15a. Spectroscopic studies of distant galaxies lensed by galaxies	32

Table 15b. Imaging studies of distant galaxies lensed by galaxies.	33

Table 16.  Other: Backup Science derived requirements	35

Table 18. Science Instrument Requirements	37

Table 19. Science Operations Requirements, Raw Data Quality	38

Table 20. Science Operations Requirements, Data Products Quality	40

Table 21. Science Operations Requirements, Archiving and Retrieval	41

Table 21. Facility Requirements	42

Table 22. Telescope and Done Environment Requirements	45

Table 23. Observatory Science Instrument Requirements	46

Table 24. Observatory Operational Requirements	47

Table 25. Observatory Implementation Requirements	49

# Introduction

This document describes the requirements for the Next Generation Adaptive Optics (NGAO) system to be built for the W. M. Keck Observatory (WMKO).

The requirements in this document are intended to be at a level appropriate for the system design phase.  Further development of the requirements will take place in the next phase of the project (preliminary design).  In particular, parametric performance requirements given at this stage are intended to indicate the scope and format of the requirements, but do not in all cases establish final values for the specified parameters.  In some cases values for these parameters have yet to be established and are given as TBD.

A more generic set of requirements for new WMKO instrumentation is described in the Observatory’s “Instrumentation Baseline Requirements Document.”  These requirements are also applicable to NGAO.  The NGAO System Requirements Document will take precedence over the “Instrumentation Baseline Requirements Document” in the event of a conflict.

It is important to understand that at this stage of development the requirements provide a basis for identifying the parameters that will be part of the system’s specifications, but the values given are subject to change as the development continues.  During the next phases of the project work will be done to refine the requirements for review at the preliminary design review.  The final requirements to be reviewed at the detailed design review will form the basis for the acceptance test criteria for the instrument.

The purpose of this document is to define and communicate the requirements for the NGAO-specific design and implementation in terms of the needed scientific and technical performance.  The document also expresses specific requirements for implementation or design where those requirements are essential to satisfactory integration and interoperation of NGAO with the observatory systems.  The document avoids prescribing specific design or implementation solutions except for solutions that embody the Observatory’s unique knowledge or experience.  The document establishes requirements for the NGAO that will guide the design through the detailed design phase.

# Scope and Applicability

This document establishes requirements for all aspects of NGAO beyond those already specified in the “Instrumentation Baseline Requirements Document”.  This document also establishes requirements for changes to related Keck telescope subsystems and software where required.

This document does not address the requirements for the science instruments that will work with NGAO, although it does cover the NGAO interfaces to these instruments.  Separate system requirements documents will need to be prepared for each of these instruments as part of their design process.

# References

## Related Documents

- Instrumentation Baseline Requirements Document.

- KAON 153.  Coordination and Use of Laser Beacons for AO on Mauna Kea.

- KAON 399.  NGAO Proposal Executive Summary.

- KAON 400.  NGAO Proposal.

- KAON 428.  Implications and requirements for Interferometry with NGAO.

- KAON 455.  NGAO Science Case Requirements Document v1.0.

- KAON 476.  NGAO Science Operations Observing Model Trade Study

- ANSI Z136.1 Safe Use of Lasers Indoors (2000).

- ANSI Z136.6 Safe Use of Lasers Outdoors (2000).

- Architecture_reqments_summary_v7.xls

- 070709_NGAO_RollUp_v1.xls (July 9, 2007) [We need a more official document for this. This is a revised version of the Table in Section 2.5 of KAON 455.  Don.]  [This table will be updated as part of the next official version of the SCRD.  Claire]

## Referenced Drawings

Table X lists the drawing numbers, revisions and date, source and title for all drawings referenced in this document.

# Revision History

| Version | Date | Author | Reason for revision / remarks |
| --- | --- | --- | --- |
| 0.1 | Jan. 16, 2007 | Wizinowich | Initial version |
| 0.3 | Feb. 1, 2007 | Wizinowich | Multiple edits |
| 0.4 | Feb. 6, 2007 | Wizinowich | Multiple edits. Included seeing & telescope environment in section 6.1.2 |
| 1.0 | Feb. 21, 2007 | Wizinowich | Added Dekany performance requirements input |
| 1.1 | Apr. 17, 2007 | Wizinowich | Edited section 6 |
| 1.2 | May 15, 2007 | Wizinowich | Incorporated input from NGAO team meeting 6 |
| 1.3 | May 22, 2007 | Wizinowich | Miscellaneous |
| 1.4 | May 22, 2007 | Wizinowich | Mods to the table in section 6.1.2 |
| 1.5 | May 25, 2007 | Wizinowich | Additions to 6.1.4 based on KAON 476 |
| 1.6 | May 30, 2007 | Wizinowich | Science requirement changes based on SRD telecom |
| 1.7 | May 31, 2007 | Wizinowich | Solar system science requirement input with Le Mignant & Marchis + minor mods to science operation requirements |
| 1.8 | June 1, 2007 | Wizinowich | Galactic Center & QSO host science cases. Mods to Observatory operational requirements |
| 1.9 | June 6, 2007 | Wizinowich | AGN science case from Adkins, Le Mignant, Max & McGrath |
| 1.10 | June 19, 2007 | Wizinowich | Added planets around low mass stars & gravitational lensing requirement tables. First cut at asteroid shape & size table. Added H-band to GR Galactic Center case. Minor edits to asteroid companions table. |
| 1.11 | June 29, 2007 | Wizinowich | Minor edits to science performance requirements |
| 1.12 | August 15, 2007 | Gavel | Review and edits, highlighting gaps in the flow down from the ScRD. Added a requirements numbering system. |
| 1.13 | August 22, 2007 | Gavel | Continued edits, through all the science requirements tables and instrument requirements tables (to Table 13). Comments and questions are in red font. |
| 1.14 | Sept 4 thru 7, 2007 | Max, McGrath, Le Mignant | Revised science requirements tables. We have scrutinized Tables 1, 2, 3, 4, and 6 thoroughly. Things to follow up on are highlighted in yellow. |
| 1.15 | Oct. 1 – 3rd, 2007 | D. Le Mignant | Revised observatory requirements (section 6). Things to follow up on are highlighted in yellow. Updated Table numbering and TOC. Removed the obsolete performance requirement tables from optical performance section 7. Consolidated Telescope and Dome Environment requirements in Table 22 and included in section 6.2.3 |

# Background

## Purpose

The purpose of the background section of this document is to provide context and related information for the requirements defined in later sections of this document.

## Motivation for the Development of NGAO

The Keck telescopes are the world’s largest optical and infrared telescopes.  Because of their large apertures the Keck telescopes offer the highest potential sensitivity and angular resolution currently available.  WMKO has already demonstrated scientific leadership in high angular resolution astronomy with the first NGS and LGS AO systems on 8-10 m telescopes.  The importance of achieving the full potential of the Keck telescopes is recognized in the Observatory’s strategic plan which identifies leadership in high angular resolution astronomy as a key long-term goal.

In order to maintain our leadership in this field we must pursue new AO systems and the instrumentation to exploit them.  We have examined, and are continuing to examine, a broad range of key science goals in order to identify the most compelling future science goals of our community and to determine what is needed to realize these goals.  As a result we have identified that NGAO should provide the following suite of capabilities:

- Near diffraction-limited performance at near infrared wavelengths, producing a point spread function with unprecedented precision, stability and contrast;

- Increased sky coverage and a multiplexing capability, enabling a much broader range of science programs; and

- AO correction in the red portion of the visible spectrum (0.6-1.0 µm), delivering the highest angular resolution images available for filled aperture telescopes.

NGAO will be a broad and powerful facility with the potential to achieve major advances in astrophysics.  It will provide dramatic gains in solar system and galactic science where AO has already demonstrated a strong scientific impact.  NGAO will also allow for extraordinary advances in extragalactic astronomy, far beyond the initial gains being made with the Observatory’s current AO systems.

To be clear NGAO need not be a single facility.  It may be that the requirements are best met with multiple AO systems.

The NGAO proposal (KAON 400) and NGAO proposal executive summary (KAON 399) provide more background on the motivation for the development of NGAO.  Further scientific motivation is provided in the NGAO science case requirements document (KAON 455).

## Overview

The scientific and technical requirements for NGAO result in the following basic systems:

- AO system.  The AO system will likely consist of an AO enclosure, an opto-mechanical system, and software and electronics for both non real-time and real-time control.

- Laser facility.  The laser facility will likely consist of a laser enclosure, the laser(s), the launch facility including a beam transport system and launch telescope, safety systems and laser system control electronics and software.

- Science operations facility.  The science operations facility will primarily include the software and computers required to support operation of the AO system and science instruments.  This includes operating the systems for nighttime observing as well as pre- and post-observing activities.

- Science instruments.  The three highest priority instruments are currently a near-IR imager, a visible imager and a deployable near-IR integral field unit (IFU).  Three lower priority instruments have also been identified including a near-IR IFU, a visible IFU and an L and M-band imager.  There is also a requirement that the NGAO project be designed so as to allow the continued AO support of the Interferometer and the fiber injection module used for the ‘OHANA (Optical Hawaiian Array for Nanoradian Astronomy) project.

The AO and laser facilities and the science instruments will have to interface with the telescope structure.  Figure 1 shows a schematic view of a Keck telescope.  The most likely location for the NGAO system and science instruments is on one of the Nasmyth platforms of the telescope.  Nominally we have chosen the left Nasmyth platform of the Keck II telescope as our starting point.  The most likely location for the projection telescope is behind the f/15 secondary mirror in the top end of the telescope.

Figure 1 Keck telescope structure

# Overall Requirements

## Science Requirements

### Purpose and Objectives

The purpose of the science overall requirements section is to summarize and convey requirements that apply generally to the overall NGAO system and its accessories.  These are based on the Science Case Requirements Document (SCRD) (KAON 455) and, in this initial release, on the NGAO proposal (KAON 399) and on a summary architecture spreadsheet that was developed during NGAO team meetings in the spring and summer of 2007 (Architecture_reqments_summary_v7.xls).

### Science Performance Requirements

The performance requirements developed in the SCRD are summarized in the following tables.  These will be updated as the science case requirements and performance budgets become better defined.

The relevant source of the requirement from the Science Requirements Document (SCRD), KAON 455, is referenced by section and paragraph number. Some of the requirements have a source other than the SCRD.  In these cases the additional source is listed.  Some of the requirements reference recommendations from the Observing Model Trade Study, KAON 476 [Don].

We have categorized the various science cases into two classes: those that push the limits of AO system, instrument, and telescope performance (designated “Key Science Drivers”), and those that are less technically demanding than the “Key Science Drivers” but still place important requirements on available observing modes, instruments, and PSF knowledge.  We shall call the latter category “Science Drivers.”

In the remainder of this section we present tables showing the requirements that flow down from both of these categories of science cases.  The cases we shall discuss are listed below, together with their designation as Key Science Drivers (KSD) or “Science Drivers” (SD).

Reference in

Title and Number in This Report				Summary Spreadsheet

Architecture_reqments_summary_v7.xls

- Asteroid companions survey (KSD)					S1a

- Asteroid companions orbit determination (KSD)			S1b

- Asteroid size, shape, and composition (SD)				S2

- Moons of the Giant Planets (SD)					S3

- Uranus and Neptune (SD)						S4

- Planets around low-mass stars (KSD)					G1

- Debris disks (SD)							G3

- Young stellar objects (SD)						G4

- Place-holder for astrometry in sparse fields (SD)			GX1

- General Relativity at Galactic Center (KSD)				G2

a. Astrometric							G2a

b. Radial velocity						G2b

- Resolved stellar populations in crowded fields (SD)			GX2

- Nearby AGNs: black hole mass measurements (KSD)		X3

- QSO host galaxies (SD)						X1

- High-redshift galaxies (KSD)						X2

- Gravitationally lensed galaxies (SD)

a.  Spectroscopy						X4a

b. Imaging							X4b

- Other: backup science (SD)						O

#### Asteroid Companions Survey

The requirements for the asteroid companions survey science case are summarized in the following table (see the Multiplicity, Size and Shape of Minor Planets section of KAON 455 (V1) and the Multiple Asteroids 1: Survey mode to find new systems Observing Scenario).

Table 1. Asteroid Companions Survey driven requirements

General comment about this table: need to clarify and motivate choice of wavelength band at which contrast requirements are stated.  Current version switches from H to J to I to R bands frequently; hard to understand why.

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 1.1 | The companion sensitivity shall be ΔH ≥ 5.5 mag at 0.5” separation for a V ≤ 17 asteroid specify J magnitude (asteroid size < 0.2”) with a proper motion of ≤ 50 arcsec/hour [ScRD §2.1.4.6, RollUp_v1 E3] | The asteroid can be used as tip/tilt guidestar (proper motion of ≤ 50 arcsec/hour). The AO system has sufficient field of view for objects and for their seeing disks (>3 arcsec, see # 1.8). The tip-tilt residual error will be less than 10 mas while guiding on one V=17 (specify J mag) object at 50 arcsec/hr (14 mas/sec). The AO system has sufficient Strehl to achieve this contrast ratio and sensitivity in 15 min exposure time. (need further simulations to state wavefront error requirement. ~170nm?) | Near-IR imager |
| 1.2 | H-band relative photometric accuracy (between primary and companion) of 5% at 0.6” for ΔH = 3 for a V ≤ 17 specify J magnitude asteroid (asteroid size < 0.2”) with a proper motion of ≤ 50 arcsec/hour [ScRD §2.1.4.4, RollUp_v1 H3] | System has facilities and tools for calibrating the PSF to within xxx percent during one observation. Need to think about what it means to calibrate the PSF to a specific number of percent. Is this the energy in the peak? The width of the peak? Some other parameter? | Near- IR imager (no coronagraph because many asteroids will be resolved) |
| 1.3 | Target sample ≥ 300 asteroids in ≤ 4 yr [SCRD §2.1.3.2 ¶4] Insert in SCRD doc. Leads to requirement of ≥ 25 targets per 11 hour night [SCRD §2.1.5.2, RollUp_v1 M3] | Assumes 3 good nights per year. Needs high observing efficiency: Able to slew to new target and complete the entire observation within 26 minutes on average. |  |
| 1.4 | Observing wavelengths = J and H-band what about R band? [SCRD 2.1.6.5, RollUp_v1 B3] | Transmit R through z to visible imager, z through H to IR imager. | Visible and IR imagers. |
| 1.5 | Spatial sampling ≤ Nyquist at the observing wavelength. [SCRD §2.1.6.1] |  | Spatial sampling ≤ Nyquist at the observing wavelength (R through H bands? - explain). |
| 1.6 | Field of view ≥ 3” diameter [SCRD §2.1.6.2] | AO system passes a >3” unvignetted field of view | Imager fields of view ≥ 3” |
| 1.7 | The following observing preparation tools are required: guide star finder for asteroids too faint to use as the only TT star, PSF simulation as function of wavelength and seeing conditions. [ Need to retrofit into SCRD] | Guide star finder tool. PSF simulation tool (predict energy and width of central core to within 10%). |  |
| 1.8 | The following data products are required: Access to archive with proper identification in World Coordinate System (to within 1 arc sec or better) and with associated PSF calibrated to within xxx %. [Need to retrofit into SCRD] | Calibrated PSF capability (to what accuracy? Why) Ability to collect AO telemetry data to support the required PSF calibration. | FITS header system is capable of handling non-sidereal offsets in reporting object coordinates in the World Coordinate System to within 1 arc sec or better. |
| 1.9 | Observing requirements: Observer present either in person or via remote observing rooms, because real-time observing sequence determination is needed. | Classical observing mode or service mode with active observer participation. Remote observing capabilities must allow frequent real-time decisions by observer. |  |

#### Asteroid Companions Orbit Determination

The requirements for the asteroid companions orbit determination science case are summarized in the following table (see the Multiplicity, Size and Shape of Minor Planets section of KAON 455 (V1) and the Multiple Asteroids 2: Orbits Determination for the Discovered System Observing Scenario in KAON ???).

Table 2. Asteroid Companions Orbit Determination driven requirements

need to clarify what wavelengths are needed and why

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 2.1 | Same as #1.1 | Same as #1.1 | Near-IR imager |
| 2.2 | The companion sensitivity shall be ΔI ≥ 7.5 mag at 0.75” separation for a V ≤ 17 asteroid (asteroid size < 0.2”) with a proper motion of ≤ 50 arcsec/hour [need source] |  | Imager [what wavelength range?] |
| 2.3 | Photometric accuracy: Same as #1.2 | Same as #1.2 |  |
| 2.4 | H-band relative astrometric accuracy of ≤ 1.5 mas for a V ≤ 17 [what J mag?] asteroid (asteroid size < 0.2”) with a proper motion of ≤ 50 arcsec/hour [SCRD §2.1.4.5 RollUp_v1 I3] | Non-sidereal tracking accuracy <10 mas. Thoughput, emissivity to the guider sufficient to achieve SNR ≥ 400 on the primary in <1 second. [Do we need this additional requirement on SNR if the non-sidereal tracking accuracy is < 10 mas? And where does the <1 sec requirement come from?] | Uncalibrated detector distortion < 1.5 mas (Is this an edge to edge number? Over xxx pixels? Need to be more explicit.) |
| 2.5 | Target sample size of ≥ 100 asteroids in ≤ 4 years [SCRD §2.1.3.2 ¶4]. Leads to requirement of ≥ 25 targets in an 11 hour night. [SCRD §2.1.5.2, RollUp_v1 M3] | Assumes 3 good nights per year, need 3 epochs per target. Needs high observing efficiency: Able to slew to new target and complete the entire observation within 25 minutes on average. Will generally only observe at one wavelength (the one that gives the best astrometric information). |  |
| 2.6 | Observing wavelengths = J and H-band [SCRD §2.1.6.5, RollUp_v1 B3] | Same as #1.6 (Note: R-band may become a future requirement if R-band Strehl > 15%) | Near-IR imager (R-band imager) |
| 2.7 | Spatial sampling same as #1.5 |  | Same as #1.5 |
| 2.8 | Field of view ≥ 2” diameter [why not 3”?] [SCRD §2.1.6.2] | AO system passes a >3” unvignetted field of view | Imager fields of view ≥ 3” |
| 2.9 | Same as #1.7 | Same as #1.7 |  |
| 2.10 | Same as #1.8 | See #1.8 |  |
| 2.11 | Observing requirements: 8 epochs per target per semester. [SCRD §2.1.3.2 ¶4] |  |  |

#### Asteroid Size, Shape, and Composition

The requirements for the asteroid size, shape, and composition (characterize surface and orbital parameters) science case are summarized in the following table (see the Multiplicity, Size and Shape of Minor Planets section of KAON 455 (V1)).  In addition to the requirement of a high resolution visible imager, the slope of the visible spectrum is needed to determine the asteroid age or surface type.  This case requires a spectral resolution of R ~ 100 for 0.7 – 1.0 µm wavelength with Nyquist sampling.  Since the slope of the spectra is of importance this could be accomplished with a visible IFU, with a visible slit spectrograph, or with multiple narrow-band filters.

Table 3. Asteroid size, shape, and composition derived requirements

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 3.1 | The companion sensitivity shall be ΔI ≥ 7.5 mag at 0.75” separation for a V ≤ 17 asteroid (asteroid size < 0.2”) with a proper motion of ≤ 50 arcsec/hr [RollUp_v1 F2] | The asteroid can be used as the tip/tilt guidestar (proper motion of ≤ 50 arcsec/hr). The AO system has sufficient field of view for both objects and their seeing disks (>3 arcsec, see requirement 1.8). The tip-tilt residual error will be less than 10 mas while guiding on one V=17 (specify J mag) object at 50 arcsec/hour (14 mas/sec). The AO system has sufficient Strehl to achieve this contrast ratio and sensitivity in 15 minutes of exposure time. (need further simulations to state wavefront error requirement.) | I-band imager. |
| 3.2 | Target sample size of ≥ tbd asteroids in ≤ tbd years [needs to be specified]. ≥ 10 targets in an 11 hour night [SCRD §2.1.5.2, RollUp_v1 M2] | #2.5 is stricter requirement. |  |
| 3.3 | Observing wavelengths 0.7 – 1.0 µm [SCRD §2.1.6.5, RollUp_v1 B2] | AO system must pass 0.7 to 1.0 micron wavelengths with sufficient sensitivity to satisfy #3.1 | Visible imager (R through I band) with narrow-band filters or slit spectrograph (R~100), or possibly visible IFU (R~100). |
| 3.4 | Spatial sampling same as #1.5 | Same as #1.5 | Same as #1.5 |
| 3.5 | Field of view ≥ 2” diameter [SCRD §2.1.6.2] | Same as #2.8 | Same as #2.8 |
| 3.6 | Ability to measure the spectral slope with R ~ 100 at 0.85-1.0 m [SCRD §2.1.6.6, §2.5] |  | Spectroscopic imaging at R ~ 100 across R and I-band, potentially with narrow-band filters (how many? at what spacing?) |
| 3.7 | Ability to measure the SO2 frost bands at R=1000 (R=4000 is acceptable) at 1.98 and 2.12 m [SCRD §2.1.6.6, §2.5] |  | Spectroscopic imaging at R ~ 1000 or 4000 in the K band. Are there other lines of interest possibly in H or J? |
| 3.8 | Same as #1.7 | Same as #1.7 |  |
| 3.9 | Same as #1.8 | Same as #1.8 |  |

#### Moons of the Giant Planets

The requirements for the moons of the giant planets science case are summarized in the following table (see the TBD section of KAON 455 (V1) and the Io Observing Scenario). Don: Note: moons of giant planets science case is not written up in the ScRD, and is essentially blank in the roll-up tables in section 2.5 in the ScRD. I’ve seen subsequent roll-up charts, but I’m not sure where the “official” version is stored (070709_NGAO_RollUp_v1.xls posted on Meeting 8 page? Still a blank row on moons.)  Claire: this topic will be written up as part of the next SCRD release.

Table 4. Moons of giant planets derived requirements

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 4.1 | Able to map intensity distribution with a spatial resolution ≤ 20 mas at 0.7 µm (are 0.8 or 0.9 microns OK? what science would be lost?) for a target with V = 5 to 8 with proper motion ≤ 30” per hour Need to lay out in SCRD what unique info about each moon can be gained with NGAO in the visible, as opposed to the NIR. | The moon will be used as the tip/tilt and NGS guidestar (at a non-sidereal rate of 30 arcsec/hr). Tracking (tip-tilt) accuracy must be better than 5 mas. Need to simulate and calculate what wavefront error would give the required “spatial resolution” at 0.7 or 0.8 or 0.9 microns. | Visible imager. |
| 4.2 | Able to acquire Io within 5” of Jupiter and to track it to within 2.5” of Jupiter. Note that this is a goal but perhaps not a rigid requirement: we know we can acquire within 10” today. | May require either a diaphragm or a filter to attenuate the light from Jupiter. | See AO derived requirement at left. |
| 4.3 | Relative photometric accuracy of ≤ 0.05 mag on target. (What science is possible at a variety of levels of accuracy? Need to rephrase photometric requirement as a function of observing wavelength, and take into account possible cirrus) | Puts a requirement on Strehl and PSF knowledge at each wavelength. What is this requirement? | Detector flat-fielding requirements, linearity, etc will flow down from required photometric accuracy. Need to specify. |
| 4.4 | Absolute photometric accuracy ≤ 0.05 mag on target. (What science is possible at a variety of levels of accuracy? Need to rephrase photometric requirement as a function of observing wavelength, and take into account possible cirrus) | Puts a requirement on Strehl and PSF knowledge at each wavelength. What is this requirement? | Detector flat-fielding requirements, linearity, etc will flow down from required photometric accuracy. Need to specify. |
| 4.5 | Relative astrometric accuracy ≤ 2 mas on target to determine positions of features within disk of a moon (diameter ≤ 1.2 arc sec). |  | Well-calibrated detector distortion (see # 2.4) |
| 4.6 | Observing wavelengths 0.7 to 2.4 µm (how crucial is 0.7 microns to the science, compared with 0.8 or 0.9?) | AO system must pass these wavelengths to science instruments. | Visible imager. Near-IR imager. May require a number of (how many?) specific narrow-band filters; implies needed flexibility in filter wheels. |
| 4.7 | Spatial sampling ≤ Nyquist at all wavelengths |  |  |
| 4.8 | Moons are very bright: do not allow saturation. Typical brightness: 5 mag per square arc sec. Check this. |  | Either need to use neutral density filters, or have a fast shutter, or have a detector with large wells or very short exposure times (and low read noise). Note: these observations will have high overhead. |
| 4.9 | Field of view ≥ 2” diameter. | Must pass ≥ 2” | Minimum FOV 2” |
| 4.10 | The following observing preparation tools are required: PSF simulation, target ephemeris, exposure time calculator to enable choice of ND filter and exposure time. |  |  |
| 4.11 | The following data products are required: Calibrated PSF. (to what accuracy?) |  |  |
| 4.12 | Io and Titan are time domain targets; Io requires ≤ 1 hr notification of volcano activity. Typical timescales for clouds on Titan are of order days or weeks. |  |  |

#### Uranus and Neptune

Text and tables not yet written

Table 5.  Uranus and Neptune derived requirements

#### Planets Around Low-Mass Stars

The requirements for the planets around low mass stars science case are summarized in the following Table (see the Imaging and Characterization of Extrasolar Planets around Nearby Stars section of KAON 455 (V1)).  The key area in which NGAO will excel is the detection of planets around low-mass stars and brown dwarfs, because Keck, unlike GPI, will be able to use laser guide stars.  NGAO will also be able to search for planets around young solar-type stars where dust extinction is significant.  JWST will have coronagraphic capability in the 3 to 5 m window, but will have significantly lower spatial resolution than Keck NGAO.  In terms of the types of solar systems that can be studied, this means that JWST will focus on older, nearby main sequence stars (since older giant planets will remain visible in 3 to 5 m for a longer time).  JWST may be more limited than NGAO in doing large surveys, because of its longer slewing time and possibly a lifetime limit on the total number of slews.

Note: this science will usually be in LGS mode. Depending on the target magnitude, may choose to use the parent star as one of the required tip-tilt stars.

Table 6. Planets Around Low Mass Stars derived requirements

|  | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 6.1 | Target sample 1: Old field brown dwarfs out to distance of 20 pc. Sample size several hundred, desired maximum survey duration 3 yrs (practical publication timescales). | Observe 20 targets per night (each with e.g. 20 min integration time). Need a requirement on tip-tilt stars (what magnitude are the parent stars?) | Near infrared imager (possibly with coronagraph). Survey primary stars at J- and H-band. |
| 6.2 | Target sample 2: Young (<100 Myr) field brown dwarfs and low-mass stars to distance of 80 pc. Sample size several hundred, desired maximum survey duration 3 yrs. | Observe 20 targets per night (each with e.g. 20 min integration time). | Near infrared imager (possibly with coronagraph). Survey primary at J- and H-band. Could benefit from dual- or multi-channel mode for rejecting speckle suppression, but not essential for this program. |
| 6.3 | Target sample 3: solar type stars in nearby star forming regions such as Taurus and Ophiuchus, and young clusters @ 100 to 150 pc distance. Bright targets (on-axis tip-tilt generally possible: V=14-15, J=10-12). Sample size several hundred, desired maximum survey duration 3 yrs. | (May not require LGS if there is a good enough near-IR wavefront sensor available). | Possible dual- or multi-channel mode for speckle suppression. Alternatively an IFU would help, provided it is Nyquist sampled at H and has FOV > 1 arc sec. Min. IFU spectral resolution is R~100. May need IR ADC for imaging or coronagraphic observations (J or H bands); typical airmass is 1.7 for Ophiuchus. |
| 6.4 | Companion Sensitivity Sample 1: assume no companions beyond 15 AU. Targets at 20 to 30 pc; companion distribution peaks at 4 AU = 0.2"; this yields 2 MJupiter planets at a 0.2" separation with -contrast ΔJ = 10. Will be retrofitted into SCRD. Should put in apparent magnitudes | Excellent (5 nm rms? needs simulation) calibration of quasi-static non-common path aberrations, especially at mid-spatial-frequencies. Needs algorithms such as phase retrieval or speckle nulling (on a fiber source + good stability). Small servo-lag error (how small?) to avoid scattered light at 0.2 arc sec. Source: Bruce Macintosh’s error budget and simulations. Need to document. | 6 to 10 λ/D general-purpose coronagraph. Need to describe the coronagraph more explicitly. Speckle suppression capability (multi-spectral imaging) Simulations needed. |
| 6.5 | Companion Sensitivity - Sample 2: a 1 MJupiter planet is at 300K, J=22, ΔJ = 11 (2 MJupiter is 2.5 mags brighter). This distribution could have a wider distribution of binaries a) 0.1" separation, ΔJ = 8.5 (2MJ) b) 0.2" separation, ΔJ = 11 (1MJ) c) Goal ΔJ = 11 at 0.1" separation (1MJ) based on properties of the planets you want to look for. Put in apparent magnitudes | Excellent (5 nm rms? needs simulation) calibration of quasi-static non-common path aberrations, especially at mid-spatial-frequencies. Needs algorithms such as phase retrieval or speckle nulling (on a fiber source + good stability). Small servo-lag error (how small?) to avoid scattered light at 0.2 arc sec. Source: Bruce Macintosh’s error budget and simulations. Need to document. | a) 6 λ/D general-purpose coronagraph b) 6 or 10 λ/D general-purpose coronagraph c) (Goal) Not achievable with a general purpose coronagraph May need small Inner Working Distance coronagraph. Non-redundant aperture masking is an interesting approach for this, limits currently unknown, probably requires low read noise in science detector. |
| 6.6 | Goal: Companion Sensitivity Case 3 (need to do simulations): at 5 Myr , 1 Msun primary; a) goal ΔJ = 13.5 to see 1 MJupiter or b) goal ΔJ = 9 for 5 MJupiter. 0.2" is uninteresting, 0.07" is needed. Should put in apparent magnitudes (probably J = 22 or 23) | Excellent (5 nm rms? needs simulation) calibration of quasi-static non-common path aberrations, especially at low-spatial-frequencies. Needs algorithms such as phase retrieval or speckle nulling (on a fiber source + good stability). Small servo-lag error (how small?) to avoid scattered light at 0.2 arc sec. Source: Bruce Macintosh’s error budget and simulations. Need to document. | a) Not achievable with a general purpose coronagraph (requires a very small IWA coronagraph, non-redundant masking – maybe have a conversation with Ireland, Lloyd or Tuthill). b) 6 λ/D general-purpose coronagraph |
| 6.7 | Sensitivity of H=25 for 5-sigma detection in 20 minutes, at 1 arcsec separation from primary star. [RollUp_v1 G5] | Sufficiently high throughput and low emissivity to permit detecting H=25 in 20 minutes at 5 sigma above background. (need to derive throughput and emissivity numbers) Sufficiently low wavefront error to permit the coronagraph to suppress scattered light from primary to this level. TBD. 140nm meets goal; simulations by Ralf Flicker of 170nm underway. May be in KAON. [SCRD §2.2.5.1] |  |
| 6.8 | H-band relative photometry (between primary and companion): accuracy ≤ 0.1 mag for recovered companions; goal of measuring colors to 0.05 mags (0.03 mag per band) [where did this requirement come from? Needed to measure temperatures and surface gravities sufficiently accurately. For what?] | Diagnostics on AO data to measure Strehl fluctuations if it takes a while to move on and off the coronagraph (a possible more attractive solution is a specialized coronagraph that simultaneously images the primary) | Induced ghost images of primary; or rapid interleaving of saturated and unsaturated images; or a partially transparent coronagraph |
| 6.9 | Requirement: Astrometric precision 2 mas (~1/10 PSF) relative between primary and planet, for initial rejection of background objects. [SCRD §2.2.5.5, RollUp_v1 I5] Goal: For measuring orbits of nearby field objects, want 0.5 mas to measure masses to 10%. Note this gives you mass of primary star. Difficult goal: To get mass of companion you need absolute astrometry at very low level (0.1 mas?). It’s unclear if we can meet this goal. Could be combined with Doppler measurements if that’s practical for the brighter objects. Need to elaborate this. All of this has to be retrofitted into SCRD | Ways to do this: a) Position stability requirement for star behind coronagraph. 2) Induced ghost image method. Needs a wire grating ahead of the coronagraph, or use DM to induce ghost images. (papers by Marois et al. and Sivaramakrishnan et al ApJL 2006). Get reference | Distortion requirement TBD (similar to asteroid science case). Also want ghost images of primary (as for photometry #5.7) in order to locate it accurately relative to planet. |
| 6.10 | Efficiency: 20 targets per night (30 goal) [SCRD §2.2.6.2] | AO system must be able to absolutely steer objects so they land on the coronagraph. This implies 0.005" reproducibility of field steering –or lock the tip/tilt to this accuracy relative to coronagraph field stop. How did we calculate this required accuracy number? |  |
| 6.11 | Observing wavelengths JHK bands (strong goal: Y and z for companion characterization) [SCRD §2.2.4, §2.2.7.1, RollUp_v1 B5] | Transmit JHK to science instrument. Goal: Y and z. | Methane band filters for rapid discrimination, Y and z, and/or a custom filter for early characterization. |
| 6.12 | Able to register and subtract PSFs (with wavelength, time, etc.) for post-processing to get rid of residual speckles. (To what accuracy do you need to subtract PSFs?) | What is the derived requirement here? | 1.5 x better than Nyquist sampled at J (goal Y) |
| 6.13 | Field of view: must see companions at 100 AU scales at 30 pc (goal 20 pc) [SCRD §2.2.5.3] |  | Field of view 3" radius (goal 5" radius) |
| 6.14 | Characterization of companion [SCRD §2.2.4, §2.5] |  | a) R ~150 IFU, sub-Nyquist sampling spectrograph, or if above not available, b) Nyquist spatial sampling IFU, R ~ 4,000, OH suppressing). Both must be sensitive to J = 22 or 23 (in how long an integration time?). |
| 6.15 | Sky Coverage >30%. How was this derived? [SCRD §2.2.7.3, RollUp_v1 N5] | Technical field for low-order wavefront guidestar pickoff at least 150 (180?) arcseconds to achieve this at high galactic latitude. Ability to acquire and track 3 tip/tilt stars. (More lenient if parent star can be used as one of the three TT stars.) |  |
| 6.16 | The following observing preparation tools are required: guide star finder for high proper-motion stars |  |  |

#### Debris Disks

Text and tables not yet written.  Will be a driver for coronagraph design.

Table 7.  Debris Disks derived requirements

#### Young Stellar Objects

Text and tables not yet written.  May be a driver towards having an infrared wavefront sensor.

Table 8.  Young Stellar Objects derived requirements

#### Astrometry Science in Sparse Fields

Text and tables not yet written.  Will be a driver for low and/or very well calibrated instrument distortions, compensation for atmospheric differential refraction, and good temperature control of the AO system.

Table 9.  Astrometry Science in Sparse Fields derived requirements

#### General Relativity Effects in the Galactic Center

The requirements for the General Relativity Effects in the Galactic Center science case on both precision astrometry and radial velocities are summarized in the following two tables, respectively (see the Precision Astrometry: Measurements of General Relativity Effects in the Galactic Center section of KAON 455 (V1)).

Table 10a. General relativity effects in the Galactic Center derived requirements

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 10a.1 | Astrometric accuracy ≤ 100 µas for objects ≤ 5” from the Galactic Center [SCRD §2.3.8, RollUp_v1 I6] | High Strehl to reduce confusion limit: rms wavefront error ≤ 170 nm. IR tip/tilt sensors. Lower background than current AO system (not quantified how much lower) |  |
| 10a.2 | Observing wavelengths K-band [SCRD §2.3.11, RollUp_v1 B6] | Transmit K band to science instrument |  |
| 10a.3 | Spatial resolution ≤ Nyquist at H and K Not called out. Perhaps derived from #6.1 |  |  |
| 10a.4 | Spatial resolution stable over field of view to 100 mas [SCRD §2.3.9] | Means of aligning tip/tilt sensors so that their mean reference position distorts the field by less than 100 mas in closed loop operation. Means of preventing WFS-blind field-distortion modes (if multi-DMs are in series). Error budgets are needed for these. |  |
| 10a.5 | Field of view ≥ 10” x 10” for imaging [SCRD §2.3.11] | Science path shall allow an unvignetted 10” x 10” field. Note: there is no specification on allowable isoplanatic error |  |
| 10a.6 | Target drift should be ≤ TBD mas Not sure what this means. See #6.4. |  |  |
| 10a.7 | Ability to construct 40x40” mosaic to tie to radio astrometric reference frame [SCRD §2.3.5] To what accuracy? |  |  |
| 10a.8 | The following observing preparation tools are required: not specified |  |  |
| 10a.9 | The following data products are required: Not specified |  |  |

Table 10b. Radial velocity measurements derived requirements

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 10b.1 | Radial velocity accuracy ≤ 10 km/sec for objects ≤ 5” from the Galactic Center [SCRD §2.3.10, RollUp_v1 J6] | High Strehl, low background, differential atmospheric refraction correction, PSF estimation [suggestions from SCRD §2.3.10] (need to derive required accuracies. Need to make an error budget for ADC accuracy across band) | Spectral resolution ≥ 4000 |
| 10b.2 | Observing wavelengths K-band [SCRD §2.3.6] |  |  |
| 10b.3 | Spatial resolution ≤ 20 or 35 mas [SCRD §2.3.6] |  | 20 and 30 mas plate scales, R=4000 IFU spectrometer with… |
| 10b.4 | Field of view ≥ 1” x 1” [SCRD §2.3.6] |  | …1” x 1” spatial field of view |
| 10b.5 | Target drift should be ≤ TBD mas Not sure what this means. See #6.4. |  |  |
| 10b.6 | The following observing preparation tools are required: not specified |  |  |
| 10b.7 | The following data products are required: IFU pipeline for wavelength/flux calibration [SCRD §2.3.6] |  |  |

#### Resolved Stellar Populations in Crowded Fields

Text and tables not yet written.

Table 11.  Resolved Stellar Populations in Crowded Fields derived requirements

#### Nearby AGNs: Black Hole Mass Measurements

The requirements for the Nearby AGNs science case are summarized in the following table (see the Nearby AGN section of KAON 455 (V1)) (to be written).  The typical AGN that we are considering is at redshift <0.1, and if a Seyfert 1 galaxy has a magnitude yyy point source in the center, with a host galaxy of magnitude zzz per square arcsec.

The scientific goals are to measure the black hole mass using stellar kinematics in the cores of AGNs.  In order to accomplish this, PSF subtraction will be crucial for Seyfert 1 galaxies.

The region of interest for spatially resolved spectroscopy is within the gravitational sphere of influence of the central black hole: generally we will need at least two resolution elements across this distance.  [convert to arcsecs as function of z]

Table 12. Nearby AGNs derived requirements

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 12.1 | Number of targets required: sample size of TBD galaxies in TBD nights or years | Requirement on sky coverage fraction may be implied here |  |
| 12.2 | Required wavelength range 0.85 – 2.4 microns |  |  |
| 12.3 | Required spatial sampling at least two resolution elements across gravitational sphere of influence | 80% enclosed energy radius < ½ gravitational sphere of influence. This implies a total wavefront error of not more than TBD nm over a range from TBD to TBD microns. | Spectral and imaging pixels/spaxels < ½ gravitational sphere of influence (in the spatial dimension) |
| 12.4 | Required field of view for both spectroscopy and imaging > 10 radii of the gravitational sphere of influence [fill this in] |  | Will need to get sky background measurement as efficiently as possible. For IR, consider using d-IFU on the sky; for visible, need solution |
| 12.5 | Required SNR for spatially resolved spectroscopy of the central black hole region using stellar velocities > 30 per resolution element | AO Strehl ratio > TBD at 0.85 microns (Ca infrared triplet). This implies a total wavefront error of TBD nm at 0.85 microns. PSF stability and knowledge, temporal and field of view [uniformity trade]; | Spectral resolution R ~ 3000-4000 (TBD) with two pixels per resolution element; detector limited SNR performance; Spatial sampling at least two resolution elements across the gravitational sphere of influence |
| 12.6 | Required signal to noise ratio for imaging of the region around the central black hole [is this a contrast requirement?] | AO Strehl ratio > TBD, this implies a total wavefront error of not more than TBD nm over a range from TBD to TBD microns. PSF stability and knowledge, temporal and field of view [uniformity trade]; | Spatial sampling at least two resolution elements across the gravitational sphere of influence |
| 12.7 | Photometric accuracy required for imaging the central point source and possible cusp: TBD at TBD wavelengths | AO Strehl ratio > TBD, this implies a total wavefront error of not more than TBD nm over a range from TBD to TBD microns. AO background level < TBD; over TBD wavelength range; PSF stability and knowledge, temporal and field of view [uniformity trade]; | Calibration stability and accuracy, Zero-point stability and knowledge, Quality of flat-fielding; |
| 12.8 | Velocity determined to ≤ TBD km/sec for spatial resolutions of TBD mas | PSF intensity distribution known to ≤ xxx% per spectral channel. Spatial and spectral model fitting valid to ≤ TBD |  |
| 12.9 | Required observation planning tools (e.g. guide stars); PSF simulation tools to plan for observations of Seyfert 1 galaxies which have strong central point sources |  |  |
| 12.10 | Required data reduction pipeline for IFU |  |  |

#### QSO Host Galaxies

The requirements for the QSO Host Galaxies science case are summarized in the following table (see the QSO Host Galaxy section of KAON 455 (V1)).  The typical QSO that we are considering is at redshift xxx, had a magnitude yyy point source in the center, with a host galaxy of magnitude zzz per square arc sec.  The galaxy is TBD arc sec across.  [Say something about the profile of the galaxy]  The scientific goals are the following: 1) measure colors and magnitudes for the point source; 2) measure morphology and surface brightness profile for the galaxy; 3) obtain spectrum of point source; 4) obtain spatially resolved spectrum of galaxy in order to study its kinematics and stellar populations.  In order to accomplish these things, PSF subtraction will be crucial.  [This is a derived requirement.]  Claire: The following requirements table was developed in a lengthy telecon with members of the science team.  It will be documented in the next version of the SCRD.

Don: ***QSO Host galaxies are not covered in the SCRD, nor are there any entries in the AO Requirements RollUp spreadsheet. These specs should be put on hold until the science requirements are defined.***

Table 13. QSO Host galaxies derived requirements

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 13.1 | Number of targets required: sample size of TBD galaxies in TBD nights or years | Requirement on sky coverage fraction may be implied here, particularly if data from space missions or radio surveys is required |  |
| 13.2 | Required signal to noise ratio for imaging of central point source: TBD at TBD wavelength | AO Strehl ratio > TBD AO background level < TBD; over TBD wavelength range | Instrument background level; Need to be background limited (implications for detector) |
| 13.3 | Photometric accuracy required for imaging the central point source: TBD at TBD wavelengths. | PSF stability and knowledge, temporal and field of view [uniformity trade]; Target drift should be ≤ TBD mas per hour | Calibration stability and accuracy, Zero-point stability and knowledge, Quality of flat-fielding |
| 13.4 | Required SNR for spatially resolved spectroscopy of the point source > TBD at TBD wavelengths | AO Strehl ratio > TBD AO background level < TBD; over TBD wavelength range; PSF stability and knowledge, temporal and field of view [uniformity trade]; Target drift should be ≤ TBD mas per hour | [What spatial sampling is optimum for good PSF subtraction?], spectral resolution R ~ TBD; detector limited SNR performance; FOV = TBD arc sec [e.g. larger for sky subtraction] |
| 13.5 | Required signal to noise ratio for imaging of host galaxy in presence of point source: TBD at TBD wavelength at TBD distance from the point source | AO Strehl ratio > TBD AO background level < TBD; over TBD wavelength range; PSF stability and knowledge, temporal and field of view [uniformity trade]; | Spatial sampling TBD |
| 13.6 | Photometric accuracy required for imaging the host galaxy: TBD at TBD wavelengths at TBD distance from point source. | AO Strehl ratio > TBD AO background level < TBD; over TBD wavelength range; PSF stability and knowledge, temporal and field of view [uniformity trade];Target drift should be ≤ TBD mas per hour | Calibration stability and accuracy, Zero-point stability and knowledge, Quality of flat-fielding; FOV = TBD arc sec [e.g. larger for sky subtraction] |
| 13.7 | Required SNR for spatially resolved spectroscopy of the host galaxy > TBD; spatial resolution = TBD; velocity resolution and accuracy = TBD; wavelength range TBD | AO Strehl ratio > TBD AO background level < TBD; over TBD wavelength range; PSF stability and knowledge, temporal and field of view [uniformity trade]; AO 80% enclosed energy radius = TBD; Target drift should be ≤ TBD mas per hour | FOV = TBD; spectral resolution = TBD; spatial sampling (spaxel size); detector-limited performance; Calibration stability and accuracy, Zero-point stability and knowledge, Quality of flat-fielding; Requirement on persistence? Required minimum trade between single-shot spectral coverage and field of view |
| 13.8 | Required observation planning tools (e.g. guide stars); PSF simulation tools to plan for whether PSF subtraction will be good enough to see the host galaxy |  |  |
| 13.9 | Required data reduction pipeline for IFU |  |  |

#### High-Redshift Galaxies

The requirements for the high-redshift galaxies science case are summarized in the following table (see the Galaxy Assembly and Star Formation History section of KAON 455 (V1)).

Table 14. High-Redshift Galaxies derived requirements

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 14.1 | Sensitivity. SNR ≥ 10 for a z = 2.6 galaxy in an integration time ≤ 3 hours for a spectral resolution R = 3500 with a spatial resolution of 50 mas [SCRD §2.4.4] This is not explicitly called out like this in the SCRD, but is reasonably consistent with it. The SNR of 10 is not called out. | Sufficiently high throughput and low emissivity of the AO system science path to achieve this sensitivity. Background due to emissivity less than 20% of sky + telescope. [SCRD §2.4.5.1 and SCRD Figure 6] |  |
| 14.2 | Target sample size of ≥ 200 galaxies in ≤ 3 years (assuming a target density of 4 galaxies per square arcmin) [SCRD §2.4.3] Again, not explicitly stated in this manner, but consistent with 10 targets / night over 20 observing nights. | Multi-object AO system: one DM per arm, or an upstream MCAO system correcting the entire field of regard. 6-12 arms on 5 square arc minutes patrol field. Minimum Spacing: TBD | Multiple (6-12) IFUs, deployable on the 5 square arc minute field of regard |
| 14.3 | Observing wavelengths = J, H and K (to 2.4 µm) What portions of these bands must be detected in a given single observation? [SCRD §2.4.5.2, RollUp_v1 B11] | AO system must transmit J, H, and K bands | Infrared IFUs designed for J, H, and K. |
| 14.4 | Spectral resolution = 3000 to 4000 [SCRD §2.4.5.2] |  | Spectral resolution of >3000 in IFUs |
| 14.5 | Spatial resolution 50 mas [SCRD §2.4.5.1] RollUp_v1 refers to 70mas when stating encircled energy spec. Is this consistent? | See encircled energy spec. #9.5 | 50 mas IFU spaxal scale |
| 14.6 | Encircled energy 50% in 70 mas [RollUp_v1 E11, N11] Note SCRD §2.4.5.1 seems garbled. (50% instead of 50mas (typo?)? What is meant by optimal tip/tilt stars?) Is this consistent with #9.5. Is there an EE spec for 50 mas? | Wavefront error sufficiently low (~170 nm) to achieve the stated requirement in J, H, and K bands. |  |
| 14.7 | Velocity determined to ≤ TBD km/sec for spatial resolutions of 50 to 70 mas (not specified in SCRD) | PSF intensity distribution known to ≤ 10% per spectral channel. Spatial and spectral model fitting valid to ≤ TBD (how did this get derived?) |  |
| 14.8 | Field of view ≥ 1” x 3” [SCRD §2.4.5.1] | Each MOAO channel passes a 1”x3” field plus a seeing disk diameter (~1”), assuming uncorrected field is at multiobject pickoff plane. Can be tightened appropriately if plane is partially or fully AO corrected. | Each IFU unit’s field of view is 1” x 3” |
| 14.8opt | Field of view ≥ 1” x 1” [SCRD §2.4.5.1] | Each MOAO channel passes a 1”x1” field plus a seeing disk diameter (~1”), assuming uncorrected field is at multiobject pickoff plane. Can be tightened appropriately if plane is partially or fully AO corrected. An additional MOAO channel is dedicated to sky observations. | Each IFU unit’s field of view is 1” x 1” One extra IFU for dedicated sky observations |
| 14.9 | Simultaneous sky background measurements within a radius of 3” with the same field of view as the science field | See choice of #9.8 or #9.8opt above. |  |
| 14.10 | Relative photometry to ≤ 5% for observations during a single night [RollUp_v1 H11] (left blank in SCRD) | Knowledge of ensquared energy in IFU spaxel to 5%. Telemetry system that monitors tip/tilt star Strehl and other real-time data to estimate the EE vs time. |  |
| 14.11 | Absolute photometry ≤ 0.3 mag (can’t find this requirement) |  |  |
| 14.12 | Sky coverage ≥30% to overlap with data sets from other instruments and telescopes [SCRD §2.4.5.3, RollUp_v1 N11] | Infrared tip/tilt sensors with AO correction of tip/tilt stars (specify order of correction?) |  |
| 14.13 | Should be able to center a galaxy to ≤ TBD of science field of view |  |  |
| 14.14 | Should know the relative position of the galaxy to ≤ TBD of spaxel size |  |  |
| 14.15 | Target drift should be ≤ TBD of spaxel size |  |  |
| 14.16 | The following observing preparation tools are required: TBD |  |  |
| 14.17 | The following data products are required: calibrated spectral data cube in World Coordinate System with absolute TBD accuracy |  |  |

#### Gravitational Lensing

The requirements for the gravitational lensing science case are summarized in the following two tables.  (there is no write-up on gravitational lensing in the present version of the SCRD).

Table 15a. Spectroscopic studies of distant galaxies lensed by galaxies

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 15a.1 | SNR ≥ 10 for a z = 1 – 2 galaxy in an integration time ≤ 3 hours for a Gaussian width 20 km/sec Gaussian width (50 km/sec FWHM) with a spatial resolution of 50 mas |  | R ~ 5000 (or whatever is needed to achieve 20 km/sec sigma on these targets) |
| 15a.2 | Target sample size of ≥ 50 galaxies, with density on the sky of 10 per square degree. Survey time ~ 3 years. | Number of IFUs: at least one, plus preferably one to monitor the PSF and one to monitor the sky. The extra two IFUs could be dispensed with if there were other ways to monitor the PSF and the sky background. |  |
| 15a.3 | Observing wavelengths = J, H and K (to 2.4 µm). Emphasis is on shorter wavelengths. Would use z and I bands if available. |  |  |
| 15a.4 | Spectral resolution: whatever is needed to get 20 km/sec radial velocity Gaussian sigma |  | Spectral resolution |
| 15a.5 | Spatial resolution 50 mas | [need to make the spatial resolution and the enclosed energy requirements consistent with each other] |  |
| 15a.6 | Velocity determined to ≤ 20 km/sec Gaussian sigma for spatial resolutions of 50 mas | PSF intensity distribution known to ≤ TBD% per spectral channel. |  |
| 15a.7 | Field of view > 4” diameter (lenses are bigger than galaxies). Typical lens is 3 to 4 arc sec diameter. |  |  |
| 15a.8 | Simultaneous sky background measurements |  |  |
| 15a.9 | Relative photometry to ≤ 0.1 mag for observations during a single night |  |  |
| 15a.10 | Absolute photometry ≤ 0.3 mag |  |  |
| 15a.11 | Sky coverage at least 50% with enclosed energy radius within TBD mas. Sky coverage should be better than with current LGS. |  |  |
| 15a.12 | Should be able to center a galaxy to ≤ TBD of science field of view |  |  |
| 15a.13 | Should know the relative position of the galaxy to ≤ TBD of spaxel size. (Whatever works for high z galaxy case is OK here as well) |  |  |
| 15a.14 | Target drift should be ≤ TBD of spaxel size (Whatever works for high z galaxy case is OK here as well) |  |  |
| 15a.15 | The following observing preparation tools are required: |  |  |
| 15a.16 | The following data products are required: calibrated spectral data cube |  |  |

Table 15b. Imaging studies of distant galaxies lensed by galaxies.

Goal: screen potential lensed-galaxy targets for more detailed and lengthy spectroscopic study. (there is no write-up on distant lensed galaxies in the present version of the SCRD).

| # | Science Performance Requirement | AO Derived Requirements | Instrument Requirements |
| --- | --- | --- | --- |
| 15b.1 | SNR ≥ 3 per pixel (100 per source) for a z = 1 – 2 galaxy in an integration time ≤ 1/2 hour. Strehl > 0.3 at J band. |  |  |
| 15b.2 | Target sample size of ≥ 200 galaxies, with density on the sky of 10 per square degree. Survey time ~ 3 years. | Overhead less than 10 min between targets. | 10 per square degree implies that you will only be able to observe one target at a time – average of 1 in every ~19x19’ patch. |
| 15b.3 | Observing wavelengths = J, H and K (to 2.4 µm). Emphasis is on shorter wavelengths. Could use z if available. Thermal part of K band less important. |  |  |
| 15b.4 | Spatial resolution better than 50 mas | Need a good model of the PSF or a simultaneous image of a PSF star. Need a figure of merit for goodness of the PSF: how well the model fits the “real” PSF in two dimensions. | Need Nyquist sampling of pixels at each wavelength. |
| 15b.5 | Field of view > 15” diameter for survey. Bigger is better. Some degradation between center and edge of field is tolerable. (Need to quantify.) |  |  |
| 15b.6 | Relative photometry to ≤ 0.1 mag for observations during a single night |  |  |
| 15b.7 | Absolute photometry ≤ 0.3 mag |  |  |
| 15b.8 | Sky coverage at least 50% with enclosed energy radius within 0.1 arc sec at H or K. Sky coverage should be better than with current LGS. |  |  |
| 15b.9 | Should know the relative position of the galaxy to ≤ TBD of spaxel size. (Whatever works for high z galaxy case) |  |  |
| 15b.10 | The following observing preparation tools are required: |  |  |
| 15b.11 | The following data products are required: accurate distortion map (to 1% of the size of the galaxy, or 0.01 arc sec rms) |  |  |

Notes:

May want to repeat this exercise with galaxies lensed by clusters.

Wide field: Typical size of the highly magnified region of a cluster is 1 arc min.  Need low background: lens arcs from z ~ 7 are at most Vega magnitude 23 or 24 in H (brightest arcs).  Typical size small (half light radii 0.1 arc sec).

Closer galaxies with giant arcs: deployable IFU application.  Several arc sec long.  Want field of regard of about an arc min.  Usually 3 to 5 multiple arcs within a square arc min.  (But each might be long, and require more than 1 IFU unit.)  Less than 10 IFU units needed in a square arc min.

Is a requirement needed on achieving a particular astrometric accuracy in a particular time needed for survey-type science (i.e., 1 mas in 10 min).

Aren’t encircled energy requirements needed for the IFUs?

#### Other: Backup Science

This will primarily be NGS science that can be done when the lasers cannot be propagated (e.g. due to cirrus), or less-demanding examples of LGS science that can be done when the laser power available is lower than nominal due to hardware problems.  The derived requirements for Backup Science will largely involve science preparation and operations issues.

Table 16.  Other: Backup Science derived requirements

#### Atmospheric Seeing Assumptions: (perhaps this should be an appendix?)

The atmosphere and telescope parameters assumed for achieving these numbers are summarized below.

The NGAO system shall provide its nominal performance when the atmospheric seeing is characterized by the following conditions.  An evaluation of existing seeing data has been performed (KAON 303). The KAON 303 profile was modified to include a stronger ground layer and the standard r0 value was lowered from 20 to 18 cm. The resultant baseline median Cn2 profile is presented in Table 1.  From this model we calculate the following turbulence parameters for 0.5 µm wavelength (note that r0, θ0 and 1/fG increase as λ6/5):

- Fried’s seeing parameter 	r0 = 18 cm

- Isoplanatic angle 	θ0 = 2.5 arcsec

- Turbulence characteristic frequency 	fG = 39 Hz

In addition, we have adopted a standard deviation for r0 of r0  = 3 cm with a characteristic evolution time of t = 3 min.

Table 17. NGAO baseline Mauna Kea Cn2 Profile

| Altitude (km) | Fractional C | Wind Speed (m/s) |
| --- | --- | --- |
| 0.0 | 0.471 | 6.7 |
| 2.1 | 0.184 | 13.9 |
| 4.1 | 0.107 | 20.8 |
| 6.5 | 0.085 | 29.0 |
| 9.0 | 0.038 | 29.0 |
| 12.0 | 0.093 | 29.0 |
| 14.8 | 0.023 | 29.0 |

#### Science Instrument Requirements

The NGAO system must be capable of supporting the following science instruments (in rough order of priority), based on the NGAO proposal and SCRD:

- Visible imager.  Wavelength range = 0.6 to 1.1 µm.  Field of view = 20”x20”.  Image sampling = 6 mas pixels.

- Near-IR imager.  Wavelength range = 1.0 to 2.45 µm.  Field of view = 20”x20”.  Image sampling = 10 mas.

- Deployable near-IR Integral Field Unit (IFU).  Wavelength range = 1.0 to 2.45 µm.  Field of Regard = 1.5’x1.5’ with 3x3.4” fields of view for each IFU.  Image sampling = 100 mas.

- Near-IR IFU.  Wavelength range = 1.0 to 2.45 µm.  Field of view from 2”x1.25” to 16”x5”.  Image sampling = 20 to 100 mas.

- Visible IFU.  Wavelength range = 0.6 to 1.1 µm.  Field of view from 1.2”x1.36” to 12”x6.8”.  Image sampling = 20 to 100 mas.

- L and M-band imager.  Wavelength range = 3.0 to 5.3 µm.  Field of view = 25”x25”.

- Image sampling = 25 mas.

- Future science instruments from the above list or completely new instruments.  These future science instruments would need to be designed so as to fit at a movable port or to replace a fixed first generation instrument.

Table 18. Science Instrument Requirements

| # | Science Instrument Requirement | Discussion | Based on Red means unspecified or vague |
| --- | --- | --- | --- |
| 18.1 | Visible Imager: the field of view shall be ≥ 2 arcsec (diameter?) |  | SCRD §2.1.6.2, §2.5: asteroid shapes and companions |
| 18.2 | NIR Imager: the field of view shall be ≥ 2 arcsec |  | SCRD §2.1.6.2, §2.5: Io, Titan, debris disks and QSO host galaxies |
| 18.3 | NIR Imager: the field the field of view shall be ≥ 3 arcsec |  | SCRD §2.2.5.3: planets around low mass stars |
| 18.4 | NIR Imager: the field the field of view shall be ≥ 10 x 10 arcsec |  | SCRD §2.3.11: Galactic Center |
| 18.5 | NIR Imager: shall provide a coronagraph |  | SCRD §2.2.4, RollUp_v1 G5 : planets around low mass stars, debris disks and QSO host galaxies |
| 18.6 | NIR Imager: wavelength coverage shall be at least 0.9 to 2.4 µm |  | SCRD §2.2.4, RollUp_v1 B5: planets around low mass stars |
| 18.7 | NIR IFU: field of view shall be ≥ 2 arcsec | Only 1” is required for the Galactic Center | SCRD §2.1.6.2: Asteroids, Titan |
| 18.8 | Visible IFU: field of view shall be ≥ few (specify) arcsec |  | SCRD §2.5: Nearby AGNs |
| 18.9 | Visible IFU: spectral resolution shall be R ~ 4000 | What range is acceptable? Note: Asteroid size and shape prefers R=1000 but tolerates 4000 [SCRD §2.1.6.6] | SCRD §2.3.6: Galactic center radial velocity, Nearby AGNs |
| 18.10 | NIR deployable IFU: field of view shall be ≥ few arcsec | 1”x3” or larger required for high z galaxies. (see #9.8opt for an option to this) Galactic Center needs a dIFU.don’t see this requirement for GC | SCRD §2.4.5.1: High z galaxies, gravitational lensing |
| 18.11 | NIR deployable Imager: field of view shall be ≥ few (specify) arcsec |  | SCRD §2.4.5.1:: high z galaxies (no discussion of deployable imager here, although I know we talked about it) |
| 18.12 | NIR dImager: field of view shall be ≥ 10 arcsec |  | KAON 455 (v1) Table 5: resolved stellar populations |

It is TBD whether the NGAO system will need to support any of the existing science instruments (NIRC2, NIRSPEC or OSIRIS).  The Interferometer and ‘OHANA requirements are discussed in section 6.2.4.

It is a goal for NGAO to support visitor science instruments.

### Science Operations Requirements

The top-level science operations goals for the NGAO system including the science instruments are the following (see KAON 476):

- Science-grade quality of the raw data (i.e., image quality and completeness of observations).

- Science-grade quality of the data products (i.e., photometry, astrometry, etc.).

- Science impact from a given data product (i.e., number of publications and citations).

The requirements that support these top-level goals are defined in the following tables.

#### Science-grade quality of the raw data:

Table 19. Science Operations Requirements, Raw Data Quality

| # | Science Operations Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 19.1 | Provide an extensive set of tools for instrument performance simulation and observing preparations | The requirement is that an observer should be able to prepare the observing sequences and simulate the science performance in terms of SNR, EE, SR and total observing time given a set of observing parameters and observing conditions. The suite of functions from the planning tools includes: resolve target, check observability, find suitable NGS, select and simulate instrument setup, simulate AO and instrument performance for various parameters, simulate observing sequence and efficiency, exposure time calculator, save target information, save observing setup. The tools should be designed with the end user in mind: user friendly, configurable, stand-alone and available at the observing site. | Required by most Science Cases. See e.g., #1.7, #6.16, #12.9 KAON 476, section 6.2.1 Discussions with science teams (see Observing Timeline document under Science Operations) |
| 19.2 | Document the instrument performance at an appropriate level to support observing preparations | This is both a development and operations requirement, since a continued effort will be required to characterize, monitor and document the performance. This should be based on an extensive science verification phase. | KAON 476, section 6.2.1 and 6.2.3 |
| 19.3 | Provide a semi-real-time level 1 data reduction pipeline for each instrument to at minimum perform background subtraction, cosmetic correction and shift-and-add of images. | This is an operation requirement to support # 14.2 and # 14.4 for any science instrument as well as a science requirements for IFU instruments. | KAON 476, section 6.2.1 |
| 19.4 | Provide semi-real-time tools to perform an assessment of the image quality on the level 1 data including SNR, Strehl and encircled energy. | It is required that the observing support team as well as the observer have the tools to check the image performance at the focal plane of the science instrument. | KAON 476, section 6.2.1 |
| 19.5 | Provide a science operations paradigm that optimizes the completion rate for a significant fraction (TBD) of observing programs. | It is ** not required ** to develop a plan for queue scheduling and service observing. Yet, it is required to provide the necessary tools to 1) simulate, prepare and run the observing sequences; 2) assess the science-grade quality of the data on-the-fly and subsequently 3) decide (or not) to switch to a different observing sequence, a different observing program either by the same observer or not. Particularly, a goal is to develop a science operations paradigm that allows for (and encourages?) flexible scheduling per TAC, to optimize the match between required observing conditions for a science program and real conditions. | KAON 476, section 6.2.1 |

#### Science-grade quality of the data products:

Table 20. Science Operations Requirements, Data Products Quality

| # | Science Operations Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 20.1 | Provide the required calibration methods and tools to achieve the astrometry performance requirements | This requires four important steps: 1) the tools and methods to calibrate for the astrometry are specified and designed, 2) the tools and methods are implemented, demonstrated, and optimized during the science engineering and verification phase, 3) the tools and methods are handed over to the observing support team, and 4) the performance is regularly documented and posted using the tools and methods. | KAON 476, section 6.2.2 |
| 20.2 | Provide the required calibration methods and tools to achieve the photometry performance requirements | Same as above | KAON 476, section 6.2.2 |
| 20.3 | Provide the required calibration methods and tools to achieve the PSF characterization requirements | This is in support of # 15.1, 15.2. as well as 14.1 to 3; | KAON 476, section 6.2.2 |

#### Science impact from a given data product:

Table 21. Science Operations Requirements, Archiving and Retrieval

| # | Science Operations Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 20.1 | Develop a plan for data archival | Data archival is critical for 1) science programs such as proper motion studies, transient phenomena (GRB, SN, Titan, etc), 2) the semi-automated monitoring of instrument performance and 3) the long-term visibility of Keck NGAO (either from direct science impact, or the use of the archive by the general public, e.g. CADC, google sky). It’s very likely that the archive may not be supported at first, but it’s essential to develop a plan for it and understand the requirements for the FITS information, the data storage format, etc | Science Case Requirement #1.8 KAON 476, section 6.2.3 |
| 20.2 | Develop a plan for data retrieval from the data archive | Same as above | KAON 476, section 6.2.3 |
| 20.3 | Document the on-sky science performance of each science instrument with NGAO | This is redundant with #14.2 | KAON 476, section 6.2.3 |

## O

## bservatory Overall Requirements

### Purpose and Objectives

The purpose of the overall requirements section is to convey requirements that apply generally to the overall instrument and its accessories based on the Observatory’s requirements.

Note that the Observatory’s standard requirements for all new instrumentation are summarized in the Instrumentation Baseline Requirements Document.  Additional Observatory requirements specific to NGAO are listed in the following sections.

### Facility Requirements

The following are requirements imposed by the nature of the existing facility.

Table 21. Facility Requirements

| # | Facility Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 21.1 | The NGAO system must be facility-class | Facility-class has many implications on safety, operability, reliability, maintainability, lifetime, documentation, configuration management, etc. |  |
| 21.2 | The NGAO system & science instruments should be located on the Nasmyth platform of one of the Keck telescopes | The Keck telescope foci and Nasmyth deck storage locations are already heavily utilized. The current AO systems occupy the left Nasmyth platform locations of both telescopes. HIRES occupies the right Nasmyth on Keck I while DEIMOS and NIRSPEC share the right Nasmyth on Keck II. The Cassegrain foci are occupied by LRIS (and MOSFIRE in the future) on Keck I and by ESI on Keck II. The bent Cassegrain ports are believed to have inadequate space and weight capacities. The prime focus could potentially be available but there would be many constraints on an instrument at this location. The most viable option is in the location of an existing AO system. Alternatives would be to decommission HIRES or for the existing AO system and the NGAO system to be able to share the same platform. |  |
| 21.3 | The NGAO system should accommodate the entire Keck pupil | The Keck primary has a maximum edge-to-edge diameter of 10.949 m. |  |
| 21.4 | If the existing f/15 or f/40 secondary mirrors are used then the NGAO system will be constrained by the resultant f/#, focal plane and pupil location | Both telescopes have f/15 secondary mirrors, as well as chopping secondary units that can accommodate f/25 and f/40 secondary mirrors. The choice of f/15 secondary mirrors for the current AO systems was largely driven by the resultant reduced size of the AO systems and the availability of PCS (Phasing Camera System) via a rotation of the tertiary mirror. The inability of the current systems to chop at the telescope pupil has been a limitation for thermal IR observations. The focal length of the telescope with the f/15 secondary mirror is 150 m. The 10.949 m primary corresponds to an f/13.66 beam with an exit pupil diameter of 1.460 m located 19.948 m in front of the focal plane. | Is this a facility requirement? Should it be an optical requirement |
| 21.5 | The NGAO facility must not compromise the performance of a non-NGAO instrument when that instrument is being used for science or engineering | This requirement is intended to ensure that the NGAO facility, when not in use, does not introduce vibrations or stray light that might impact the performance of another science instrument or an engineering instrument such as PCS. |  |
| 21.6 | The NGAO facility must not compromise the performance of the telescope when the telescope is used for non-NGAO observations. | The NGAO system should not impact the dynamic performance of the telescope through vibrations or different telescope dynamics. If NGAO hardware is to be mounted in the top-end then it must be designed/implemented not to compromise the secondary mirror performance. |  |
| 21.7 | The NGAO facility / science instrument combination should provide compensation for science field rotation | The Alt-Az telescope design requires compensation for field rotation in order to keep the science field fixed on the science instrument. |  |
| 21.8 | The NGAO facility / science instrument combination should provide compensation for pupil rotation | The Alt-Az telescope design and the irregular shape of the Keck primary mirror require that NGAO system provide appropriate compensation for pupil rotation. Examples: The existing Keck AO system updates the reconstructor to compensate for pupil orientation. The NIRC2 coronagraph mask rotates to match the rotating pupil. |  |
| 21.9 | The NGAO facility should provide compensation for LGS projector rotation | The Alt-Az telescope design will cause the laser projector (and the resultant LGS asterism) to rotate with respect to the Nasmyth platform. Compensation will be required to maintain the off-axis LGS on the corresponding wavefront sensor. |  |
| 21.10 | The NGAO facility should accommodate access for routine maintenance of the telescope | For example, access to the elevation journal, elevation wrap, bent Cassegrain platform and stairwell to the mirror cell |  |
| 21.11 | The NGAO facility should not routinely require more than 30 minutes of daytime telescope restriction on an NGAO science night. |  |  |
| 21.12 | The NGAO system must operate within specifications under the normal summit temperature and humidity conditions | See the conditions specified in the Instrument Baseline Requirements Document. |  |

### Telescope and Dome Environment Requirements:

The NGAO system shall provide its nominal performance when the telescope and dome environment can be characterized by the following conditions.

Table 22. Telescope and Done Environment Requirements

| # | Observatory Instrument Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 22.1 | Dome & telescope seeing less than 0.1” | The Keck dome and telescope environment degrades the intrinsic seeing by less than 0.1 arcsec, in quadrature, as measured from the effective increase in image FWHM (this change corresponds to decreasing the r0 parameter from 18 cm to 17.8 cm). |  |
| 22.2 | The phasing errors will be 10 nm rms wavefront or less before NGAO correction. | Standard performance is 60 nm rms currently. Currently available algorithms have demonstrated 10 nm rms. This error interacts with the segment figure error discussed next. We may want to place an error on the overall telescope wavefront figure PSD instead. |  |
| 22.3 | Segment figure : The wavefront error of the 36 segments will be less than 80 nm rms wavefront after warping, but before NGAO correction. | This number is an average over all 36 segments segment. As a goal the wavefront error shall be 80 nm rms over each segment. |  |
| 22.4 | Stacking: The segment stacking errors will contribute less than 20 nm rms wavefront to the overall wavefront error before NGAO correction. |  |  |
| 22.5 | Line of sight jitter: The aggregate line of sight jitter (wavefront tip and tilt) resulting from motion of the primary, secondary and tertiary mirrors will be less than 0.020 arc seconds rms before correction by the NGAO. | This vibration is known to currently be largely dominated by a narrow resonance at ~29 Hz. |  |
| 22.6 | Segment motion: The motion of each segment as a solid body will be less than 0.015 arc seconds rms before correction by the NGAO. | This vibration is known to currently be largely in a narrow resonance at ~29 Hz. |  |

### Observatory Science Instrument Requirements

In addition to the science instrument requirements specified in section 6.1.2.18 the NGAO facility must allow the Observatory to continue supporting Interferometer science with the two Keck telescopes. Should we move this paragraph to section 6.1..3 then? The requirements for these instruments are developed in KAON 428 and are further specified in later sections.

Table 23. Observatory Science Instrument Requirements

| # | Observatory Instrument Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 23.1 | The NGAO facility should support the Keck Interferometer (KI) with performance as good or better than provided by the pre-NGAO Keck AO systems | The KI dual star modules (DSM) currently move into both AO enclosures on rails to feed the KI. The requirement to feed the KI requires that collimated and f/15 light can be fed to the DSM and that the field rotation, pupil rotation, longitudinal dispersion and polarization from the NGAO system and the AO system on the other telescope be identical. See KAON 428 Implications and Requirements for Interferometry with NGAO. | Same idea: should these requirements be part of the Science Instrument Requirements (6.1.3)? |
| 23.2 | The NGAO system should support the ‘OHANA interferometer with performance as good or better than provided by the pre-NGAO Keck AO systems | Injection modules are currently placed on each AO bench to feed an optical fiber that goes to the KI. In future the output from these fibers will be interfered with those from multiple telescopes. | Same as above |

### Observatory Operational Requirements

The purpose of this section is to document the Observatory’s overall requirements for the support of science operations.  These requirements can be divided into the following categories:

- Percent of time collecting science quality data.

- Capability to support a certain number of nights per year of science operations.

- Operational costs.

- Impact on daytime and nighttime operations.

- Compliance with regulations, including safety, Mauna Kea policies, FAA and U.S. Space Command.

Each of these categories is represented in one of the following tables.  In addition, all of the above categories require a facility-class NGAO system and science instruments, and this requirement is reflected in the final table.

Table 24. Observatory Operational Requirements

| # | Observatory Operational Requirement | AO Derived Requirements | Instrument Derived Requirements |
| --- | --- | --- | --- |
| 24.1 | Assuming a classical observing model and adequate observing conditions for the science program, more than 80% of the observing time is spent on collecting science-quality data for the deployable science instruments. | Observing model should allow flexible scheduling and quick-and-easy switching of observing modes and instrument. Observing overhead are minimized, particularly by being able to center multiple targets and reference guide stars with a < 0.3” accuracy. Observing reliability is critical to avoid canceling long integration: there should be less than 2 faults per night (one system fault is equivalent ~ 20-30 min lost time on faint targets for the worst cases). Observing tools should allow for automated observing sequences. | The deployable instrument are used for faint object spectroscopy with much longer time on target, hence less observing setup overhead. One should be able to stop, abort and restart an integration on command. |
| 24.2 | Assuming a classical observing model and adequate observing conditions for the science program, more than 70% of the observing time is spent on collecting science-quality data for the narrow-field science instruments. | Same as above. In addition, observing efficiency is critical for survey mode (see Science Case 1) Acquisition sequences must be fully automated. Overhead must be reduced by allowing parallel sequences between telescope, AO, laser and science instrument. | The narrow field science instruments are used on brighter targets in imaging, coronagraphy and spectroscopy modes. The science instrument interface command must allow for parallel sequences, e.g, start a nod/dither sequence immediately after the pixel readout. |
| 24.3 | The NGAO system must be capable of supporting 200 nights/year of science operations and keep the total annual operational personnel within the 5-year plan Observatory operation budget, including non-personnel costs. | Observatory currently willing to support 140 nights/year. |  |
| 24.4 | The NGAO facility should not require more than TBD engineering nights per year for system maintenance. |  |  |
| 24.5 | The Mauna Kea laser projection requirements must be satisfied | This includes requirements on laser power, wavelength, laser traffic control participation, aircraft safety and space command. See KAON 153. The current policy only accepts sodium wavelength lasers, and requires that a single laser beacon not exceed 50W and that a maximum of 200 W be projected from a single facility, and that laser beacons not be projected below 70º zenith angle. | Mauna Kea laser projection requirement |

Discussion: Should this efficiency number include any of the weather statistic? The Keck Strategic Planning document calls for a 90% efficiency goals in 2020.

### Observatory Implementation Requirements

Table 25. Observatory Implementation Requirements

| # | Observatory Implementation Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 25.1 | The NGAO system and instruments must complete the Observatory standard design review process. |  |  |
| 25.2 | The time between decommissioning an AO capability on the telescope where the NGAO system is to be installed and making NGAO available for limited shared-risk science must be agreed upon with the SSC. (it used to be: must not be longer than 6 months) | Down time impact on Interferometer and AO science at the Observatory. |  |
| 25.3 | The telescope downtime required to implement NGAO must not be longer than 5 days |  | This should be agreed on and should be retrofitted in the Instrument Baseline Requirements |
| 25.4 | The NGAO system must complete an operations transition review where operational responsibility is transferred from development to operations | This has implications on defining transition requirements and on training. |  |

# Optical Requirements

## Purpose and Objectives

The purpose of this section is to describe optical requirements for the performance, implementation and design of the NGAO optical system.

## Performance Requirements

The following performance requirements are duplicated from the Science Performance Requirements in section 6.1.2 since these are direct optical performance requirements.

The following performance requirements are derived from the Science Performance Requirements in Section 6.1.2 and the relevant performance budgets.

## Implementation Requirements

| # | Optical Implementation Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | The NGAO optical axis shall be coincident to the telescope’s elevation axis to ≤ TBD. | This is intended to minimize pupil and image motion. |  |

## Design Requirements

NGAO is required to provide an AO corrected beam to each of the science instruments, including the Interferometer and ‘OHANA fiber injection module.  This may be accomplished with one or more AO systems.  The following requirements are valid for all of these science instruments.

| # | Optical Design Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | The focal ratio provided to the science instruments should be TBD |  |  |
| 2 | The exit pupil location provided to the science instruments should be TBD |  |  |
| 3 | The NGAO system and science instrument combination should be capable of keeping the field or pupil fixed on the science instrument. |  |  |
| 4 | The NGAO system shall have ≤ TBD of non-common path aberration delivered to the science instruments. |  |  |
| 5 | The NGAO system should be capable of correcting ≥ TBD nm of low spatial frequency (Zernikes 4 to 15) non-common path aberration in the science instruments. | Image sharpening can be used to correct for aberrations in the science instruments. New science instruments should be designed to have small optical aberrations. The interferometer and possible legacy instruments such as NIRC2 or OSIRIS may be allowed larger aberration budgets. |  |
| 6 | The peak-to-peak range of tip/tilt correction provided by NGAO shall be ≥ 3” on sky. | The existing Keck AO systems have a peak-to-peak range of 1.6” which has proven to be inadequate in windy conditions. |  |

The following design requirements are duplicated from the Science Performance Requirements in section 6.1.2 since these are direct optical performance requirements.

| # | Science Instrument Optical Design Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | Unvignetted contiguous fields shall be provided to the NIR and visible science imagers and single field IFUs. These fields should be centered to within 1” of the telescope’s optical axis. The maximum field size is 20”x20”. | Driven by the science requirements on the contiguous field science instruments. |  |
| 2 | Multiple unvignetted contiguous fields shall be provided to the NIR d-IFU. Each field should be ≥ 4” in diameter. | Driven by science requirements on the NIR d-IFU. |  |
| 3 | The unvignetted field of regard provided to the NIR d-IFU shall have a total area of ≥ 6 arcmin2 and shall have a maximum off-axis distance of ≤ 1.5’ with respect to the telescope’s optical axis. | Driven by science requirements on the NIR d-IFU. |  |
| 4 | NGAO shall provide appropriate outputs to ≥ TBD science instruments with a switching time of ≤ 5 min. |  |  |

The following design requirements are imposed by the non-interferometric science instruments.

| # | Science Instrument Optical Design Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | A wavelength range of 0.7 to 1.0 µm must be provided to the visible science instruments (imager and IFU) |  |  |
| 2 | A wavelength range of 1.0 to 2.45 µm must be provided to the NIR science instruments (imager, IFU and deployable IFU) |  |  |
| 3 | A wavelength range of 3.0 to 5.3 µm must be provided to the thermal NIR imager | This is currently low priority and should be discussed if this drives the design. |  |
| 4 | An unvignetted field of view ≥ 20”x20” must be provided to the science imagers (visible and NIR) | Should this only be 10x10” for the visible imager? |  |
| 5 | An unvignetted field of view ≥ 12” diameter must be provided to the visible IFU |  |  |
| 6 | An unvignetted field of view ≥ 16” diameter must be provided to the NIR IFU |  |  |
| 7 | A field of regard of ≥ 1.5’ diameter must be provided to the NIR deployable IFU with vignetting by all sources ≤ TBD over TBD % of the field of regard |  |  |
| 8 | A field of view of 25”x25” must be provided to the L and M-band imager | Low priority. |  |
|  | An unvignetted science target field ≥ 3” x 3.4” shall be provided to each channel of the NIR deployable IFU |  |  |
|  | The NGAO + deployable IFU system shall support simultaneous observations of at least two science targets separated by ≤ 5” | There is no requirement for simultaneous nearest-neighbor target distances less than 5”. |  |
|  | The NGAO + deployable IFU system shall support simultaneous observations of at least six science target fields inscribed within a 30” diameter field |  |  |
|  | NGAO shall be capable of compensating for focus changes due to changing filters or modes |  |  |

The following design requirements are imposed by the Interferometer and/or ‘OHANA support requirements.

| # | Interferometry Optical Design Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | A wavelength range of 1.1 to 14 µm must be provided to the Interferometer |  |  |
| 2 | A wavelength range of 1.1 to 2.45 µm must be provided to the OHANA injection module |  |  |
| 3 | A field of view of ≥ 1’ diameter must be provided to the Interferometer |  |  |
| 4 | A field of view of 5” diameter must be provided to the OHANA injection module |  |  |
| 5 | NGAO must be able to support a chopping mode for the interferometer. | The nuller requires small amplitude chopping with the AO loops closed at each end of the chop for alignment purposes. |  |
| 6 | The interferometer output of NGAO must be polarization matched to the interferometer output of the AO system on the other telescope in order to produce ≤ 3º of differential s-p phase shift | The current KI achieves polarization matching by keeping the number, angle and coatings of all reflections the same in the beam trains from each telescope. The differential s-p phase shift in the current KI is measured at 6º resulting in a loss in V2 of 0.003. |  |
| 7 | The interferometer output of NGAO must have the same image rotation as the interferometer output of the AO system on the other telescope |  |  |
| 8 | The interferometer output of NGAO must have the same pupil rotation as the interferometer output of the AO system on the other telescope |  |  |
| 9 | The interferometer output of NGAO must have the same longitudinal chromatic dispersion as the interferometer output of the AO system on the other telescope | Transmissive optics fabricated from different materials can have different amounts of longitudinal chromatic dispersion resulting in the loss of fringe visibility |  |
| 10 | The ratio of the Strehls from the interferometer output of NGAO and the interferometer output of the AO system on the other telescope must be ≤ 1.2 and ≥ 0.9. | A Strehl mismatch of 22% or an intensity ratio of 1.22 results in a V2 loss of 0.010. |  |
| 11 | NGAO must be able to accommodate the accelerometers needed to support the Interferometer | On the current AO bench one accelerometer is placed near the telescope focus and another near the output to the DSM. These are used to measure vibration along the optical path and are used in the fringe tracker control system. The accelerometer acquisition system is housed in an electronics rack in the AO electronics vault. |  |
| 12 | NGAO or NGAO in combination with a modified DSM must be capable of supporting the laser metrology beams from the interferometer | These metrology beams are a potential source of background light on the wavefront sensors |  |
| 13 | NGAO must incorporate the required tools and tolerances to support alignment to the interferometer | For example, the current AO bench hosts a corner cube to aid in aligning the interferometer to the optical axis of the AO system and telescope |  |
| 14 | NGAO or NGAO in combination with a modified DSM must provide a collimated 100 mm diameter beam to the interferometer | In the current AO system a removable (on a translation stage) dichroic beamsplitter, located between the deformable mirror and second off-axis parabola, folds the collimated beam to the DSM |  |
| 15 | The rms residual tilt at the NGAO system output to the interferometer should be ≤ 0.007” for TBD guide star. |  |  |

# Mechanical Requirements

## Purpose and Objectives

The purpose of this section is to describe mechanical requirements for the performance, implementation and design of the NGAO mechanical systems.

## Performance Requirements

| # | Mechanical Performance Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | The NGAO AO system shall not exceed a thermal dissipation budget, into the dome environment, of 100 W |  |  |
| 2 | The NGAO laser system shall not exceed a thermal dissipation budget, into the dome environment, of 100 W |  |  |
| 3 | The NGAO AO system shall not exceed a thermal dissipation budget at the top-end of the telescope, into the dome environment, of 50 W |  |  |

## Implementation Requirements

| # | Mechanical Implementation Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | The NGAO facility must allow a means to install new science instruments | For example, if a new instrument can only be installed with the overhead crane then a means of removing the appropriate portion of the AO enclosure roof would be required. |  |

## Design Requirements

| # | Mechanical Design Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | The maximum weight of the AO system on the Nasmyth platform shall not exceed 10,000 kg | The weight requirements are imposed by limits on what the telescope can support at various locations without changing its performance. |  |
| 2 | The maximum weight of the laser facility on the azimuth rotating part of the telescope shall not exceed 10,000 kg |  |  |
| 3 | The maximum weight of the beam transport system on the elevation portion of the telescope shall not exceed 150 kg |  |  |
| 4 | The maximum weight of the laser launch facility in the top-end module shall not exceed 150 kg |  |  |
| 5 | If mounted behind the f/15 secondary mirror, the launch telescope facility must allow for the removal, storage and installation of the f/15 secondary module |  |  |
| 6 | If mounted behind the f/15 secondary mirror, the launch telescope facility must not extend beyond the module in the x,y-directions and must not extend more than 1 m beyond the top of the telescope structure. |  |  |
| 7 | The NGAO facility must fit within the mechanical constraints of a Nasmyth platform (nominally the Keck II left Nasmyth platform) |  |  |
| 8 | The NGAO facility must provide access and space for the installation of each science instrument and a mechanical interface on which to mount each instrument |  |  |
| 9 | The NGAO facility must provide access for routine maintenance of the elevation bearing, elevation wrap, bent Cassegrain platform and mirror cell stairwell |  |  |
| 10 | The required glycol flow rate and pressure for cooling the NGAO facility shall not exceed TBD and TBD, respectively |  |  |

# Electronic/Electrical Requirements

## Purpose and Objectives

The purpose of this section is to describe electronic and electrical requirements for the performance, implementation and design of the NGAO electronic and electrical systems.

## Performance Requirements

## Implementation Requirements

## Design Requirements

| # | Electrical Performance Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | The entire NGAO facility must not exceed a total electrical power requirement of 30 kW |  |  |
| 2 | The NGAO facility cabling through the azimuth wrap must not require an area of more than TBD |  |  |
| 3 | The NGAO facility cabling through the left Nasmyth elevation wrap must not exceed an area of more than TBD |  |  |
| 4 | The NGAO facility cabling through the right Nasmyth elevation wrap must not exceed an area of more than TBD |  |  |

# Safety Requirements

## Purpose and Objectives

Safety is the paramount concern for all activities at the observatory.  The purpose of this section is to provide requirements related to specific safety concerns during the operation and handling of NGAO.

## Scope

The general Observatory safety requirements that are also applicable to NGAO are already contained in the Instrumentation Baseline Requirements Document.  This section covers the additional laser safety and laser projection safety requirements.

Adequate earthquake restraints are also required for all systems.

## Laser Safety Requirements

A safety system will be implemented to ensure the safe use of the laser.  This safety system will include both engineering and administrative/procedural controls to assure safe operations.  The system will apply ANSI Z136.1 and Z136.6 standards for safe use of laser for indoor and outdoor.

The NGAO system will conform to OSHA and local codes in addition to codes specified for each subsystem.

## Laser Projection Safety Requirements

This section covers the additional requirements on safety for projecting the laser beams outside the dome.

### Aircraft Safety

An aircraft safety system compliant with FAA requirements must be implemented and approved by the FAA.

### Space Command

A system must be implemented to facilitate effective communication with U.S. Space Command of projection dates and targets, and to ensure that no projection occurs on a target without Space Command approval.  This requirement may be deleted if it is determined that Space Command approval is no longer required.

# Software Requirements

## Purpose and Objectives

The software requirements section describes requirements for performance, implementation and design.

## Scope

Unless otherwise indicated all of the requirements of this section apply to all software components of NGAO.

## Performance Requirements

## Implementation Requirements

## Design Requirements

| # | Software Design Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | The NGAO system must be able to support the external interfaces supported by the existing AO systems | The external keywords and EPICS channels used by the existing AO systems are documented in KAON 315, Summary of External Interfaces in the Current WFC and Implications for the NGWFC Design” |  |
| 2 | The NGAO system must be able to accept tip/tilt offloads from the science instrument |  |  |
| 3 | The NGAO system must be able to acquire and configure the AO system based on the star magnitude and spectral type provided in the target list |  |  |
| 4 | NGAO must be able to support a chopping mode for the Nuller mode of the interferometer |  |  |

# Interface Requirements

## Purpose and Objectives

This section is reserved for interface requirements that are not addressed by other portions of the document.

## Performance Requirements

## Implementation Requirements

## Design Requirements

### Optical Interface

### Mechanical Interface

| # | Mechanical Interface Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | An agreed upon kinematic interface between the NGAO opto-mechanical systems and the telescope structure must be provided | Kinematic interfaces are proposed so that thermal changes do not distort optical benches. |  |
| 2 | NGAO must be compatible with the interferometer dual star module (DSM) or replicate its functionality | The current Keck AO systems interface to the interferometer via an opto-mechanical system known as the DSM. |  |
| 3 | An agreed upon mechanical interface between the NGAO electronics and the telescope must be provided |  |  |
| 4 | An agreed upon mechanical interface between any NGAO enclosures and the telescope must be provided. |  |  |
| 5 | An agreed upon mechanical interface between any NGAO glycol cooled systems and the telescope instrument and/or facility glycol systems must be provided |  |  |
| 6 | An agreed upon mechanical interface between any NGAO CCR-cooled systems must be provided |  |  |

### Electrical/Electronic Interface

| # | Electrical Interface Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | An agreed upon electrical interface for power between NGAO systems and the Observatory /telescope must be provided | The Observatory is responsible for providing power to the various NGAO system locations. |  |
| 2 | An agreed upon interface for communication between NGAO systems and the Observatory/telescope must be provided | The Observatory is responsible for the implementation of all cables between NGAO system locations. The Observatory must in particular approve all cabling required to go in the elevation or azimuth wraps. |  |

### Software Interface

| # | Electrical Interface Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | An interface should be provided to offload tip/tilt errors to telescope pointing through the telescope drive and control system (DCS) |  |  |
| 2 | An interface should be provided to offload focus errors to the secondary mirror piston through DCS |  |  |
| 3 | An interface should be provided to offload coma errors to the secondary mirror tilt through DCS |  |  |
| 4 | An interface should be provided to offload segment stacking errors to the active control system (ACS) |  |  |
| 5 | An interface should be provided to write all appropriate NGAO keywords to the science instrument FITS header |  |  |
| 6 | An interface should be provided to allow NGAO to read/write NGAO science instrument keywords |  |  |

# Reliability Requirements

## Purpose

A process should take place to confirm that the NGAO system will provide a high level of reliability for a 10 year lifetime.

## Scope

Unless otherwise indicated all of the requirements of this section apply to all components of NGAO.

## Performance

System downtime should be minimized by a combination of component reliability, ease of repair, maintenance and appropriate sparing.

| # | Reliability Performance Requirement | Discussion | Based on |
| --- | --- | --- | --- |
| 1 | ≤ 5% of observing time lost to problems | This includes any loss to an exposure in progress and the time to start the next exposure after recovering from a fault |  |
| 2 | The median time between faults during observing time should be ≥ 3 hrs | Frequent short duration faults are not acceptable since they have a high impact on science productivity |  |

# Spares Requirements

TBD

# Service and Maintenance Requirements

TBD

# Documentation Requirements

The documentation requirements are defined in the Instrumentation Baseline Requirements Document.

## Documentation Package

## Drawings

## Electrical/Electronic Documentation

## Software

# Glossary

Table 2 defines the acronyms and specialized terms used in this document.

Table 2 Glossary of Terms

| Term | Definition |
| --- | --- |
| ACS | Active Control System |
| ANSI | American National Standards Institute |
| AO | Adaptive Optics |
| DCS | Drive and Control System |
| DSM | Dual Star Module |
| FAA | Federal Aviation Administration |
| FOV | Field Of View |
| FWHM | Full Width at Half Maximum. |
| IFU | Integral Field Unit |
| KAON | Keck Adaptive Optics Note |
| KI | Keck Interferometer |
| LGS | Laser Guide Star |
| MTBF | Mean Time Between Failures |
| NGAO | Next Generation Adaptive Optics |
| NGS | Natural Guide Star |
| NIR | Near InfraRed |
| NIRC2 | NIR Camera 2 |
| NIRSPEC | NIR SPECtrometer |
| ‘OHANA | Optical Hawaiian Array for Nanoradian Astronomy |
| OSHA | Occupational Safety and Health Administration |
| OSIRIS | OH-Suppression InfraRed Integral field Spectrograph |
| TBC | To Be Completed |
| TBD | To Be Determined |
| WMKO | W. M. Keck Observatory |