Keck Adaptive Optics Note 971

Near-Infrared Tip-Tilt Sensor System:

Pre-Ship Review Report

R. Dekany, R. Smith, D. Hale, J. Fucik, H. Rodriguez, J. Cromer

July 29, 2013

DRAFT Version 0.12

# 1

# Optical Performance

# 1.1

# “As-Built” Optical Design

Figure 1: NIRTTS Camera Optical Layout.  FL1 and FL2 are assumed to be operating at a temperature of 273K at 0.6atm.  C1 is assumed to be at an operating temperature of 100K in vacuum.  C2, C3, and C4 are assumed to be at an operating temperature of 110K in vacuum.  There are three element groups in the optical prescription:  G1 is the field lens group that reforms the Adaptive Optics (AO) System pupil onto the back surface of the filter, G2 is the air space triplet camera that re-images the AO image plane that is 100mm behind the vertex of FL1 onto the Teledyne H2RG HgCdTe detector, and G3 is the field flattener that corrects for the field curvature in G2.

The NIRTTS camera is an f/7 diffraction limited re-imager that operates in the H (1.5μm – 1.8μm) and Ks (2.0μm – 2.3μm) Infra-Red (IR) photometric bands.  The nominal camera design gives a mean Strehl ratio of 92% over the 120” Field of View (FOV) and >85% for an optical system aligned to better than ±25μm for lateral/axial displacements and ±2.4’ for angular displacements.

# The camera is composed of three lens groups: field lens (G1), air-spaced triplet camera (G2), and field flattener (G3) with a total of six refractive optical elements with power. Figure 1 shows the optical layout of the NIRTTS camera where the optical elements are all spherical surfaces manufactured from Barium Fluoride (BaF2) or Infrasil fused silica glass material.  The change in the index of refraction and radius of curvature as a function of temperature for the various materials was accounted for in the design of the camera, which is detailed in KAON890.

# 1.1.1  ISP Optics Manufacturer’s Report

# The six-element camera optics was polished and Anti-Reflection (AR) coated by ISP Optics Corporation.  Table 1 gives the Quality Assurance (QA) inspection report supplied by the vendor where the measurements where made in lab ambient conditions (i.e. T=293K and P=1atm).  Figure 2 shows the as measured AR coating reflectance curves for both the fused silica and BaF2 substrates.  The fused silica optical elements have an average reflectance of 0.3% and the BaF2 optical elements have an average reflectance of 0.4% across the instrument passband of 1490nm to 2310nm.  Both AR coatings meet the requirement of the reflectance to be less than 0.5%.  Both coatings were measured at room temperature, but it is expected that the edges of the AR coating will shift 60nm towards the red part of the spectrum at cryogenic temperatures.  This shift in the AR coating edges at operational temperatures was accounted for in the design by oversizing the passband of the AR coating.

Table 1:  ISP Optics QA Inspection Report Summary

| #### Element | Radius [mm] | CT [mm] | ET [mm] | Dia. [mm] | Glass | Wedge [arcsec] | Irregularity PV / RMS [λ@633nm] | RMS Roughness [nm] | Scatch/Dig |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| FL1 | S1: 288.615CX S2: 8361.572CX | 20.058 | 14.784 | 108.030 | BaF2 | 19.1 | S1: 0.23 / 0.03 S2: 0.50 / NaN | NaN | 40/20 |
| FL2 | S1: 178.411CX S2: 411.774CC | 15.134 | 10.058 | 108.026 | BaF2 | 19.1 | S1: 0.19 / 0.03 S2: 0.30 / 0.06 | S1: 4.5 S2: 2.5 | 40/20 |
| C1 | S1: 46.975CX S2: 355.217CC | 6.454 | 3.909 | 32.090 | BaF2 | 32.1 | S1: 0.21 / 0.02 S2: 0.15 / 0.03 | S1: 2.9 S2: 2.5 | 40/20 |
| C2 | S1: 139.337CC S2: 39.578CX | 13.056 | 8.855 | 38.204 | BaF2 | 43.2 | S1: 0.17 / 0.02 S2: 0.10 / 0.02 | S1: 4.3 S2: 1.7 | 40/20 |
| C3 | S1: 24.727CC S2: 34.535CX | 13.004 | 9.801 | 41.482 | Infrasil | 49.7 | S1: 0.06 / 0.01 S2: 0.15 / 0.03 | S1: 1.3 S2: 1.3 | 40/20 |
| C4 | S1: 41.686CC S2: 162.924CC | 9.834 | 17.623 | 45.990 | Infrasil | 26.9 | S1: 0.15 / 0.04 S2: 0.15 / 0.03 | S1: 1.5 S2: 1.9 | 40/20 |

Figure 2:  ISP Optics Corporation AR coating reflectance curves for fused silica and BaF2 substrates.  Both AR reflectance curves were measured at 6° Angle Of Incidence (AOI) at room temperature.  The edges of the AR reflectance curve are estimated to shift about 60nm to the red part of the spectrum at cryogenic temperatures, but this was accounted for in the coating design for which the AR coating passband was oversized.  The fused silica AR coating has an average reflectance of 0.3% and the BaF2 AR coating has an average reflectance of 0.4% across the instrument passband of 1490nm to 2310nm.

# 1.1.2  Asahi Filter Manufacturer’s Report

# Asahi Spectra supplied the two IR photometric filters, which are multilayer dielectric coatings on an Infrasil fused silica glass substrate that is 38mm in diameter and 5.0mm thick.  The H band IR photometric filter has a Full Width of Half Max (FWHM) of 287nm and a Center Wavelength (CWL) of 1634nm with an average transmission of 92% over the passband.  Similarly, the Ks band IR photometric filter has a FWHM of 320nm and a CWL of 2150nm with an average transmission of 94% over the passband.

# Figure 3:  Asahi Spectra H and Ks IR Photometric Filters Transmission Curves.  The dashed line curve is the atmospheric transmission at Mauna Kea at 1.6mm precipitable water vapor at 1.5 airmass.

# 1.1.3  Off-the-Shelf Optical Components

# There are two major off the shelf optical components namely the instrument window and the fold mirror.  The instrument window substrate was purchased from CVI Melles Griot part number PW-2025-INF, which has since been discontinued.  The window is an Infrasil fused silica substrate that is 50.8mm in diameter and 6.35mm in thickness.  It has a 10-5 scratch/dig surface quality with a λ/10 transmitted wavefront error at 633nm and less than 10” of wedge.  The substrate was AR coated by ISP Optics Corporation with the same coating used for the Infrasil camera elements in Section 1.1.1.  A spare window substrate was purchased and coated for the NIRTTS project.  The fold mirror was purchased from Edmunds Optics part number NT48-177-566.  The fold mirror is a Zerodur substrate that is 76.2mm in diameter and 19.1mm in thickness.  It has a 60-40 scratch/dig surface quality with a λ/20 wavefront error at 633nm.  The fold mirror has a protected (i.e. Silicon Monoxide overcoat) gold metallic mirror coating with a 97% average reflectance from 800nm to 3000nm.

# 1.1.4  Lens Prescription

# After receiving the “as-built” optical surfaces from ISP Optics Corporation the values in Table 1 were inserted into the Zemax optical model and then the air spaces were re-optimize to deliver a maximum Strehl ratio over the FOV of the camera.  Since all the optical surfaces were manufactured to within their allowed manufacturing tolerances, no significant changes were made to the optical model from the prescription delivered at the Detail Design Review (DDR).  Table 2 gives the optical prescription of the camera at operational temperatures for the “as-built” design, where the lens elements in Table 1 have been scaled for temperatures of 100K-110K.  Figure 3 shows the performance of the nominally aligned camera at operating temperature in the H and Ks bands when the camera is behind the Keck AO system.

Table 2: NIRTTS Camera Optical Prescription at Operating Temperatures

| #### Element [#] | #### Name | #### Radius [mm] | #### CT [mm] | #### ET [mm] | #### Dia. [mm] | #### Material |
| --- | --- | --- | --- | --- | --- | --- |
| - 1 | - Field Lens 1 | #### 288.510 CX | #### 20.051 | 14.779 | #### 107.991 | #### BaF2 |
|  |  | #### 8358.537 CX | #### 12.829 | #### 21.372 | #### 107.991 | #### Al 6061 |
| - 2 | - Field Lens 2 | #### 178.346 CX | #### 15.129 | #### 10.055 | #### 107.987 | #### BaF2 |
|  |  | #### 411.625 CC | #### 190.000 | #### 186.705 | #### 103.958 | #### Al 6061 |
| - 3 | - Fold Mirror | #### Plano | #### 64.000 | #### 64.000 | #### 76.200 | #### Al 6061 |
| - 4 | - Window | #### Plano | #### 6.400 | #### 6.400 | #### 50.000 | Infrasil |
|  |  | #### Plano | #### 16.603 | #### 19.419 | #### 50.000 | #### Al 6061 |
| - 5 | - Camera 1 | #### 46.836 CX | #### 6.435 | #### 3.897 | #### 31.996 | #### BaF2 |
|  |  | #### 354.162 CC | #### 11.802a | #### 11.523c | #### 28.200 | #### SS 17-4 GP1 |
| - 6 | - Filter | #### Plano | #### 5.000 | #### 5.000 | #### 32.000 | #### Infrasil |
|  |  | #### Plano | #### 10.202b | #### 9.491d | #### 32.000 | #### Al 6061 |
| - 7 | - Camera 2 | #### 138.938 CC | #### 13.019 | #### 8.829 | #### 28.084 | #### BaF2 |
|  |  | #### 39.464 CX | #### 2.649 | #### 3.830 | #### 38.094 | #### SS 17-4 GP1 |
| - 8 | - Camera 3 | #### 24.727 CC | #### 13.004 | #### 9.801 | #### 26.084 | #### Infrasil |
|  |  | #### 34.535 CX | #### 98.568 | #### 99.192 | #### 41.482 | #### Invar |
| - 9 | Camera 4 | 41.686 CC | #### 9.834 | #### 17.623 | #### 44.060 | #### Infrasil |
|  |  | #### 162.924 CC | #### BFL* | #### 3.232 | #### 43.990 | #### Invar |
| * BFL = 4.724 |  |  |  |  |  |  |
| a = 13.002 in K band |  |  |  |  |  |  |
| b = 9.002 in K band |  |  |  |  |  |  |
| c = 12.723 in K band |  |  |  |  |  |  |
| d = 8.290 in K band |  |  |  |  |  |  |
| NOTE: Physical outer diameter of Camera 4 is 45.990 |  |  |  |  |  |  |

Figure 3:  Nominal Prescription Strehl Ratio Map when the NIRTTS Camera is behind the Keck AO system.  (Left Panel)  Gives the H band Strehl map as seen by the Teledyne H2RG detector with a peak of 95.5% and an rms over the FOV of 90.6%.  (Right Panel) Gives the Ks band Strehl map as seen by the detector with a peak of 98.6% and an rms over the FOV of 94.0%.

# 1.1.5  Roll-Pin Flexure Cryogenic Lens Mount

# One of the major changes in the design since DDR was the lens mount design for the BaF2 lens elements at cryogenic temperatures.  Do to uncertainties in understanding the stresses involved in a lens element that is bonded to a metal mechanical cell with an epoxy joint the project decided that it would be less risky to go with an all mechanical mount design for the relatively small (i.e. ~50mm in diameter) optical elements.  The design of the mount was an adaptation from the cryogenic mounts used in the WIYN High Resolution Infrared Camera (WHIRC) (Smee 2011).  The detail design of the cryogenic lens mount as it is implemented in NIRTTS can be found in KAONTBD.  Solid Concepts using a Direct Metal Laser Sintering (DMLS) manufacturing method, which is a 3D printing of metal materials technology, fabricated the lens mounts.  DMLS manufacturing is a form of rapid prototyping of parts, which saved both in cost and time.  Also, do to the complex geometry of the mount flexures it was a suited candidate for the DMLS manufacturing method that could otherwise not be accomplished with traditional CNC fabrication methods.  It should be noted that the cells could have been manufactured with a combination of CNC fabrication and wire EDM manufacturing for the flexures themselves.  Currently, wire EDM manufacturing can achieve 2.5μm feature resolution, but at a higher cost and schedule time.  However, the 25μm feature resolution of the DMLS process was accurate enough for this particular design implementation.

# After receiving the parts from Solid Concepts a few post manufacturing steps were taken prior to installation of the lens element. Since the DMLS is an additive process that is based on sintering fine powdered layers that are 30μm thick together it creates a part with a surface roughness of 15-40μm P-V, which is equivalent to a 320 – 600 grit.  The surfaces that come into contact with the lens element were polished to avoid scratching or chipping of the brittle BaF2 glass material during installation.  Also, the lens cells were designed to give a particular preload force onto the lens element, which is achieved by an interference fit between the mechanical cell and lens element.  In this particular case a 50μm radial interference fit was designed, such that the lens would be constrained from moving more than the required 25μm and the stress imparted to the lens be less than the required 5MPa.  However it is almost impossible to manufacture (by any method) the distance between two opposing springs to within 25% of the designed interference fit do to the uncertainties of how the spring will relax post machining.  For DMLS manufacturing, the relaxation of the spring is caused by the localized heating and cooling of the metal material on a thin wall section that is 380μm - 630μm thick.  In practiced it was found that the radial roll pin springs had to be pushed back in order to achieve the desired preload force.  A custom mandrel was fabricated with a shaft diameter 100μm greater than the as fabricated lens element diameter.  The mandrel was then inserted into the roll pin flexure mount and by turning the mandrel in the cell it burnishes the roll pin flexures.  The amount of burnishing was done experimentally till the desired spring separation between opposing springs was achieved for the designed interference contact between the cell and lens.

# In order for the mechanical lens mounts to achieve the desired positional accuracy in the presence of frictional forces a low coefficient of friction coating needed to be added to the surfaces of the mount that come into contact with the lens (see KAONTBD for more detail).  The mechanical lens cells were fabricated from a 17-4 Stainless Steel alloy, which limited the coating options.  The WHIRC lens mount design (Smee 2011) was fabricated from an Aluminum alloy, which used a black Teflon impregnated anodized coating.  An anodization process is not possible with a Stainless Steel alloy material, but could be coated with a Polytetrafluoroethylene (PTFE) (i.e. Teflon) material, which would be vacuum compatible.  For cost and schedule reasons, the project decided to use Kapton tape with an acrylic adhesive that has a thickness of 75μm.  The Kapton tape provided two technical advantages over the PTFE coating, which was a lower coefficient of friction and uniform thickness.  Kapton tape comes in two varieties:  silicon adhesive and acrylic adhesive.  Silicon adhesive has been known to migrate over time in cryogenic vacuum environments, which is not known to be a problem with acrylic adhesive.  With no coating the radial and axial positioning accuracy of the lens mount was calculated to be ±40 - 60μm and ±14 - 25μm respectfully and ±2 - 9μm and ±0.5 - 3μm with a low coefficient of friction coating.  Camera #1 lens cell was tested for its positional accuracy by pushing the lens a distance of 150μm with a delrin tipped set screw and then releasing the set screw to see if the lens returned to its original position.  A microscope with a 50x magnification and a fiducial scale with a resolution of 25μm/1div was used to measure the displacements.  There was no noticeable change in the position of the lens after perturbing the lens, which indicated the lens cell was functioning as designed.

# Figure 4:  Roll Pin Flexure Mount for BaF2 Lens Elements at Cryogenic Temperatures.  Mount is composed of two elements: the roll pin flexure cell that constrains the radial motion and the leaf spring disk that constrains the axial motion.

# Figure 5:  Kapton Tape applied to Camera#2 Lens Cell as a Low Coefficient of Friction Coating.

# 1.1.6  Metrology Report

# 1.1.6.1  Profilometer Measurements

# 1.1.6.2  Faro Arm (CMM) Measurements

# 1.2

# Precision Projector Lab Overview

# 1.3  Data Reduction

# 1.4  Image Quality

# 1.4.1  Extra-Focal Imaging

# 1.4.2  Strehl Map

# 1.4.3  MTF Curves

# 1.5  Expected Instrument Throughput

# 1.6  Compliance Matrix

# 1.7  Outstanding Issues

# 1.7.1  Detector Offset from Camera FOV

# 1.7.2  Stray Light

# 1.7.3  Astigmatic Corrector

# References

Smee, S. A., Barkhouser, R. H., Scharfstein, G. A., Meixner, M., Orndorff, J. D. and Miller, T. 2011, ‘Design of the WIYN High Resolution Infrared Camera (WHIRC)’, PASP 123, 87–106