Draft for

Instrument Baseline Requirements Document

Keck Adaptive Optics Note 572

Version 0.1

January 15, 2007

1	Introduction	1

2	Scope and Applicability	1

3	References	2

3.1	Related Documents	2

3.2	Referenced Standards	3

3.2.1	Industry Consensus Standards	3

3.2.2	WMKO Standards	5

3.3	Referenced Drawings	6

4	Revision History	6

5	Background	7

5.1	Purpose	7

6	Overall Requirements	7

6.1	Purpose and Objectives	7

6.2	Performance Requirements	7

6.2.1	Parametric Performance Requirements	7

6.2.1.1.1	Transportation and Shipping Environment	7

6.2.1.1.2	Non-Operating Environment	8

6.2.1.1.3	Operating Environment	9

6.2.2	Operational Performance Requirements	9

6.2.2.1	Air Borne Contaminants	9

6.2.2.2	Audible Noise	9

6.2.2.3	Telescope Reconfiguration	9

6.2.2.4	Power Failure Tolerance	10

6.3	Implementation Requirements	10

6.3.1	Common Practice Implementation Requirements	11

6.3.2	Standards Implementation Requirements	11

6.3.2.1	Shipping Containers	11

6.3.3	Regulatory Implementation Requirements	11

6.4	Design Requirements	14

6.4.1	Technological Design Requirements	14

6.4.1.1	Materials Suitability and Safety	14

6.4.2	Regulatory Design Requirements	15

6.4.3	Standards Related Design Requirements	15

6.4.4	Integration Related Design Requirements	15

7	Optical Requirements	16

7.1	Purpose and Objectives	16

7.2	Performance Requirements	17

7.3	Implementation Requirements	17

7.3.1	Common Practice Implementation Requirements	17

7.3.2	Standards Implementation Requirements	17

7.3.3	Regulatory Implementation Requirements	17

7.4	Design Requirements	17

7.4.1	Technological Design Requirements	17

7.4.1.1	Optical Component Mountings	17

7.4.1.2	Alignment Tolerancing	17

7.4.2	Regulatory Design Requirements	17

7.4.3	Standards Related Design Requirements	18

7.4.4	Integration Related Design Requirements	18

7.4.4.1	Focal Position	18

8	Mechanical Requirements	19

8.1	Purpose and Objectives	19

8.2	Performance Requirements	20

8.2.1	Parametric Performance Requirements	20

8.2.1.1	General	20

8.2.1.2	Vacuum Pump	20

8.2.1.3	Power Dissipation	20

8.2.2	Operational Performance Requirements	20

8.2.2.1	Operating Temperature Range	20

8.2.2.2	Vibration	20

8.3	Implementation Requirements	21

8.3.1	Feature Implementation Requirements	21

8.3.1.1	Instrument Dewar Mechanisms	21

8.3.1.2	Handler	21

8.3.1.3	Rotator	22

8.3.1.4	Tractor Drive	22

8.3.1.5	Access and Covers	22

8.3.1.6	Entrance Window	23

8.3.1.7	Glycol Cooling	24

8.3.1.8	Vacuum Systems	24

8.3.1.8.1	Pressure Control	24

8.3.1.8.2	Gettering	25

8.3.1.9	Cryogenic Systems	25

8.3.2	Common Practices Implementation Requirements	25

8.3.2.1	Fit and Finish	25

8.3.2.2	Continuity of Shielding and Grounding	26

8.3.2.3	Corrosion resistance	26

8.3.2.4	Fasteners	26

8.3.2.5	Lubricants	27

8.3.2.6	Lubricated Components	27

8.3.3	Standards Implementation Requirements	27

8.3.3.1	Structural	27

8.3.3.2	Vacuum Systems	27

8.3.3.3	Cryogenic Systems	27

8.3.4	Regulatory Implementation Requirements	27

8.4	Design Requirements	28

8.4.1	Technological Design Requirements	28

8.4.1.1	Vacuum and Cryogenic Components	28

8.4.1.2	Opto-Mechanical Assemblies	28

8.4.1.3	Electrical/Electronic Assemblies and Enclosures	28

8.4.1.4	Mechanisms	29

8.4.1.5	Drive Couplings	30

8.4.1.6	Component Ratings	30

8.4.2	Regulatory Design Requirements	30

8.4.3	Standards Related Design Requirements	30

8.4.4	Integration Related Design Requirements	30

8.4.4.1	Mounting Position	30

8.4.4.2	Handling	30

9	Electronic/Electrical Requirements	32

9.1	Purpose and Objectives	32

9.2	Performance Requirements	32

9.2.1	Parametric Performance Requirements	32

9.2.1.1	Power Dissipation	32

9.2.1.2	Compatibility	32

9.2.1.3	Temperature and Humidity	32

9.2.1.4	Cable and Wire Ratings	32

9.2.2	Operational Performance Requirements	32

9.3	Implementation Requirements	33

9.3.1	Feature Implementation Requirements	33

9.3.1.1	Emergency Stop Input	33

9.3.1.2	Rotator	33

9.3.1.3	Host Computer	33

9.3.1.4	Data Storage	33

9.3.1.5	Target and “Embedded” Computers	34

9.3.1.6	Instrument Connection Panel	34

9.3.1.7	Printed Circuit Boards	34

9.3.2	Common Practices Implementation Requirements	34

9.3.2.1	Stray Light	34

9.3.2.2	Digital Control and Status Communications	35

9.3.3	Standards Implementation Requirements	35

9.3.3.1	Electrical Safety	35

9.3.3.2	Electromagnetic Compatibility	36

9.3.4	Regulatory Implementation Requirements	36

9.3.4.1	AC Line Connections	36

9.3.4.2	Covers	37

9.3.4.3	Wiring	37

9.3.4.4	Overcurrent Protection	37

9.3.4.5	Grounding and Shielding	38

9.3.4.6	Terminations	38

9.3.4.7	Altitude Derating	38

9.4	Design Requirements	39

9.4.1	Technological Design Requirements	39

9.4.1.1	Motion Control Systems	39

9.4.1.2	Power Ratings	39

9.4.1.3	Wiring and Interconnections	39

9.4.1.3.1	Connector and Cable Mounting	39

9.4.1.3.2	Cable and Wire Routing	39

9.4.1.3.3	Labeling of Interconnections	39

9.4.1.3.4	Interconnections	40

9.4.1.3.5	Data communications – connectors & formats	40

9.4.2	Regulatory Design Requirements	41

9.4.3	Standards Related Design Requirements	41

9.4.4	Integration Related Design Requirements	41

9.4.4.1	Rotator	41

10	Safety Requirements	42

10.1	Purpose and Objectives	42

10.2	Scope	42

10.3	Performance Requirements	42

10.3.1	Parametric Performance Requirements	42

10.3.2	Operational Performance Requirements	42

10.4	Implementation Requirements	42

10.4.1	Feature Implementation Requirements	42

10.4.1.1	Local Control	42

10.4.1.2	Mechanical	43

10.4.1.3	Entrance Window Cover	43

10.4.1.4	Electrical	43

10.4.2	Common Practice Implementation Requirements	43

10.4.3	Standards Implementation Requirements	43

10.4.4	Regulatory Implementation Requirements	44

10.5	Design Requirements	44

10.5.1	Technological Design Requirements	44

10.5.1.1	Instrument	44

10.5.1.2	Rotator	44

10.5.2	Regulatory Design Requirements	45

10.5.3	Standards Related Design Requirements	45

10.5.4	Integration Related Design Requirements	45

11	Software Requirements	46

11.1	Purpose and Objectives	46

11.2	Scope	46

11.3	Performance Requirements	46

11.3.1	Parametric Performance Requirements	46

11.3.1.1	Reliability	46

11.3.1.2	Fiber Optic Data Links	47

11.3.1.3	Data Transfer Performance	47

11.3.1.4	Display Updates	47

11.3.2	Operational Performance Requirements	47

11.3.2.1	Overhead	47

11.3.2.2	Error Recovery	47

11.3.2.2.1	Loss of Network Connections	47

11.3.2.2.2	Detector Controller Aborts	48

11.3.2.2.3	Data Disk Full	48

11.3.2.3	Execution Speed and Command Latency	48

11.4	Implementation Requirements	49

11.4.1	Feature Implementation Requirements	49

11.4.1.1	User interfaces	49

11.4.1.2	Image Display	49

11.4.1.3	CSU Configuration	49

11.4.1.4	Data Reduction Pipeline	49

11.4.1.5	MOSFIRE Target Computers	50

11.4.1.6	Software Licenses	50

11.4.2	Common Practice Implementation Requirements	50

11.4.3	Standards Implementation Requirements	50

11.4.4	Regulatory Implementation Requirements	50

11.5	Design Requirements	51

11.5.1	Technological Design Requirements	51

11.5.1.1	Client-Server Architecture	51

11.5.1.2	Communications Protocols	51

11.5.1.3	Keywords	51

11.5.1.4	Target Software	51

11.5.1.5	Host Software	52

11.5.1.6	Science Data File Formats	52

11.5.2	Regulatory Design Requirements	52

11.5.3	Standards Related Design Requirements	52

11.5.4	Integration Related Design Requirements	52

12	Interface Requirements	53

12.1	Purpose and Objectives	53

12.2	Performance Requirements	53

12.2.1	Parametric Performance Requirements	53

12.2.2	Operational Performance Requirements	53

12.2.2.1	Handling	53

12.3	Implementation Requirements	53

12.3.1	Feature Implementation Requirements	53

12.3.1.1	Optical Requirements	53

12.3.1.2	Mechanical	53

12.3.2	Common Practice Implementation Requirements	53

12.3.2.1	Glycol Cooling	53

12.3.2.2	Vacuum and Cryogenics	53

12.3.2.3	Stray Light	53

12.3.3	Standards Implementation Requirements	54

12.3.4	Regulatory Implementation Requirements	54

12.4	Design Requirements	54

12.4.1	Technological Design Requirements	54

12.4.2	Regulatory Design Requirements	54

12.4.3	Standards Related Design Requirements	54

12.4.4	Integration Related Design Requirements	54

12.4.4.1	Optical Interface	54

12.4.4.2	Mechanical Interface	54

12.4.4.3	Electrical/Electronic Interface	54

13	Reliability Requirements	55

13.1	Purpose	55

13.2	Scope	55

13.3	Procedure for Reliability Determination	55

14	Spares Requirements	55

15	Service and Maintenance Requirements	56

16	Documentation Requirements	57

16.1	Documentation Package	57

16.2	Drawings	58

16.2.1	Drawing Standards	58

16.2.2	Required Drawings	59

16.3	Electrical/Electronic Documentation	60

16.4	Software	60

17	Glossary	63

Figure 3: Keck Telescope Equipment Vibration Limits	21

Table 1: Referenced Standards	3

Table 2: WMKO Standards	5

Table 4: Transportation and Shipping Environment	7

Table 5: Non-Operating Environment	8

Table 6: Operating Environment	9

Table 7: Materials not Suitable for use in Equipment at the Summit of Mauna Kea	14

Table 17: Software Latencies	48

Table 19: Glossary of Terms	63

# Introduction

This document describes the baseline requirements for W. M. Keck Observatory (WMKO) instruments.

The requirements in this document are at a draft level appropriate for the preliminary design phase of the instrument.  Further development of the requirements will take place in the next phase of the project (detailed design).  In particular, parametric performance requirements given at this stage are intended to indicate the scope and format of the requirements, but do not in all cases establish final values for the specified parameters.  In some cases values for these parameters have yet to be established and are given as TBD.

It is important to understand that at this stage of development the requirements provide a basis for identifying the parameters that will be part of the instrument’s specifications, but the values given are subject to change as the development of the instrument continues.  During the next phase of the project work will be done to refine the instrument’s specifications into final specifications that will be reviewed at the detailed design review.  The final specifications will also form the basis for the acceptance test criteria for the instrument.

The purpose of a requirements document is to define and communicate the Observatory’s expectations for the design and implementation of a new scientific instrument for the Observatory.  As the procuring organization, WMKO authors the requirements document in collaboration with the instrument design team.

A requirements document describes the new instrument in terms of the needed scientific and technical performance.  The document also expresses specific requirements for implementation or design where those requirements are essential to satisfactory integration and interoperation of the instrument with the observatory systems.  The requirements document also references consensus standards approved by recognized standards organizations for specific guidance on technical matters related to implementation, compatibility and safety.

The document avoids prescribing specific design or implementation solutions except for solutions that embody the Observatory’s unique knowledge or experience.  The document establishes requirements for the new instrument that will guide the design of the instrument through the detailed design phase.

# Scope and Applicability

This document establishes requirements for all future WMKO instruments.  This document also establishes requirements for changes to related sub-systems and software of the Keck telescopes where required.

This revision of the document is the first release.

# References

## Related Documents

## Referenced Standards

### Industry Consensus Standards

Table 1 lists the industry consensus standards referenced in this document in alphabetical order by standardizing organization.  Unless otherwise noted all references to standards are included because compliance with some part of each standard may be required.

Table 1: Referenced Standards

| Source (Organization or Standardizing Body) | Number | Title |
| --- | --- | --- |
| ANSI | Y14.5M-1994 (R1999) | Dimensioning and Tolerancing |
| ANSI | Y14.1-1995 (R2002) | Decimal Inch Drawing Sheet Size And Format |
| ANSI | Y14.34-2003 | Parts Lists, Data Lists, And Index Lists: Associated Lists |
| ANSI | Y14.3M-1994 | Multi And Sectional View Drawings |
| ANSI / ASME | Y14.18M-1986 | Optical Parts (Engineering Drawings and Related Documentation Practices) |
| ASME | HPS-2003 | High Pressure Systems |
| ASME | Y14.100-2000 | Engineering Drawing Practices |
| ASME | Y32.10-1967 (R1994) | Graphic Symbols for Fluid Power Diagrams |
| ASTM | E595-93 (2003)e1 | Standard Test Method for Total Mass Loss and Collected Volatile Condensable Materials from Outgassing in a Vacuum Environment |
| ATA | Spec 300-2001.1 | Specification for Packaging of Airline Supplies |
| CENELEC | EN 50082-1:19971 | Electromagnetic compatibility – Generic immunity standard – Part 1: Residential, commercial and light industry |
| Council of the European Communities | EMC 89/336/EEC1 | Council Directive 89/336/EEC of 3 May 1989 on the approximation of the laws of the Member States relating to electromagnetic compatibility (EMC Directive) |
| County of Hawaii | 1995 edition | Hawaii County Code 1983 (1995 edition) |
| Department of Defense | MIL- STD-171E | Finishing of Metal and Wood Surfaces |
| Department of Defense | MIL-HDBK-217F-21 | Reliability Prediction of Electronic Equipment |

1. This reference for information only.

Table 1: Referenced Standards, continued

| Source (Organization or Standardizing Body) | Number | Title |
| --- | --- | --- |
| Department of Defense | MIL-STD-810F | Test Method Standard for Environmental Engineering Considerations and Laboratory Tests |
| EIA | EIA-310-D | Cabinets, Racks, Panels, and Associated Equipment |
| EIA | EIA-6491 | National Consensus Standard For Configuration Management |
| FCC | Title 47 CFR Part 151 | Radio Frequency Devices |
| IEEE | 802.3U revision 95 | Carrier Sense Multiple Access with Collision Detection (CSMA/CD) Access Method & Physical Layer Specifications: Mac Parameters, Physical Layer, Medium Attachment Units and Repeater for 100 Mb/S Operation (Version 5.0) |
| IEEE | 1012-2004 | Standard for Software Verification and Validation |
| International Code Council (ICC) | IBC-2006 | 2006 International Building Code® |
| ISO/IEC | ISO / IEC 12207:1995 | Information Technology - Software life cycle processes |
| National Electric Manufacturers Association | 250-1997 | Enclosures for Electrical Equipment (1000 Volts Maximum) |
| National Fire Protection Association (NFPA) | NFPA 55, 2005 edition | Standard for the Storage, Use, and Handling of Compressed Gases and Cryogenic Fluids in Portable and Stationary Containers, Cylinders and Tanks |
| NFPA | NFPA 70, 2005 edition | National Electric Code |
| NFPA | NFPA 99C, 2005 edition | Standard on Gas and Vacuum Systems |
| Naval Surface Warfare Center | NSWC 98/LE11 | Handbook of Reliability Prediction Procedures for Mechanical Equipment |
| OSHA | Title 29 CFR Part 1910 | Occupational Safety And Health Standards |
| Telcordia | GR-63-CORE | NEBS™ Requirements |

1. This reference for information only.

Table 1: Referenced Standards, continued

| Source (Organization or Standardizing Body) | Number | Title |
| --- | --- | --- |
| TIA/EIA | TIA/EIA-568-B | Commercial Building Telecommunications Cabling Standards |
| Underwriters Laboratories Inc. | Standard for Safety 508 | Industrial Control Equipment |

1. This reference for information only.

### WMKO Standards

WMKO software standards are also referenced in this document.  References to these standards are included because compliance with some part of each standard may be required.

Table 2: WMKO Standards

| Source (Organization or Standardizing Body) | Number | Title |
| --- | --- | --- |
| WMKO | KSD 3 | Software Items for Acceptance Review |
| WMKO | KSD 8 | KTL: the Keck Task Library |
| WMKO | KSD 46a | DCS Keyword Reference Manual (partial) |
| WMKO | KSD 50 | Keck II C Style and Coding Standards |
| WMKO | KSD 201 | How to Set Up a config.mk Build |
| WMKO | KSD 210 | WMKO Software Standards |

## Referenced Drawings

# Revision History

| Version | Date | Author | Reason for revision / remarks |
| --- | --- | --- | --- |
| 0.0 | Apr. 4, 2006 | SMA | MOSFIRE Requirements Document |
| 0.1 | Jan. 15, 2007 | PW | Rework as a more generic document |

Due to the difficulties in documents with moderately complex formatting such as this one, the Microsoft Word “Track Changes” feature is not useable.  To see the changes in this document since the previous version, use the “Tools, Track Changes, Compare Documents” drop down menu sequence and compare this document to the previous version.  It is not recommended that you attempt to print the results.  Subsequent versions of this document will include the filename and date for the previous version.

# Background

## Purpose

The purpose of the background section of this document is to provide context and related information for the requirements defined in later sections of this document.

# Overall Requirements

## Purpose and Objectives

The purpose of the overall requirements section is to convey requirements that apply generally to the overall instrument and its accessories.

## Performance Requirements

### Parametric Performance Requirements

##### Transportation and Shipping Environment

The Instrument shall continue to meet all of the performance requirements without repair after a single shipment to the delivery location by any combination of air or surface transportation.  For information, the expected conditions to be encountered during shipping are given in Table 3.

Table 3: Transportation and Shipping Environment

| Parameter | Min. | Typ. | Max. | Units | Notes |
| --- | --- | --- | --- | --- | --- |
| Altitude | 0 | - | 4,572 | m | 1 |
| Temperature | -33 | - | 71 | ºC | 2, 3 |
| Temperature shock | -54 | - | 70 | ºC | 4 |
| Humidity | 0 | - | 100 | % | 5 |
| Gravity orientation | - | - | - | NA | 6 |
| Vibration | - | - | 0.015 | g2/Hz | 7, 8 |
| Shock | - | - | 15 | g | 9 |
| Acceleration |  |  |  |  |  |
| Due to transport | - | - | 4 | g | 10 |
| Due to seismic activity | - | - | 2 | g | 12 |

Notes:

    - See MIL-STD-810F Method 500 §2.3.1.

    - Maximum is for induced conditions, see MIL-STD-810F Method 501 Table 501.4-I.

    - Minimum is for induced conditions, see MIL-STD-810F Method 502 Table 502.4-II.

    - See MIL-STD-810F Method 503.

    - Relative, condensing.

    - Packaged equipment may be subjected to all possible gravity orientations during transportation and shipping.

    - 10 Hz to 40 Hz, -6dB/oct. drop-off to 500 Hz, all axes.

    - See MIL-STD-810F Method 514.

    - 0.015 second half-sine, all axes.

    - All axes.

    - 0.5 Hz to 100Hz, all axes.

##### Non-Operating Environment

The Instrument shall meet all of the performance specifications without repair or realignment after being subjected to any number of cycles of any of the non-operating environment conditions defined in Table 4.  These represent environments associated with normal non-operating telescope activities including but not limited to storage and handling within the facility and installation and removal from the telescope.

Table 4: Non-Operating Environment

| Parameter | Min. | Typ. | Max. | Units | Notes |
| --- | --- | --- | --- | --- | --- |
| Altitude | 0 | - | 4300 | m |  |
| Temperature |  |  |  |  |  |
| Range | -10 | 0 | 30 | ºC | 1 |
| Rate of change | -0.8 | - | 0.8 | ºC/h |  |
| Humidity | 0 | - | 90 | % | 2 |
| Gravity orientation | - | -1 | - | g | 3 |
| Vibration | - | - | 8.0x10-4 | g2/Hz | 4 |
| Shock | - | - | 15 | g | 5 |
| Acceleration |  |  |  |  |  |
| Due to handling | - | - | - | g | 6 |
| Due to seismic activity | - | - | 2 | g | 7 |

Notes:

- Typical value is the average annual temperature.

- Relative, non-condensing.

- Normal to the earth’s surface.

- 20 Hz to 1000 Hz, 6db/oct drop- off to 2000 Hz.

- 0.015 second half-sine, all axes.

- 2 g vertical, 1 g fore/aft, 0.5 g lateral

- 0.5 Hz to 100Hz, all axes.

##### Operating Environment

The operating environment is the ensemble of all conditions experienced under normal telescope operation when the Instrument is installed at the Keck  telescope.  All performance requirements shall be met while the Instrument is subjected to the operating environment conditions given in Table 5.

Table 5: Operating Environment

| Parameter | Min. | Typ. | Max. | Units | Notes |
| --- | --- | --- | --- | --- | --- |
| Altitude | 0 | - | 4300 | m |  |
| Temperature |  |  |  |  |  |
| Range | -10 | 0 | 20 | ºC | 1 |
| Rate of change | -0.8 | - | 0.8 | ºC/h |  |
| Humidity | 0 | - | 90 | % | 2 |
| Gravity orientation | - | -1 | - | g | 3 |
| Vibration | - | - | 1x10-5 | g2/Hz | 4 |
| Acceleration | - | - | 1 | g | 5 |

Notes:

- Typical value is the average annual temperature.

- Relative, non-condensing.

- Normal to the earth’s surface.

- 20 Hz to 1000 Hz, 6db/oct drop- off to 2000 Hz.

- All axes, due to telescope drive system fault conditions.

### Operational Performance Requirements

#### Air Borne Contaminants

The weather conditions at the summit of Mauna Kea include frequent high winds resulting in some air borne contaminants, particularly dust and insects.  Instruments must be protected during installation and handling against the entry of these contaminants, in particular care must be taken with optical surfaces, precision mechanisms and fine pitch or fiber optic connectors.

#### Audible Noise

Unless otherwise specified or accepted any pumps, motors, outboard electronics or computers should not at any time produce audible noise in excess of 50 dBA at a distance of 1 meter.  This is a standard office operating environment maximum noise level.  This includes intermittent noises from pumps and variable speed cooling fans.  Audible warning signals for emergency or fault conditions are exempt from this requirement, but they must be provided with a silence after delay feature or a manual silencing switch.

#### Telescope Reconfiguration

None.

#### Power Failure Tolerance

The observatory summit facilities provide backup power to the instrument electronics.  The first level of backup is the Keck I or II dome UPS, an industrial uninterruptible power supply (UPS) shared with the other instruments.  This UPS has a hold up time of 30 minutes.  A separate UPS is provided for the Keck I or II computer room, and this UPS provides backup power for the instrument computers.  The computer room UPS also has a hold up time of 30 minutes.

Under normal conditions the observatory summit standby generator will start within 1 minute of the power failure and begin supplying primary power to the Keck dome UPS, Keck computer room UPS and the other UPS units at the summit.

During a power failure the glycol cooling system pumps and chiller will be inoperative, so instrument electronics dependent on glycol cooling require either flow switches or temperature sensors to ensure that the electronics are shut down even though the electronics will be powered from the Keck dome UPS and the observatory summit standby generator.

The CCR compressors and CCR heads are powered from the generator but they require glycol cooling for continuous operation.  During a power failure the CCR compressors will experience momentary power interruptions of less than 1 minute duration and will then continue to operate on the generator until their thermal protection systems shut them down.

Under normal conditions the observatory summit standby generator has sufficient fuel for 18 hours of continuous operation at full load.  With only two exceptions in over 10 years of operation, the longest power failures to date that WMKO has experienced at the summit have been less than 1 hour in duration.

The worst-case conditions to be experienced by the instrument can be understood to occur under conditions where the observatory summit standby generator fails to start.  In this case the CCR compressors and CCR heads will cease to operate, and within 30 minutes the dome UPS and computer room UPS will be exhausted resulting in a total instrument power failure for a further 30 minutes based on the majority of the worst case power failures to date.

Because of the possibility of power failures, and also the necessity of disconnecting instruments from services during telescope reconfiguration, instruments should be designed so that power failures of up to 1 hour in duration affecting the electronics, glycol cooling and CCR systems will not result in permanent loss of performance or damage to the instrument’s detectors or other components.

## Implementation Requirements

None.

### Common Practice Implementation Requirements

None.

### Standards Implementation Requirements

#### Shipping Containers

All shipping containers must be designed to provide adequate protection for the equipment during transport.  Unless otherwise specified single use containers suitable for the size, weight and shipment method to be employed are acceptable.  For guidance in the design of suitable containers consult Air Transport Association (ATA) Spec 300, 2001.1 edition, “Specification for Packaging of Airline Supplies”.

### Regulatory Implementation Requirements

The Instrument shall comply in all respects with the applicable requirements of the Occupational Safety and Health Administration (OSHA) as established by Code of Federal Regulations (CFR) 29 Part 1910 “Occupational Safety And Health Standards”, particularly subpart O, section 1910.212 and subpart S sections 1910.302 through 1910.304.

The requirements of Subpart O, section 1910.212 that are applicable to an Instrument are summarized as follows:

- Machine guarding must be provided to protect the operator and other employees from hazards such as those created by ingoing nip points or rotating parts.

- Guards shall be affixed to the machine.

- Revolving barrels and drums shall be guarded by an enclosure that is interlocked with the drive mechanism so that the barrel or drum cannot revolve unless the guard is in place.

The requirements of Subpart S, sections 1910.302 through 1910.304 that are applicable to MOSFIRE may be summarized as follows:

- Listed or labeled equipment shall be used or installed in accordance with any instructions included in the listing or labeling.

- Conductors shall be spliced or joined with splicing devices suitable for the use or by brazing, welding, or soldering with a fusible metal or alloy.  Soldered splices shall first be so spliced or joined as to be mechanically and electrically secure without solder and then soldered.  All splices and joints and the free ends of conductors shall be covered with insulation equivalent to that of the conductors or with an insulating device suitable for the purpose.

- Parts of electric equipment which in ordinary operation produce arcs, sparks, flames, or molten metal shall be enclosed or separated and isolated from all combustible material.

- Electrical equipment may not be used unless the manufacturer’s name, trademark, or other descriptive marking by which the organization responsible for the product may be identified is placed on the equipment.  Other markings shall be provided giving voltage, current, wattage, or other ratings as necessary.  The marking shall be of sufficient durability to withstand the environment involved.

- Each disconnecting means for motors and appliances shall be legibly marked to indicate its purpose, unless located and arranged so the purpose is evident.

- Live parts of electric equipment operating at 50 volts or more shall be guarded against accidental contact by approved cabinets or other forms of approved enclosures.

- A conductor used as a grounded conductor shall be identifiable and distinguishable from all other conductors.  A conductor used as an equipment grounding conductor shall be identifiable and distinguishable from all other conductors.

- No grounded conductor may be attached to any terminal or lead so as to reverse designated polarity.

- A grounding terminal or grounding-type device on a receptacle, cord connector, or attachment plug may not be used for purposes other than grounding.

- Conductors and equipment shall be protected from overcurrent in accordance with their ability to safely conduct current.

- Overcurrent devices may not interrupt the continuity of the grounded conductor unless all conductors of the circuit are opened simultaneously.

- Overcurrent devices shall be readily accessible to each employee or authorized building management personnel.  These overcurrent devices may not be located where they will be exposed neither to physical damage nor in the vicinity of easily ignitable material.

- Fuses and circuit breakers shall be so located or shielded that employees will not be burned or otherwise injured by their operation due to arcing or suddenly moving parts.

- Circuit breakers shall clearly indicate whether they are in the open (off) or closed (on) position.

- The path to ground from circuits, equipment, and enclosures shall be permanent and continuous.

- Metal enclosures for conductors shall be grounded.

- Exposed, non-current-carrying metal parts of fixed equipment, which may become energized, shall be grounded.

- Exposed non-current-carrying metal parts of cord and plug connected equipment, which may become energized, shall be grounded.

- Non-current-carrying metal parts of fixed equipment, if required to be grounded, shall be grounded by an equipment grounding conductor, which is contained within the same raceway, cable, or cord, or runs with or encloses the circuit conductors.  For DC circuits only, the equipment grounding conductor may be run separately from the circuit conductors.

For the purposes of the foregoing approved means acceptable to the authority enforcing the applicable subpart.  The authority enforcing the applicable subpart is the Assistant Secretary of Labor for Occupational Safety and Health.  The definition of ‘‘acceptable’’ indicates what is acceptable to the Assistant Secretary of Labor, and therefore approved within the meaning of the applicable subpart.  Approved for the purpose means approved a specific purpose, environment, or application described in a particular standard requirement.  Suitability of equipment or materials for a specific purpose, environment or application may be determined by a nationally recognized testing laboratory, inspection agency or other organization concerned with product evaluation as part of its listing and labeling program.

Note that the preceding text is reproduced verbatim from the referenced CFR and any grammatical errors or typographical errors are part of that text.

## Design Requirements

### Technological Design Requirements

#### Materials Suitability and Safety

Certain environmental conditions (low temperature and pressure) at the summit of Mauna Kea make certain materials unsuitable for use in instrument construction.  Materials used in the construction, lubrication or packaging of instruments must not produce hazardous by-products such as gases or other contaminants under the conditions of operation and use at the summit of Mauna Kea.  No mercury may be used in any component.

Table 6 lists specific materials that should not be used.  Note that this table applies to portions of the instrument normally open to the atmosphere.  See §8.4.1.1 for materials considerations for vacuum cryostats and similar environments.

Table 6: Materials not Suitable for use in Equipment at the Summit of Mauna Kea

| Material Type | Common Name | Reason(s) for Unsuitability |
| --- | --- | --- |
| Adhesive, insulator | RTV silicone rubber1 | Outgases during curing |
| Adhesive | Cyanoacrylates | Outgases during curing, subject to hydrolytic degradation |
| Conductor | Mercury2 | Reactive, salts formed are toxic |
| Insulator | Acrylic4 | Outgases, hygroscopic, brittle at low temperatures |
| Plated finish | Cadmium2 | Outgases, reactive, hazardous |
| Insulator | Cellulose Acetate Butyrate | Hygroscopic |
| Insulator | Glass-Reinforced Extruded Nylon | Outgases, hygroscopic |
| Insulator | Kapton | Subject to hydrolytic degradation |
| Insulator | Neoprene | Outgases, subject to degradation by ozone and UV exposure |
| Insulator | Nylon5 | Outgases, subject to degradation by ozone and UV exposure |
| Insulator | Phenolic3 | Hygroscopic |
| Insulator | Polychlorinated Biphenyls2 | Combustion produces highly toxic gases |

Notes:

- Neutral cure RTV silicones may be acceptable provided that the cured silicone and the surrounding area are cleaned after assembly.

- Use is or soon will be highly regulated.

- Electrical grade phenolic is not hygroscopic.

- Cast acrylic resin

- Cable ties of weather resistant Nylon 6/6 (carbon black additive) are acceptable.

### Regulatory Design Requirements

None.

### Standards Related Design Requirements

None.

### Integration Related Design Requirements

None.

# Optical Requirements

## Purpose and Objectives

The purpose of this section is to describe requirements for the performance, implementation and design of the optical system.

## Performance Requirements

## Implementation Requirements

### Common Practice Implementation Requirements

None.

### Standards Implementation Requirements

None.

### Regulatory Implementation Requirements

None.

## Design Requirements

### Technological Design Requirements

#### Optical Component Mountings

All optical components should be mounted so that alignment is maintained during cool down and warm up cycles, where appropriate.  Mountings must ensure that excessive stress is not placed on the optical components due to thermal differentials between the optical component and the mount.  Mountings must also ensure that alignment of optical components without excessive stress is maintained at all rotator angles and telescope elevations.

Materials used in optical component mountings, particularly elastomers and adhesives must be compatible with the coatings on the associated optical components.  All materials used within the dewar must be compatible with vacuum and cryogenic environments, see §8.4.1.1.

#### Alignment Tolerancing

Before assembly all optical components and systems must have a documented optical alignment tolerance budget.  During assembly measurements must be made as required to ensure that the stack-up of tolerances does not exceed the tolerance budget.

### Regulatory Design Requirements

None

### Standards Related Design Requirements

Drawings for optical components should conform to American National Standards Institute (ANSI) / American Society of Mechanical Engineers International (ASME) standard Y14.18M-1986 “Optical Parts (Engineering Drawings and Related Documentation Practices)”.

### Integration Related Design Requirements

#### Focal Position

# Mechanical Requirements

## Purpose and Objectives

The purpose of this section is to describe requirements for the performance, implementation and design of the Instrument mechanical systems.  In many cases these requirements reflect the preliminary mechanical design of the instrument.

The mechanical requirements address issues of design, reliability and maintainability.  Based on experience with previous instruments the observatory is sensitive to certain aspects of performance, implementation and design that have proven to be important factors in the up time of its instruments.  The mechanical requirements section has as a main objective the description of the expected performance, features and configuration of the instrument’s mechanical systems.  A secondary objective is to identify specific areas where experience indicates particular attention is required with respect to performance, implementation or design.

In this revision of the document some of the mechanical requirements are very detailed, and others are broader.  These broader requirements may need to be broken down into more detailed requirements in a further revision of this document.  In the case of parametric requirements, many of the values given are starting points for a more detailed analysis that will take place in the next phase of the instrument’s design.

## Performance Requirements

### Parametric Performance Requirements

#### General

#### Vacuum Pump

If required to protect the instrument in the event of an unintended warm up, an on instrument vacuum pump may be provided.  This pump must be provided with remote control facilities, and must tolerate all orientations when not operating without impairment of performance or leakage of oils or other fluids.

#### Power Dissipation

The Instrument must not radiate more than 50 watts of heat into the telescope dome ambient environment.  All heat generated by the Instrument in excess of this amount must be carried away by a glycol based cooling system.

This requirement does not apply to equipment not used during normal operations.

### Operational Performance Requirements

#### Operating Temperature Range

The Instrument should be designed for operation over the ambient temperature range given in §6.2.1.1.3.

#### Vibration

Vibration isolation should be employed as required to isolate sources of vibration within the Instrument due to moving components such as fans, pumps and motors.

The Instrument should meet all performance and operating requirements when installed in a vibration environment that conforms to the Generic Vibration Criteria Curve “C” as shown in Figure 1.  The Instrument should not produce vibrations that result in rms velocities in excess of those given in curve “C” of Figure 1.

Figure 1: Keck Telescope Equipment Vibration Limits

## Implementation Requirements

### Feature Implementation Requirements

#### Instrument Dewar Mechanisms

Mechanisms internal to the Instrument dewar that are difficult to access for service, features should be provided that maximize the reliability of the mechanisms and provide as much information as possible about the status and performance of each mechanism.

All Instrument dewar mechanisms should provide a positive indication that the requested move(s) have been completed.  The use of a relative position indicating means in conjunction with limit switches is preferred.

Mechanisms should operate properly with reduced speed over the ambient temperature range given in §6.2.1.1.3.  This is essential to permit servicing and verification of proper operation prior to evacuation and cooling of the instrument dewar.

#### Handler

A handler must incorporate structural components that will maintain its integrity and ensure secure mounting during an earthquake with the Instrument installed as required by seismic standards for a zone 4 earthquake zone (see §8.3.3.1 below).

The handler must incorporate seismic restraint provisions for use when the handler is parked at the storage position.

The handler should be equipped with a removable tractor drive assembly compatible with the existing tractor drive assembly and drive method.

#### Rotator

A rotator required to rotate the Instrument about the telescope’s optical axis in order to compensate for the image rotation that occurs as the telescope follows the sidereal motion of the sky.

The rotator should incorporate a mechanical lockout feature that locks the Instrument in place so that it cannot rotate.  This feature will ensure that the instrument will not move due to an imbalance caused by removal of a component for service.  Mechanical lockout features should activate an electrical lockout consisting of one or more non-defeatable switches that disable the drive system when the mechanical lockout is active and provide a remote indication that the mechanical lockout is active.  The electrical lockout will protect the rotator drive system components as well as prevent unintended drive activation.

Where appropriate, the rotator must incorporate a defining point system that is compatible with the existing Keck defining point system.  In particular the components of the defining system on the rotator must incorporate sufficient adjustment range to allow alignment of the Instrument.

#### Tractor Drive

Where appropriate, a removable tractor drive assembly compatible with the existing tractor drive assembly and drive method should be provided to move the handler on the Keck   platform and Nasmyth deck rail system.

#### Access and Covers

Components requiring routine service or maintenance should be accessible by removing a single cover secured by no more than 8 fasteners.  Covers that may be removed in a location where fasteners could fall into the interior of the enclosure or the instrument may be equipped with captive fasteners.  Covers that may be removed in a location where fasteners could fall into the interior of the enclosure should be equipped with captive fasteners.  Captive fasteners shall be of the threaded type and shall not captivated by swaged sleeve fittings.  Quarter turn fasteners engaging spring hooks are specifically discouraged for reasons of fit and reliability.

Whenever possible service access provisions should be provided that do not require disassembly of the entire instrument to access motors or switches for replacement.

All electronics systems of the Instrument (not including the science detector and ASIC inside the instrument dewar) must be accessible for service without returning the instrument to atmospheric pressure.

#### Entrance Window

Where appropriate, a remotely operated cover should be provided that protects the Instrument optics from dust and from damage due to glancing or direct blows or impacts while in the storage position or moving from storage to the telescope.  A typical scenario for the calculation of forces involved is as follows:

A person moving at a normal walking pace (~1.3 m/s) carrying a 3 meter length of schedule 80 1-1/4” pipe (~14 kg) walks directly towards the front of MOSFIRE.  The pipe strikes the cover.  The person carrying the pipe does not loose his grip on the pipe and for the purposes of this analysis the v in the collision is –1.3 m/s.

The cover should be able to resist the resulting force without damage to the Instrument.

The cover should be interlocked to the instrument and telescope interlocks so that the cover is prevented from opening except when the Instrument is defined at the operating position.  Special provisions for local operation of the entrance window cover while the instrument is in other positions may be required, but remote operation of the entrance window cover should only be possible when the instrument is defined at the Keck I Cassegrain position.

#### Glycol Cooling

When glycol cooling is required the following implementation requirements apply.

All glycol cooling should be plumbed with braided stainless steel hose and stainless steel fittings.  Custom manifolds should be used rather than arrangements of “T” fittings and hose.  Permanent connections should be made with JIC 37º flare compression fittings or SAE straight thread O-ring fittings.  Teflon tape should not be used to seal threaded connections.

Removable connections should be made with ½” Parker Hannifin series FS quick disconnect fittings.  The instrument supply coupler is male and the return coupler is female.

Where required King Instrument Company flow meters and needle valves are preferred for flow metering and control applications.  Where variable gravity orientations are encountered a spring loaded variable area flow meter, such as the in-line flow meters manufactured by the Hedland Division of Racine Federated Inc. should be employed.  The Hedland T303 stainless steel models are preferred.

All glycol cooling systems should be provided with a flow switch, Proteus Industries Inc. type 100B110 is preferred, to generate a loss of coolant alarm.  This flow switch should interrupt power to the affected system unless a separate over-temperature detection system is provided to remove power from the affected system.

#### Vacuum Systems

When vacuum systems are required the following requirements apply.

Vacuum system implementations must prevent contamination of the dewar from back streaming of oil or other contaminants.  Oil free pumps are preferred.

NW 40 size KF flanges are preferred.  All vacuum system fittings, including valves and piping and flexible couplings should be stainless steel.

##### Pressure Control

Vacuum systems must be equipped with at vacuum gauge facilities capable of accurately measuring the pressure in the dewar.  This should consist of at least one low vacuum gauge and one high vacuum gauge.  A back-up high vacuum gauge is also desired.

Dewars must be equipped with pressure relief valves to protect against over pressure due to the liberation of adsorbed gasses during the warm up process.

##### Gettering

Vacuum systems must be equipped with passive gettering for the reduction of water and gasses adsorbed by the dewar walls and internal components.

Where molecular sieves such as Zeolite are used to perform gettering the sieves must be able to be removed and replaced without returning the instrument dewar to atmospheric pressure.  Regeneration of the sieves after a warm up must be accomplished in a manner that removes all adsorbed water from the sieve without contaminating the dewar or other components with water.  The grain size of molecular sieve material should be selected to minimize the potential for migration of sieve material from the sieve holder.  Electropolished stainless steel mesh should be used for the sieve holder.  All components of the sieve holder must withstand baking at temperatures up to 350 ˚C without damage, outgassing (except for adsorbed water) or deterioration.

#### Cryogenic Systems

Where auto-fill systems are employed for LN2 cryogen they should be compatible with the auto-fill systems currently in use at the observatory.  In the event of auto-fill failure, manual fill must be possible.

Cryogenic systems should provide adequate cryopumping capability to completely condense all residual gasses remaining at initial cool down.

CCR heads should be vibration isolated from the instrument dewar.

Manifolds should be provided for the distribution of helium to the CCR heads according to the capacity of the associated compressors in order to minimize the number of instrument interconnections required.

### Common Practices Implementation Requirements

#### Fit and Finish

All steel or iron components should be plated or painted to prevent rust.  This includes fasteners and rivets.  Welds not ground to the surface or joint profile should be of dress quality.  All welds and castings must be stress relieved prior to painting and assembly.

Machined components should be free of tool marks, scratches and material flaws such as inclusions or voids.

Unless otherwise specified all external enclosure and exposed structural elements should be finished in TBD epoxy paint applied in accord with the manufacturer’s instructions.

All burrs and sharp edges shall be removed from all fabricated components unless the function of the component requires a sharp edge.

Mild steel surfaces that cannot be painted for functional reasons (such as accurate interface surfaces) shall be protected by a non-tracking anti-corrosion dry film lubricant.

#### Continuity of Shielding and Grounding

Dissimilar metals in contact under conditions where electrolytic corrosion may occur will be isolated by a dielectric finish or insulating spacers.  Not withstanding this requirement all components of enclosures that are required to provide protective grounding or EMI shielding must be electrically bonded at multiple points by threaded fasteners, finger stock, or a continuous conductive elastomeric gasket.  If grounding straps are used they must be tin plated copper braids not less than 6 mm in width.  Anodized aluminum parts must be free of anodizing at the points where electrical contact is required.  Painted metal parts must be free of paint at the points where electrical contact is required.

#### Corrosion resistance

All metal components should be finished to prevent corrosion in the operating environment (see Table 5) over a normal 10 year lifetime of operation including handling, maintenance and repair.

All removable fasteners must be plated or treated to prevent corrosion.

Internal components may be plated or paint finished.  A contractor who can show conformance to the requirements of MIL-STD-171E “Finishing of Metal and Wood Surfaces” or equivalent should perform any required painting, plating or anodizing.

#### Fasteners

Press fit studs or threaded inserts must be installed in the correct material (i.e. no aluminum inserts in steel) according to the manufacturer’s instructions.  Samples of such fasteners installed in the actual material should be obtained and subjected to pull out tests prior to use in an actual design.  Self tapping screws should not be used for removable covers or to secure components that will have to be removed for repair or replacement.

Fasteners should have either Phillips or hex socket heads.  Hex socket button head fasteners should not be used except where space or specific function requires them.  Undercut machine screws should not be used except in special cases where there is no other appropriate design alternative.

Prevailing torque locknuts or lock washers are preferred to thread locking compounds.  Soft insert locknuts should have Kel-F or Vespel inserts, and should only be used where subsequent removal is not anticipated.

#### Lubricants

Lubricants must be suited for the low temperature environment encountered at the summit.  The base oil in a grease lubricant should have a high viscosity index, a low pour point temperature and a low viscosity at the average operating temperature (based on a 0 ˚C ambient).  Greases using synthetic base oils such as Fluoroether or Silicone are preferred.

#### Lubricated Components

Exposed lubricated components such as gear trains or lead screws should be enclosed in a shroud or boot to prevent the collection of dust and dirt and also to prevent accidental contact that may result in the transfer of the lubricant to other surfaces.

### Standards Implementation Requirements

#### Structural

The structure of MOSFIRE should meet the zone 4 earthquake survival requirements of Telcordia Standard GR-63-CORE, “NEBS™ Requirements”.

#### Vacuum Systems

Vacuum systems should be implemented in conformance with the requirements of ASME HPS-2003, “High Pressure Systems” and NFPA 99C, “Standard on Gas and Vacuum Systems”, 2005 edition.

#### Cryogenic Systems

Cryogenic systems should be implemented in conformance with the requirements of NFPA 55, “Standard for the Storage, Use, and Handling of Compressed Gases and Cryogenic Fluids in Portable and Stationary Containers, Cylinders and Tanks”, 2005 edition.

### Regulatory Implementation Requirements

None.

## Design Requirements

### Technological Design Requirements

#### Vacuum and Cryogenic Components

Materials used in the construction of components for vacuum environments should have a total mass loss (TML) of ≤ 1%.  Materials used in the construction of components for vacuum environments should have a collected volatile condensable materials (CVCM) value of ≤ 0.1%.  Values for TML and CVCM should be determined in accord with the methods of ASTM standard E595-93 (2003)e1 “Standard Test Method for Total Mass Loss and Collected Volatile Condensable Materials from Outgassing in a Vacuum Environment”.

Materials for use in vacuum and cryogenic environments must be selected for compatibility with the vacuum and the temperatures to be encountered.  Although written primarily for visual wavelength instruments, in particular to protect detector QE in the UV range, guidance in the design and integration of instrument dewars may be obtained from the CARA document “Draft Engineering Guideline for the Design and Integration of Optical Detector Cryostats”.

Where LN2 is used the fill must have an overflow shield so that loss of vacuum does not result from O-ring freezing.

#### Opto-Mechanical Assemblies

Optical and mechanical assemblies, modules or components that must be removed for service shall be provided with locating pins or other features as required to permit repeatable removal and replacement.

Handling features shall be provided on all components unless they are inherently easy to handle without risk of damage.  Handles shall be provided (preferably fixed) for components with weights greater than 1 kg up to 25 kg.  Heavier components and subassemblies shall be provided with lifting eyes or ‘A’ brackets.

#### Electrical/Electronic Assemblies and Enclosures

Service access and regulatory compliance in electronic assemblies and enclosures requires attention to the dimensions of components and the space provided for terminal access, wire bending and component mounting.

The mechanical arrangement of the electronic assemblies within enclosures should be designed using techniques that document the proposed arrangement and permit the verification of accessibility, wire bend radii and electrical spacings.  Computer aided design techniques including solid modeling may be of value in achieving these objectives.

Where possible electrical and electronic subsystems should consist of rack mounted modules conforming to the 19 inch (482.6 mm) width pattern of Electronic Industries Association (EIA) standard 310-D, “Cabinets, Racks, Panels, and Associated Equipment”, section 1.  Where rack mounted modules are used each module should be installed using rack slides.

Where rack mounted equipment can be accessed only from the front all rack slides must extend far enough to permit disconnection of any rear panel connections prior to removal of the rack module from the slides.

In systems that consist predominantly of rack mounted modules, all commercial off the shelf (COTS) modules, components and subsystems that are not available in rack mount configurations should be mounted in suitable rack module chassis or on rack mount shelves.  All rack module chassis and shelves should be mounted on slides.  Components or modules mounted on shelves must be fully enclosed as required to meet all other requirements for grounding, shielding and electrical safety.

Components or modules weighing less than 0.5 kg may be mounted on hinged or screw mounted rack panels provided that all other requirements for grounding, shielding and electrical safety are met.

Rails in 19 inch rack cabinets should be tapped or equipped with captive tapped inserts.  Clip nuts should not be used.

Enclosures for electrical and electronic components must provide a continuous shield to prevent the entry or emission of electromagnetic energy.  No openings greater than 3 mm in diameter or 3 mm in width and 15 cm in length should be permitted on the exterior of any enclosure for electrical and electronic components.  This includes gaps due to access covers, hinges or other enclosure components.  Removable covers that do not make continuous contact with the enclosure must be provided with a fastener every 15 cm or with conductive gaskets or finger stock as described in §8.3.2.2.

Thermal analysis should be performed to ensure that all components operate within their temperature limits and to ensure that excess heat is not transmitted to other components or sub-systems of the instrument.

#### Mechanisms

Mechanisms in the Instrument should be based on as few identical mechanical assemblies as possible.  Mechanisms should be designed in modular assemblies with a minimum of parts and with provisions for simple installation and removal during servicing and repair.

#### Drive Couplings

Shaft couplings for motors, encoders and other drive components should be pinned or locked so that the shaft and coupling cannot slip.  Separable couplings should be used whenever possible for motors to facilitate motor replacement.

#### Component Ratings

Structural elements and fasteners whose failure could cause injury to personnel or equipment must be selected for a safety factor of 10 over ultimate strength of the material.  All other structures and fasteners should be designed with a safety factor of at least 5.

All mechanical moving parts should be selected for a 10 year operating lifetime in the operating environment specified in Table 5.

### Regulatory Design Requirements

None.

### Standards Related Design Requirements

Enclosures for electrical/electronic components and wiring should conform to the requirements of the Underwriters Laboratories Inc. (UL) Standard for Safety 508 “Industrial Control Equipment”.  See §9.3.3.1 for references to the relevant requirements.

All electrical and electronic components should be enclosed in a manner that meets the requirements for a NEMA type 4 or better enclosure.  The requirements of a NEMA type 4 enclosure are given in the National Electric Manufacturers Association (NEMA) standards publication 250-1997, “Enclosures for Electrical Equipment (1000 Volts Maximum)”.

Mechanical drawings should conform to ANSI standard Y14.5M-19994 (R1999)  “Dimensioning and Tolerancing” and ASME standard Y14.100-2000 “Engineering Drawing Practices”.

### Integration Related Design Requirements

#### Mounting Position

#### Handling

The Instrument must be provided with all fixtures and equipment needed to disassemble the Instrument for service.  If required a crane will be provided by the observatory.  The footprint of service fixtures or stands must be minimized because storage and working space on the summit is at a premium.

The profile of all service fixtures or stands must be designed with as low of a center of gravity as possible to resist tipping.  Seismic restraints may also be required.

Handling provisions, fixtures and stands must be designed for safe operation and with consideration for ergonomic factors such as range of motion and working posture.

Any temporary clean room or dust cover facilities required for service should be provided with the instrument.

# Electronic/Electrical Requirements

## Purpose and Objectives

The purpose of this section is to describe requirements for the performance, implementation and design of the Instrument electronic and electrical systems.  In many cases these requirements reflect the preliminary electronic and electrical design of the instrument.

The electronic/electrical requirements address issues of safety, design, reliability and maintainability.  Based on experience with previous instruments the observatory is sensitive to certain aspects of performance, implementation and design that have proven to be important factors in the up time of its instruments.  The electronic/electrical requirements section has as a main objective the description of specific requirements for implementation and design.

A key consideration is the safety of personnel and equipment, and proper electrical design and implementation practices in compliance with recognized standards are an essential aspect of electrical safety.  A second consideration is the electromagnetic compatibility of the instrument with the observatory systems, and specific implementation and design requirements are given to aid in achieving the required electromagnetic emissions and susceptibility performance.

## Performance Requirements

### Parametric Performance Requirements

#### Power Dissipation

See §8.2.1.3.

#### Compatibility

The Instrument must be electrically compatible with the telescope environment.

#### Temperature and Humidity

All electronics should be designed for operation in an ambient temperature range of –10 C to 30 C and a relative humidity of 95%, non-condensing.

#### Cable and Wire Ratings

All wire and cable will be rated for an ambient temperature range of –30 C to 100 C.

### Operational Performance Requirements

None.

## Implementation Requirements

### Feature Implementation Requirements

#### Emergency Stop Input

The Instrument should be provided with an emergency stop input that stops all instrument motion (including a rotator) and closes the entrance window cover when the observatory emergency stop signal is activated.

#### Rotator

A rotator should be equipped with a local control switch to defeat remote control during service and maintenance operations.  The rotator module should be equipped with a motion stop switch to prevent motion of the mechanism during emergencies, service and maintenance.

#### Host Computer

WMKO has standardized on computers running the Sun Solaris operating system as the primary platform for instrument host computers.

This computer should be provided with a dedicated disk array for data storage.  Connections to the observatory wide “public” network and the instrument private network should be made two separate network interfaces in the host computer.  This will isolate the time critical instrument control and data communications from the Observatory wide network traffic. The host computer must be configured to ensure that there are no routes or bridges between the Observatory wide network and the instrument private network.

#### Data Storage

A local, dedicated disk array should be provided for storage of science data where appropriate.  This disk drive should be interfaced to the host computer using a high performance data transfer protocol.  The disk array should utilize a RAID 3 or 5 configurations to permit replacement of a failed disk without loss of data.  The capacity of the disk array should be sufficient to contain data from at least 5 to 7 typical observing nights.

The disk array should use disk drives approved and tested by WMKO for operation at the summit altitude.  The disk array should use front accessible, hot pluggable disk drive modules in an EIA 19 inch rack insert.

#### Target and “Embedded” Computers

If a target computer is a PC type computer located on or integral to the Instrument, it should be an industrial/server grade 1U, 19” EIA rack mount computer equipped with a flash disk as the system disk and running a WMKO approved operating system.  The computer should be equipped with local monitor, mouse and keyboard connections for test and diagnostic purposes.

If a CD-ROM drive is required it should be a removable external drive that is connected when required for maintenance.

If a remote computer is used for a target computer the computer should be a Sun workstation or server running a WMKO approved version of the Solaris operating system.

#### Instrument Connection Panel

A location should be provided with one or more instrument connection panels where all electronic and electrical connections are made.  This panel should also incorporate circuit breakers and other protective devices as required to protect the wiring of the Instrument.  Additional panel(s) for glycol and CCR helium connections should also provided at the same location.

#### Printed Circuit Boards

All removable plug-in printed circuit boards should be equipped with positive retention features.  Extractors should be provided for all circuit boards where high insertion and withdrawal forces are expected.

### Common Practices Implementation Requirements

#### Stray Light

The Instrument should not produce stray light from LED or lamp indicators, optical switches or optical shaft encoders.

LED or lamp indicators should not be used on the exterior of the Instrument.  Any indicators required for service should be concealed behind a cover or access door.  Optical switches or shaft encoders must be optically baffled or enclosed so that no stray visible or infrared light is emitted into the telescope optical path or dome environment.

All exterior parts of the Instrument should be examined for stray light emissions with a night vision device with a light gain of at least 50,000.  A person known to have normal photopic and scotopic visual sensitivity should conduct the examination under dark adapted conditions.

#### Digital Control and Status Communications

Where ever possible digital communications for control and status information between subsystems and modules should be implemented using the TCP/IP protocol over a 100Base-TX Ethernet interface.  Purpose built or custom designed electronic modules and circuits that require such communication should be designed with these protocols.

Where legacy or COTS hardware is used and only serial communications is available, RS-232 signal levels with an asynchronous 8 bit format may be used.  RS-232 data rates should be the maximum practical for the required cable length, and RS-485 levels with electrical isolation (to prevent common mode problems and ground loops) should be used for cable runs longer than 3 meters.

All RS-232 controlled devices should be interfaced to the instrument computers using a terminal server.  The Lantronix ETS8PS is the preferred terminal server at WMKO.

### Standards Implementation Requirements

#### Electrical Safety

The design and construction of the wiring for MOSFIRE should conform to the requirements of UL Standard for Safety 508 “Industrial Control Equipment”.  The relevant portions of UL 508 may be summarized as follows:

- Specific metal gauge requirements are given in tables 6.1 (page 22) and 6.2 (page 23).

- Specific details for doors and covers are given in section 6.4 (pages 24 through 27).

- Specific requirements for the design of ventilation openings are given in section 6.9 (pages 31 through 33).

- Specific details for controlling the accessibility of live parts are given in section 6.17 (pages 36 through 37 and figures on pages 38 and 39).

- Requirements for insulating material that directly supports live parts are given in section 15 (pages 42B through 43).  This includes printed circuit boards.

- Specific requirements for the protection of control circuits are given in section 18.2 (pages 47 through 48B).

- Specific requirements for internal wiring are given in section 21 (pages 50 through 56A).

- Section 34 (page 68) gives specific requirements for the separation of circuits.

- Section 35 (page 68A) gives specific requirements for optical isolators.

- Specific details for required electrical spacings are given in section 36 (pages 68A through 73).

- Specific details for grounding are given in section 40 (pages 79 through 82).

- Table 43.1 (pages 84C through 84E and explanations on pages 84E and 84F) indicates the maximum permissible temperature rises for specific materials and components.

- Table 43.2 (page 86) indicates the ampacity of various insulated conductors.

- Section 49 (pages 99 through 100A) gives the requirements for dielectric voltage-withstand testing.

- Section 62 (pages 128B and 128C) gives specific requirements for device ratings.

- Section 63 (pages 128E through 133) gives specific requirements for markings.  These are summarized in table 67.1 (pages 134A through 136B).

- Additional requirements for programmable controllers are given in sections 177 through 193 (pages 196B through 201)

The design and construction of the wiring should conform to the requirements of the National Electric Code.  The applicable local electric code is the Hawaii County Code 1983, 1995 Edition.  This code adopts the National Electric Code in its entirety and there are no additional special requirements applicable to the locations where the Instrument will be installed or operated.  The requirements given in §9.2.4 are consistent with the applicable portions of the National Electric Code.

#### Electromagnetic Compatibility

Standards exist that specify the test conditions and limits for electromagnetic emissions and electromagnetic immunity.  They do not give information on how to achieve compliance.  In the absence of such information WMKO believes that a satisfactory level of electromagnetic emission and immunity compliance can be achieved by following the requirements given in sections 8.3.2.2, 8.4.1.3 and 9.3.4.5 of this document.

For information on the permitted level of emissions and the required level of immunity the following standards may be consulted:

- The conducted and radiated emissions limits for unintentional radiators are specified in Title 47 CFR Part 15, sections 15.107 and 15.109 for class B devices.

- Electromagnetic immunity requirements are given in the Council of the European Communities Directive EMC 89/336/EEC, and the reference standard of the European Committee for Electrotechnical Standardization (CENELEC) EN 50082-1:1997 “Electromagnetic compatibility-Generic immunity standard-Part 1: Residential, commercial and light industry” published in the Official Journal of the European Community on March 1, 1998.

### Regulatory Implementation Requirements

#### AC Line Connections

All ac line connected parts shall be fully enclosed so as to prevent accidental contact with live parts.  All ac line connections shall utilize UL listed connectors and cables.

All power input connectors shall have an adjacent label indicating the voltage, frequency and current rating for which the equipment is designed.

#### Covers

Removable covers that permit access to circuits with voltages in excess of 36 volts RMS ac or 30 volts dc shall be marked with a warning label.

Removable covers that permit access to circuits of less than 36 volts RMS ac or 30 volts dc that are capable of fault currents in excess of 2 amperes shall be marked with a warning label.

#### Wiring

Internal wiring of 120/208/240 volts ac circuits shall use UL type AWM stranded wire with an insulation thickness of at least 0.8 mm.

The insulation color of internal wiring and the conductors of multi-conductor cable for ac power wiring shall conform to the requirements of the National Electric Code.  The insulation of neutral (grounded) conductors shall be white or gray in color.  Neutral conductors shall be the same size as phase conductors except in cases where two or more phases are provided and harmonic currents are expected, in which case the neutral conductors shall be 125% of the size of the phase conductors.

The insulation of grounding conductors (protective or earth ground) shall be green or green with a yellow stripe.

Grounding conductors shall be the same size as the phase conductors.

Phase, neutral and ground conductors shall be sized using table 43.2 of UL 508.

#### Overcurrent Protection

A fuse or circuit breaker shall internally protect all ac line connected equipment.  When a time delay fuse or time delay breaker is used the rating of the breaker shall not exceed 150% of the continuous full load current of the connected load.  Where a non-time delay fuse is used the rating of the fuse shall not exceed 150% of the continuous full load current of the connected load.  Where an instantaneous trip breaker is used the rating of the breaker shall not exceed 250% of the continuous full load current of the connected load.

The panel where the fuse or circuit breaker is located shall be clearly marked with the type and rating of the protective device.

#### Grounding and Shielding

The enclosures of ac line connected components shall be grounded in conformance with the requirements of the National Electric Code and any local codes.  Grounding conductors shall be continuous and bonded to the enclosure in at least one point.  The grounding point shall be specifically provided for the purpose and shall not be a screw or nut used for mounting components or covers.  Any paint or surface treatment that acts as an insulator shall be removed in order to ensure a good electrical contact for the ground connection.

All components capable of generating electromagnetic emissions in excess of the limits established in the standards referenced in 9.3.3.2 above will be shielded and the shielding grounded to limit electromagnetic emissions to the levels allowed by the standards referenced in 9.3.3.2.  All components susceptible to externally generated electromagnetic emissions in excess of the limits established in the standards referenced in 9.3.3.2 above will be shielded and the shielding grounded to protect those components from unintended operation due to external electromagnetic emissions of the levels established in the standards referenced in 9.3.3.2.

#### Terminations

Crimp terminals and compression screw terminals shall not be used to terminate more than the number of conductors specifically approved for the terminal.  All crimp terminals and screw terminals used for ac line connected wiring must be UL recognized components.  All crimp terminations shall be performed using the manufacturer’s tooling in accord with the manufacturer’s instructions.

#### Altitude Derating

The voltage ratings of relays, switches and insulated cables must be reduced to 80% of their rated value due to the altitude at the summit of Mauna Kea.  Electrical spacings must also be increased by a factor of 1.25 to compensate for the increased altitude.

The normal dielectric withstand test specification for UL approved or listed components for use in ac line connected equipment operating from 120/240 volts ac is 2500 volts AC/60 Hz for one minute.  Voltage ratings for all components should be checked for safety margin with respect to this rating using the following equation:

VI=2∗V+1000AFwhere:VI is the voltage isolation required for the altitudeAF is the altitude factor of 0.8 for 15,000 feetV is the sea level rated working voltage

The resulting value for VI must be less than the dielectric withstand test specification voltage (2500 volts AC) or a dielectric withstand test at altitude must be performed to ensure that the system is safe for the intended application.

## Design Requirements

### Technological Design Requirements

#### Motion Control Systems

Note: revisit this for AO.

The preferred motion controllers for stepper and servomotors are Galil or Pacific Scientific programmable motion controllers.  The preferred motion controller for piezo devices (tip/tilt and focus) is the 500 series from Physik Instrumente.

#### Power Ratings

All power dissipating components to be cooled by free air convection must be derated to 80% of their sea level absolute maximum average power dissipation ratings.

#### Wiring and Interconnections

##### Connector and Cable Mounting

Cable and wiring strain relieves should be designed so that strain relief and wiring integrity is not compromised by opening access doors or removing service access covers.

Connectors should not be mounted on service access covers or on access doors.

##### Cable and Wire Routing

Cables and wiring must be routed so that they do not interfere with the optical path of the instrument.  Cables and wiring must be routed so that full travel of moving or adjustable parts is not affected and does not place a strain on the mounting or connections of any cables or wiring.  Service loops should be provided when necessary, but all cables should be routed neatly and secured at regular intervals with wire ties or lacing cord.

##### Labeling of Interconnections

All external, interconnecting cables and any corresponding panel mounted connectors must be uniquely identified and labeled.  The labeling and identification should be in a clearly visible and non-removable form.  This identification scheme must be identical to that used in the system documentation.  Identification of cables by color-coding is not a substitute for clear labeling.

##### Interconnections

External interconnections of low voltage ac and dc circuits should be shielded whenever there is a reasonable possibility that those interconnections will be subject to electromagnetic interference or unwanted coupling.

Cable shields should be terminated to the connector housings and not via a wire to a connector pin.  Where it may be necessary to isolate shields due to common mode noise problems, cable shield terminations should be made at one end of the cable only, with the end selected for termination being the one that is closest to the point in the system where the zero signal reference potential is grounded.  This is normally the location of the terminating load resistance for signal inputs and the location of the signal source for outputs.

Cable shields should be electrically continuous with the connector housing, and WMKO prefers that no ground pigtails or other wire connections separate from the connector housing be used.  In cases where the design requires different practices those design requirements should be discussed with WMKO.

Where multiple connector pairs of identical type are used each connector pair should be uniquely keyed to prevent accidental interchange of the connections.

All connectors should include pre-grounding pins that make circuit common connections (dc reference or ac protective ground) before all other connections during connector insertion and break circuit common connections (dc reference or ac protective ground) after all other connections during connector removal.

##### Data communications – connectors & formats

Control, science data and guider image data communications between the Instrument and remotely located computers should be via a multi-strand fiber optic bundle.  Fiber optic bundle connections should be via panel mounted connectors equivalent in performance to connectors that conform to military specification MIL-C-38999 series IV.

Science data and guider image data communications may be via proprietary protocols such as those employed with the SDSU-III detector controllers or they may be via high bandwidth industry standard protocols such as Fibre Channel, 1000Base-SX or 100Base-TX.

Control communications between the Instrument and the Instrument target and/or host computers should employ the TCP/IP protocol over a private 100Base-TX network conforming to the Institute of Electrical and Electronics Engineers (IEEE) Standard 802.3U revision 95 “Carrier Sense Multiple Access with Collision Detection (CSMA/CD) Access Method & Physical Layer Specifications: Mac Parameters, Physical Layer, Medium Attachment Units and Repeater for 100 Mb/S Operation (Version 5.0)”.  Cabling and terminations should conform to Telecommunications Industry Association and Electronics Industry Alliance (TIA/EIA) standard TIA/EIA-568-B “Commercial Building Telecommunications Cabling Standards”.

The private network may have a number of devices.  Network devices that are physically part of the instrument should be routed to the remotely located devices in the Keck I computer room (host or target computers) via 100Base-TX switches located on the Instrument and in the Instrument computer rack.   The switches should be interconnected by a 1000Base-SX fiber optic link.

### Regulatory Design Requirements

See §9.3.4.

### Standards Related Design Requirements

Connectors used for low voltage ac and dc circuits should be types equivalent in performance to connectors that conform to military specification MIL-C-38999 series IV.

### Integration Related Design Requirements

#### Rotator

# Safety Requirements

## Purpose and Objectives

Safety is the paramount concern for all activities at the observatory.  Specific regulations apply to health and safety as described in §6.3.3, §9.3.3 and §9.3.4.  The purpose of this section is to provide requirements related to specific safety concerns during the operation and handling of the Instrument.

## Scope

Unless otherwise indicated all of the requirements of this section apply to all components of the Instrument.

## Performance Requirements

### Parametric Performance Requirements

None.

### Operational Performance Requirements

The normal operation of the Instrument must not produce any safety hazard to personnel or equipment.  Interlocks, labeling and procedures must be provided to ensure the safety of personnel and equipment during maintenance and repair.

As part of the processes for the detailed design review and the pre-shipment review the safety of the system will be reviewed.  In general it is expected that conformance to the requirements of this document and the referenced regulatory standards will ensure a safe system.

## Implementation Requirements

### Feature Implementation Requirements

#### Local Control

Revisit.  Mechanisms internal to the Instrument will not be accessible during normal operation.  However, during servicing a means must be provided to ensure that all mechanisms are under local control and remote control is locked out.

A rotator should be equipped with a local control switch to defeat remote control during service and maintenance operations.  The rotator should be equipped with a motion stop switch to prevent motion of the mechanism during emergencies, service and maintenance.  The rotator should also be connected to the Keck telescope emergency stop circuit to disable rotator motion when the emergency stop is activated.

#### Mechanical

All areas of the rotator where exposed moving parts can create a pinch hazard should be clearly marked with a hazard warning label or equipped with shrouds to prevent accidental contact.

The rotator should incorporate a mechanical lockout feature that locks the Instrument in place so that it cannot rotate.  This feature will ensure that the instrument will not move due to an imbalance caused by removal of a component for service.  Mechanical lockout features should activate an electrical lockout consisting of one or more non-defeatable switches that disable the drive system when the mechanical lockout is active and provide a remote indication that the mechanical lockout is active.  The electrical lockout will protect the rotator drive system components as well as prevent unintended drive activation.

#### Entrance Window Cover

The instrument entrance window should be equipped with a remotely operated cover that should be interlocked to the instrument and telescope interlocks so that the window cover is prevented from opening except when the instrument is defined at the operating position.  Special provisions for local operation of the entrance window cover while the instrument is in other positions may be required, but remote operation of the entrance window cover should only be possible when the instrument is defined.

The cover must incorporate safety sensor switches to prevent injury to personnel as it closes.

The cover must be designed to protect the window from damage as described in §8.3.1.6.

#### Electrical

Removable panels that expose voltages in excess of 230 Vac or 500 volts dc should be equipped with defeatable interlock switches that remove all voltages in excess of 36 volts ac or dc from all exposed connections and terminals.

See §9.3.3.1 for additional electrical safety requirements.

### Common Practice Implementation Requirements

None.

### Standards Implementation Requirements

None.

### Regulatory Implementation Requirements

See §6.3.3, §9.3.3 and §9.3.4.

## Design Requirements

### Technological Design Requirements

#### Instrument

No part of any mechanism should move when ac mains power is applied to or removed from the Instrument.  The motion control hardware should inhibit all motion during a power on/reset.

If closed loop or servo systems are used in the motion control systems these servo loops should be designed so that loss of the encoder signal or disconnection of the motor cannot result in a “wind up” of the servo position command.  Software features should be implemented to inhibit motion when the position error measured by the servo controller exceeds the smallest reasonable margin that reflects all of the expected operating conditions.

Limit switches should be closed when not actuated (N.C. contacts).  Motion control software should be designed so that a disconnected limit switch will appear to be active, inhibiting further motion towards that limit.  Motion control software should also be designed so that movement away from an active limit switch is restricted to a reasonable distance past the limit switch actuation point after which motion is stopped and an error indicated due to the apparent failure of the limit switch to open.

If used, position encoders should include a status loop through the connections to the encoder so that in the event of loss of the encoder connection (or intentional disconnection) all motion on the associated axis is inhibited.

#### Rotator

No part of the rotator should move when ac mains power is applied to or removed from the rotator.  The rotator motion control hardware should inhibit all motion during a power on/reset.

The rotator motion control system should be designed so that loss of the encoder signal or disconnection of the motor cannot result in a “wind up” of the servo position command.  Software features should be implemented to inhibit motion when the position error measured by the servo controller exceeds the smallest reasonable margin that reflects all of the expected operating conditions.

Limit switches should be closed when not actuated (N.C. contacts).  Motion control software should be designed so that a disconnected limit switch will appear to be active, inhibiting further motion towards that limit.  Motion control software should also be designed so that movement away from an active limit switch is restricted to a reasonable distance past the limit switch actuation point after which motion is stopped and an error indicated due to the apparent failure of the limit switch to open.

Position encoders should include a status loop through the connections to the encoder so that in the event of loss of the encoder connection (or intentional disconnection) all motion on the associated axis is inhibited.

### Regulatory Design Requirements

As indicated in the sections for overall, mechanical and electrical requirements the design of the Instrument must conform to all applicable regulatory requirements.

### Standards Related Design Requirements

None.

### Integration Related Design Requirements

None.

# Software Requirements

## Purpose and Objectives

The software requirements section describes requirements for performance, implementation and design.  Based on experience with previous instruments the observatory is sensitive to certain aspects of performance, implementation and design that have proven to be important factors in the up time of its instruments.  The software requirements section has as a main objective ensuring compatibility of the Instrument software with existing observatory software systems.  A secondary objective is guiding the selection of software architecture and implementation decisions towards those that fit within the software skill sets at the observatory in order to maximize the ability of the observatory to support and maintain the Instrument software.

WMKO has established a number of standards for software and these standards form an integral part of the software requirements for the Instrument.

Specific requirements are given in areas where repeated problems have affected the availability of instruments.  Among these are issues of network reliability, reliability of fiber optic data connections to detector controllers, and problems with handling errors in a manner that minimizes the loss of observing time by providing useful error messages and avoids total system resets or power cycling to restore proper operation.

## Scope

Unless otherwise indicated all of the requirements of this section apply to all software components of the Instrument.

## Performance Requirements

### Parametric Performance Requirements

#### Reliability

All software components of the Instrument should be tested under simulated operating conditions and should achieve at least 150 hours of continuous operation without a fault.  The reliability of the following software components should be tested and confirmed:

- Host OS

- Target computer(s) OS

- Host application

- Target application(s)

- Detector controller code

#### Fiber Optic Data Links

Fiber optic data links should tolerate up to 10 db of attenuation due to interconnection losses without impairment of performance or reliability.

#### Data Transfer Performance

Data transfer from the Instrument host computer to the disk storage should be at a rate sufficient to keep up with the time to readout and co-add of the minimum useful number of frames taken at the shortest practical exposure time.

#### Display Updates

A display facility for science detector readouts should be provided and this display should update as quickly as possible at the completion of each exposure.

### Operational Performance Requirements

#### Overhead

Software should permit simultaneous motion of multiple mechanisms in order to minimize the time required to complete each instrument set-up between observations.

#### Error Recovery

##### Loss of Network Connections

All Instrument software should gracefully recover from the interruption of TCP/IP network connections, fiber optic connections or USB connections any time.  This disconnection may occur due to physical interruption of the network connection, or the power cycling or hardware reset/reboot of the device at the other end of the network connection.  Software should implement reasonable timeouts and handle all TCP/IP network errors so that recovery from a network fault is as automatic as possible.  Specifically, the components that have not experienced power cycling or a hardware reset/reboot must recover from the loss of the network connection without requiring that they be reset or rebooted.

Whenever possible it is expected that the system will perform in a manner that permits recovery from any of the following conditions without requiring manual resetting of any hardware component and where practical without loss of data (except in the case of the link from detector controller to target where data loss is inevitable and even a pause in the detector readout will typically produce artifacts in the image):

- Loss of network or data connections:

  - Host to target(s)

  - Host to public network

  - Target(s) to detector controller(s)

- Power cycling:

  - Host

  - Target(s)

  - Detector controller(s)

- Hardware resets:

  - Host

  - Target(s)

  - Detector controller(s)

When recovery is not possible, and for the cases where the host computer is not the system being reset or power cycled, it is expected that the user interface software in the system will provide a useful diagnostic message or warning to the operator without crashing or locking up.

##### Detector Controller Aborts

The science detector controller should support aborts at any time during an exposure or during any readout of greater than 5 seconds duration.

##### Data Disk Full

The software will implement some version of certain well known mechanisms for avoiding this (roll-over using DISKLIST when the directory pointed to by OUTDIR is full and so on).  It is understood that there is no requirement to cope with failed NFS cross mounts.

#### Execution Speed and Command Latency

The response time requirements for the Instrument software are given in Table 7.

Table 7: Software Latencies

| Software Function | Goal | Min. | Max. | Units | Notes |
| --- | --- | --- | --- | --- | --- |
| Status requests | 0.1 | - | 0.2 | s |  |
| Motion commands | 0.1 | - | 0.2 | s |  |
| Observatory E-stop | 0.01 | - | 0.05 | s |  |
| Detector controller commands | 0.1 | - | 0.2 | s | 1 |
| Detector controller aborts | > 1 | - | 5 | s | 2 |
| Application software startup and initialization | > 10 | - | 30 | s | 3 |

Notes:

- Not including the exposure or readout times.

- Not including the time elapsed prior to the abort command for the exposure in progress or the readout in progress.

- Not including the actual time required to perform the operating system re-boot and associated initializations.

## Implementation Requirements

### Feature Implementation Requirements

#### User interfaces

Graphical user interfaces (GUIs) should be provided for all observing control functions.  These interfaces must be implemented in a manner consistent with other WMKO instruments and in conformance with KSD 210.  User interfaces based on the OSIRIS heritage are preferred for near-IR science instruments.

If the Instrument user interfaces are written in Java then they should communicate with the Instrument servers using the OSIRIS KTL to Java interface, KJava.

#### Image Display

For near-IR science instruments the image display facility should be the Quicklook II software developed for OSIRIS, and provided with a 2D mode.

Further requirements for image display software are TBD.

#### CSU Configuration

A CSU configuration program is required.  This software should be modeled on the best practices from the processes currently in use at WMKO to design slit masks for the LRIS and DEIMOS instruments.  The configuration software should also include a feature similar to the DEIMOS DSIMULATOR application for previewing slit mask configurations.

The MSCGUI should also support the generation of slit configurations during observing, although for reasons of efficiency this practice will not be encouraged except for special circumstances such as transient object follow up.

#### Data Reduction Pipeline

A data reduction pipeline (DRP) should be provided for use with science data.

While there are important differences in near-IR observing protocols (e.g.: dithering; beam-switching) to cope with OH (and detector) variability that have to be taken into account, the DRP developed for the DEIMOS instrument at WMKO offers a starting point for the development of a DRP.

Further requirements for the DRP are TBD.

#### MOSFIRE Target Computers

All Instrument target computer(s) should be configured to auto-boot their operating systems and auto-execute their target application software and at power on/reset.

#### Software Licenses

Any licensed software required for operation of the Instrument software must be supplied with an adequate number of fully paid licenses to permit operation of all Instrument software.  Node locked licenses will be required for each host or target computer (where applicable).

### Common Practice Implementation Requirements

Instrument host and target software should be written in C/C++ to run under a WMKO approved operating system.  All communications between the instrument software components and the user interfaces and the telescope systems will be based on keywords conforming to WMKO standards.

Where Java is used to develop user interfaces the implementations should be consistent with the OSIRIS implementations.  Java user interfaces must run under the current versions of the Solaris operating systems and Solaris window managers in use at WMKO.

### Standards Implementation Requirements

Instrument software should conform to the requirements of KSD 201 and KSD 210.  All communications between the Instrument target software and the Instrument host software should be via keywords conforming to the requirements of the Keck Task Library (KSD 8).

### Regulatory Implementation Requirements

None.

## Design Requirements

### Technological Design Requirements

#### Client-Server Architecture

The basic architecture of the Instrument software should be based on client-server architecture.  The server components of the system should provide keyword services compliant with the Keck Keyword Interface standards.

#### Communications Protocols

Client-server communications should be via TCP/IP using a WMKO approved protocol.  It is not required that existing message formats or services be used, provided that they are capable of supporting the Keck Task Library (KTL) as described in KSD 8.

Standard implementations of RS-232 serial communications may be used for communication with COTS hardware that does not support TCP/IP network communications.

#### Keywords

Keywords should be defined in collaboration with WMKO software staff.  Keyword values should not be modal or dependent on other values of the keyword.  Keywords should conform to the formats described in KSD 8 and 28.

#### Target Software

Target software is by definition software that provides direct low level control of electronic or electromechanical systems through direct hardware interfaces.  Target software may run on so-called “embedded” computers that are part of the instrument’s electronics hardware, or target software may run on a remote computer connected via data communications interfaces to hardware that has its own embedded computer that runs its own control software and does not directly execute the target software.  See §9.3.1.3 for operating system and computer hardware requirements.

In general target software will implement a keyword service to allow control of the instrument’s electronic or electromechanical systems.  In some cases, such as the rotator target software, the target software may also implement client functionality, for example when monitoring DCS commands to determine rotator position.  Communications with the host software should be via TCP/IP and the Keck Keyword Interface.

#### Host Software

Instrument host software should provide the user interfaces for instrument control and image display.  All host software functions should be accomplished using keywords conforming the to the Keck Keyword Interface standards.

Additional host software design requirements are TBD.

#### Science Data File Formats

Header data for the science data files will incorporate keywords that fully describe the conditions under which the data in the file was taken.

Science FPA mosaic data is to be written as a FITS format file.

### Regulatory Design Requirements

None.

### Standards Related Design Requirements

Software design and coding should comply with KSD 50 and KSD 210.

### Integration Related Design Requirements

None.

# Interface Requirements

## Purpose and Objectives

This section is reserved for interface requirements that are not addressed by other portions of the document.

## Performance Requirements

### Parametric Performance Requirements

None.

### Operational Performance Requirements

#### Handling

See §8.4.4.2.

## Implementation Requirements

### Feature Implementation Requirements

#### Optical Requirements

See §Error: Reference source not found.

#### Mechanical

See §8.3.1.

### Common Practice Implementation Requirements

#### Glycol Cooling

See §8.3.1.7.

#### Vacuum and Cryogenics

See §8.3.1.8 and §8.3.1.9.

#### Stray Light

See §9.3.2.1.

### Standards Implementation Requirements

None.

### Regulatory Implementation Requirements

None.

## Design Requirements

### Technological Design Requirements

None.

### Regulatory Design Requirements

None.

### Standards Related Design Requirements

None.

### Integration Related Design Requirements

#### Optical Interface

See §7.4.4.1.

#### Mechanical Interface

See §8.4.4.

#### Electrical/Electronic Interface

See §9.4.4.1.

# Reliability Requirements

## Purpose

A process should take place to confirm that the Instrument will provide a high level of reliability for a 10 year lifetime.

## Scope

Unless otherwise indicated all of the requirements of this section apply to all components of the Instrument.

## Procedure for Reliability Determination

A recommended procedure to determine the reliability of the Instrument is the use of the reliability prediction models for electronic components and systems given in MIL-HDBK-217F-2 “Reliability Prediction of Electronic Equipment” and the reliability prediction models for mechanical components and systems given in the Naval Surface Warfare Center “Handbook of Reliability Prediction Procedures for Mechanical Equipment”, NSWC 98/LE1.

The MTBF as determined by the prediction models should then be used to establish the operating period before failure based on a 10 year period as follows:

R(t)=exp(−tMTBF)where:R(t)= probability of operation without failure for time tt = time in hoursMTBF =1(all component failure rates)

The probability of operation without failure for the Instrument is expected to be more than 0.90 for this time period (t = 87600 hours).  Software is not included in this requirement or the requested method of reliability assessment.  The reliability of the software to be used with the Instrument can only be determined by testing.

# Spares Requirements

TBD

# Service and Maintenance Requirements

The Instrument must incorporate provisions for disassembly for servicing of internal components.  Handling fixtures and any specialized tools required for servicing must be provided with the Instrument.  A written procedure accompanied by illustrations must be provided for removal and replacement of all major sub-assemblies in the Instrument.

# Documentation Requirements

## Documentation Package

The Instrument should be provided with design, operating and maintenance documentation package including, but not limited to, the following:

- System overview and design description, including details of optical design, mechanical design (including thermal and vacuum design), electrical design and software design.  All design documents shall be supplied in revised form as required to reflect the delivered as-built instrument.

- User’s manual, including but not limited to operating instructions.

- Revised fabrication/procurement drawings, specifications, and schematics that accurately depict the as-built condition of all of the components of the instrument.  All such drawings should be detailed enough to allow fabrication of spare parts should the need arise.

- Bills of material including supplier information for all components of the instrument.

- A maintenance manual, including all information and procedures needed to maintain and operate the Instrument during its lifetime, including but not limited to the following:

  - Procedures for handling, assembly and disassembly of the instrument and all of its components accurately reflecting the as-built instrument.  All assembly instructions shall be clear, and include a tools list, parts lists and check list.

  - Routine maintenance and inspection procedures, as well as a maintenance schedule.

  - Alignment procedures.

  - Troubleshooting guide.

  - Repair procedures.

- Acceptance Test Plan documents, test procedures and all performance data and results of acceptance testing.

- Descriptions of all recommend spare parts and procedures for removal and replacement including written procedures and assembly drawings and exploded view drawings.

- All manufacturer’s manuals and documentation for COTS components.

- All software design documents and related documents including, but not limited to software build and install procedures, source code, release description document, software design document(s), software acceptance testing plans and software user’s manual..  All software design documents and related documents shall be supplied in revised form as required to reflect the delivered as-built instrument software.

- Safety plan and procedures.

## Drawings

### Drawing Standards

All instrument drawings should use the metric standard with dimensions in millimeters.

All instrument drawings should conform to the following:

- Drawings for optical components shall conform to ANSI/ASME standard Y14.18M-1986 “Optical Parts (Engineering Drawings and Related Documentation Practices)”.

- Mechanical drawings shall conform to ANSI Y14.5M-1994 (R1999)  “Dimensioning and Tolerancing” and ASME standard Y14.100-2000 “Engineering Drawing Practices”.

- Each sheet shall conform to ANSI Y14.1-1995 (R2002), “Decimal Inch Drawing Sheet Size and Format”.  Drawing size shall be determined on an individual basis.

- Each drawing shall have a title block with at least the following information:

- Development group

- Drawing number

- Title

- Designer

- Draftsman

- Scale

- Method for determining next higher assembly.

- All drawings shall include parts and materials lists in accordance with ANSI Y14.34-2003, “Parts Lists, Data Lists, And Index Lists: Associated Lists”.  All items shall be identified with an item number or other label (with reference to the drawing number if one exists) for each part or component with all information required for procurement.

- Assembly drawings shall include all relevant views required to clearly define the assembly including isometric and exploded views.

- All detail drawings shall include all views, geometry, dimensions and feature controls required to duplicate the part in accordance with ANSI Y14.5M-1994 (R1999) “Dimensioning and Tolerancing”.

- Multi and sectional view drawings shall be developed in accordance with ANSI Y14.3M-1994 “Multi and Sectional View Drawings”.

- Fluid power system schematics shall be drawn in accordance with ASME Y32.10-1967 (R1994) “Graphic Symbols for Fluid Power Diagrams”.

- Dimensions and tolerances shall be indicated in accordance with ANSI 14.5M-1994 (R1999).

- Surface finishes shall be described in accordance with ANSI 14.5M-1994 (R1999).

- The electronic drawing format shall be at least AutoCAD 2000 (or a more current release).  Drawings created with other computer aided drafting (CAD) software shall be provided in .dxf files compatible with AutoCAD 2000 (or a more current AutoDesk software release).  The preferred CAD software for 3D drawings is AutoDesk Inventor or SolidWorks.

- The electronic drawing format for electrical/electronic schematics and printed circuit board layouts and assembly drawings shall be OrCAD V9.0 or a more current release.  A less desirable alternative is to provide drawings for electrical/electronic schematics and printed circuit board layouts and assembly drawings as AutoCAD 2000 (or a more current release) drawings or as .dxf files compatible with AutoCAD 2000 (or a more current AutoDesk software release).

### Required Drawings

All drawings must be provided as specified in the formats listed above and in the native format if translated to one of the specified formats.

The following drawings should be provided:

- As-built detailed mechanical drawings for all components not commercially available.  Drawings shall provide sufficient detail to fabricate the components to original design intent.

- As-built detailed drawings for all optical components not commercially available.  Drawings shall provide sufficient detail to fabricate the components to original design intent.

- As-built assembly drawings for all assemblies not commercially available along with appropriate detail drawings and assembly tolerances and procedures.

## Electrical/Electronic Documentation

The following documentation for all electrical and electronic assemblies and modules in the instrument should be provided:

- A top level system block diagram.

- An interconnection diagram showing all interconnecting cables and connected assemblies and modules in the instrument.

- An interconnection diagram showing the external connections to the instrument.

- Pinouts and wire color codes for all internal and external connectors and cables.

- Schematics, assembly drawings, bills of material, printed circuit board designs and printed circuit board artwork for all custom printed circuit boards in the instrument.

- Programmable logic device source code for all programmable logic devices used on custom printed circuit boards in the instrument.

- Programmable logic device source code for all programmable logic devices used in COTS components where the programmable logic device source code has been modified or customized for the instrument.

- Configuration, set up and/or switch/jumper setting information for all COTS components.

## Software

The instrument software is defined as all host, target, embedded controller software (including detector controller code) and data reduction software for the instrument including the code for servo controls including DSP code, PMAC code or other motion control code and the like.  The following software data files and documentation should be provided:

- Source code for all instrument software on CD/DVD.

- Executables for all instrument software on CD/DVD.

- One copy of any and all software libraries required to build the instrument software executables on CD/DVD.

- A list of any and all code compilers required to build the instrument software.

- All makefiles required for building the instrument software on CD/DVD.

- All configuration files and all data files read by the instrument software executables at start-up time on CD/DVD.

- Any scripts required to run the instrument or the data reduction package on CD/DVD.

- Any aliases, environment variable definitions, etc. required to correctly set up the environment to build or run the instrument software on CD/DVD.

- Any models developed for simulation of the instrument including optical designs and control loop simulations should be supplied.  The preferred software for optical design is Zemax.  The preferred software for simulations is Matlab or IDL.

- Full design documentation for all control loops including block-diagrams, transfer-function models of the system, performance criteria and analyses to show how these requirements are met. Models and simulations of the control loops should also be provided.

- Documentation for the instrument software, consisting of:

  - Users Manual: a detailed tutorial describing how to use this version of the software.

  - List of Source Code: A hierarchical list of all directories, source files, include files, libraries, etc that can be used as a checklist for new releases.

  - Functional Descriptions: a description of each routine or module describing its function.

  - Startup/Shutdown procedures: descriptions of the steps necessary to cold start the system and the steps necessary to safely shut down a running system.  This document should include descriptions of any configuration files required at start-up time.

  - Installation Manual: a detailed description of the steps necessary to rebuild and install the system from sources.

  - Troubleshooting Guide: A description of the techniques for tracking down failures, checking system health, killing and re-starting portions of the system without a full reboot.

  - Software Test Procedures:  a detailed description of how to run the software acceptance tests.

  - Programmer’s Manual: This document shall include a description of the theory of operations; data and control flow and how standard functionality can be extended (e.g. add a new command to the API).

# Glossary

Table 8 defines the acronyms and specialized terms used in this document.

Table 8: Glossary of Terms

| Term |  | Definition |
| --- | --- | --- |
| ANSI |  | American National Standards Institute |
| ASME |  | American Society of Mechanical Engineers International |
| ASTM |  | ASTM International |
| ATA |  | Air Transport Association |
| CCR |  | Closed Cycle Refrigerator |
| CENELEC |  | European Committee for Electrotechnical Standardization |
| CFR |  | Code of Federal Regulations |
| CIT |  | California Institute of Technology |
| COTS |  | Commercial Off The Shelf |
| CSU |  | Configurable Slit Unit |
| CVCM |  | Collected Volatile Condensable Materials |
| dBA |  | Sound level in decibels, measured using the A contour frequency weighting network |
| DCS |  | Drive and Control System |
| DEIMOS |  | DEep Imaging Multi-Object Spectrograph |
| EIA |  | Electronic Industries Alliance |
| EMI |  | Electro Magnetic Interference |
| FOV |  | Field Of View |
| FPA |  | Focal Plane Array |
| FWHM |  | Full Width at Half Maximum. |
| IBC |  | International Building Code |
| ICC |  | International Code Council |
| ICD |  | Interface Control Document |
| IEEE |  | Institute of Electrical and Electronics Engineers |
| KSD |  | Keck Software Document |
| MTBF |  | Mean Time Between Failures |
| NEBS |  | Network Equipment Building System |
| NEMA |  | National Electric Manufacturers Association |
| NIR |  | Near InfraRed |
| OSHA |  | Occupational Safety and Health Administration |
| RT1 |  | Rail Transport position 1, (Nasmyth deck, Keck I) |
| SSC |  | Science Steering Committee |
| TBC |  | To Be Completed |
| TBD |  | To Be Determined |
| TIA |  | Telecommunications Industry Association |
| TML |  | Total Mass Loss |
| UCLA |  | University of California, Los Angeles |
| UCSC |  | University of California, Santa Cruz |
| UPS |  | Uninterruptible Power Supply |
| UL |  | Underwriters Laboratories Inc. |
| USGS | United States Geological Survey |  |
| WMKO | W. M. Keck Observatory |  |
| WRT | With Respect To |  |

Gordon, Colin G. Generic Criteria for Vibration-Sensitive Equipment. Proceedings of the SPIE Vol. 1619, pp. 71-85, Vibration Control in Microelectronics, Optics, and Metrology.  Gordon, Colin G. editor.  SPIE 1992.

This is a typical specification for generation III night vision monoculars such as the ITT 160 Night Mariner.