Keck Adaptive Optics Note 669

Next Generation Adaptive Optics:

A Slow Tip-Tilt Mount for the Narrow-Field Deformable Mirror

P. Wizinowich

July 21, 2009

1. Introduction

The concept of mounting the NGAO narrow field deformable mirror (DM) on a tip-tilt mount to steer the light beam on the science instrument was proposed in KAON 558: “One can imagine steering the light beam on and around the science array with high accuracy using a steering mirror or a probe arm installed in the science light path. Any move can be very fast (very likely < 2 sec), and the steering mirror could be used for fine guidance/centering on a slit, a spaxel, a coronagraphic spot or any location on a detector. This steering mechanism could be used to compensate for differential atmospheric refraction, differential tracking, or any effect leading to centering and positioning error on the science path. The AO loops would remained closed as long as required on the LOWFS, the high order WFS, etc.”  The purpose of this KAON is to further develop this concept and to better understand its implications on the rest of the NGAO system.

2. Preliminary Calculations

Assumptions:

- Pupil diameter on MEMS: 23 mm (slightly under the ~ 25 mm MEMS diameter; corresponds to 10 m on the telescope primary mirror).

- Focal ratio at MEMS: f/46.  Plate scale at the focal plane after the MEMS DM is therefore 46/15 * 0.727 mm/arcsec = 2.23 mm/arcsec.

Calculations:

- Ratio of image motion to DM tilt: A 1 arcsec motion in the image plane is produced by a ½ * [1 arcsec * (46/15 * 0.727 mm/arcsec)] / (46 * 23 mm) = 1.05 milli-radians = 217 arcsec tilt of the DM.

- Shift of MEMS DM actuators with respect to LGS WFS lenslets:

  - For the case of 1 arcsec image motion the actuator at the edge of the pupil is shifted by (23 mm / 2) * 1.05 milli-radians = 12 m towards/away from the focal plane and by (23/2) * [1-(1 – 0.001052)1/2] = 6.4 nm laterally (negligible).

  - For a 10 arcsec image motion the actuator at the edge of the pupil is shifted by 121 m towards/away from the focal plane and by 64 nm laterally.  The lateral displacement is insignificant with respect to the 400 m actuator spacing.

3. Dithering Comparison

The following table summarizes the differences between re-pointing the telescope and tilting the narrow field DM to perform a dither.  Using the NF DM tilt to dither has a clear observing efficiency advantage.  It may also offer some scientific advantages by keeping the optimal AO correction centered on the science target (item 2), maintaining the AO performance at least for small moves (item 5) and better maintaining the optical path to the science target (item 3b).

3. Differential Tip-Tilt Correction Comparison

The following table summarizes the differences between using the LOWFS or NGS WFS and tilting the narrow field DM to compensate for differential motion between the NGS used by the LOWFS/TWFS or NGS WFS and the science target.  These effects include differential atmospheric refraction, differential tracking (non-sidereal science target), pointing corrections for the science path atmospheric dispersion corrector and drifts of the science instrument with respect to the AO system.  Note that tilt of the narrow field DM would not be able to be used for angle tracker offloads from the interferometer in LGS mode.  The performance impacts of using the different approaches were already mentioned in the previous table.

A major advantage of using tilt of the narrow field DM to compensate for differential motion would be not requiring the LOWFS units to track.  This could potentially significantly reduce the requirements on the LOWFS pickoff arms, LOWFS tip-tilt mirrors and LOWFS field of view.  Having one tracking unit instead of three could also reduce costs.

5. Draft Tip-Tilt Stage Performance Requirements

KAON 548 provides the following relevant requirements:

- Differential tracking:  50"/hr.

- Acquisition accuracy:  10% of IFU field,  35 mas relative for IFU &  200 mas for the interferometer.

- Dither distance:  50% of field.

- Dither accuracy:  /D.

- Dither time:  3 sec.

- Micro dither distance:  0.5/D.

- Micro dither accuracy:  0.25/D.

- Micro dither time:  3 sec.

- Positioning knowledge:  0.1/D or  5 mas for the IFU.

- Science image drift:  5 mas.

Tip-Tilt Stage Requirements (based on the above requirements and discussion):

- Tip-tilt range:  2170" for the tip-tilt stage.  This corresponds to  10" radius in the focal plane which is half the maximum field radius.  Based on calculations earlier in this document a 10” radius should not impact performance. This range would allow differential tracking for  24 minutes before having to reposition the LOWFS.  Larger dithers would require a telescope move.

- Tracking rate:  10,850"/hr for the tip-tilt stage.  This corresponds to 50"/hr of image motion.  The rate on the stage is up to 3.0"/sec.

- Positioning accuracy: 4.34" for the tip-tilt stage for anywhere in a 2170" radius.  This corresponds to /D = 20 mas at 1 m in the image plane.

- Positioning time:  3 sec for moves anywhere in a 2170" radius for the tip-tilt stage.  Should consider a goal of 2 sec.

- Micro-positioning accuracy: 1.09" for the tip-tilt stage for moves of up to 2.17".  This corresponds to 0.25/D = 5 mas at 1 m for moves of 0.5/D at 1 m or 2m in the image plane.

- Positioning knowledge:  0.43" for the tip-tilt stage.  This corresponds to  0.1/D = 2 mas in the image plane at 1 m.  This likely requires a load encoder with at least 2x this resolution.

- Tracking accuracy:  1.09" for the tip-tilt stage.  This corresponds to  5 milli-arcsec in focal plane.

- Tracking update rate:  40 Hz.

- DM position stability:  0.05 of the actuator spacing in the plane of the DM.  This corresponds to 20 m at the DM assuming 400 m actuator spacing.

- Must be able to support the mass of the MEMS DM and cable connections at the DM.

- Must not introduce vibrations into the MEMS DM while science exposures are in progress.  Note that some initial vibration during a re-positioning should be acceptable if it ends during the allotted move time.

6. Conclusion

The scientific benefits, operational simplicity and reduced LOWFS complexity of mounting the narrow field DM on a slow (i.e., non-RTC) tip-tilt stage are all significant enough to warrant including this stage in the NGAO baseline design.