## radiation
FOR CFH12K ADAPTED TO P48
RADIATIVE LOAD ON CCDs
Radiative transfer from window to CCDs should be the dominant heat source, but
to this one must add radiation to sheild, conduction in supports and wiring, and heater power
Beware: holes & cracks in shield have emissivity = 1.
Number of CCDs in X | 6
Number of CCDs in Y | 2
gap in and Y | 0.003 | m
CCD pixels in X | 2048 | pix/side
CCD pixels in Y | 4096
pixel size in X and Y | 1.5e-05 | m/pix
Radiation shield opening in X | 0.20532 | m | …need to add effect of f-number
Radiation shield opening in Y | 0.13188 | m
CCD area | 0.027078 | m^2
Stefan-Boltzmann const | 5.67e-08 | W/m^2/K^4
T | 287 | K |  | assuming radiative cooling most of area by 8C. (12C was predicted over about 80%)
Total Black body flux | 384.6898 | W/m^2
CCD emissivity (wild guess) | 100%
Radiative load single detector | 0.72608 | W | example of single large detector
Radiative load on CCD mosaic | 10.4165 | W | ..should add effect of cracks which have emissivity =1
Peak cooling by CryoTiger | 35 | W | at 153K
Additional loads: |  |  | radiation to other surfaces
conduction in wiring to CCDs
conduction though focal plane supports
N2 precool tubing … doesn’t have the usual benefit of boiloff gas.
CCD & window emissivities | 0.95
Radiative load on all CCDs |  | 9.4009 | W
radiation shield | 3% | 0.982 | W | assume 19" diameter, 16" high
conduction through supports |  | 2 | W | wild guess; do the calculation for the f.g. posts
wiring | 0.15 | 1.8 | W | based on experience, CFH desin may differ
output FETs | 0.03 | 0.36 | W | 2 source followers in series, 0.5 and 1.0mA
heater power |  | 4.7004 | W | arbitrarily set to half frontside radiative load
Total |  | 19.2433 | W
Load without heater |  | 14.5429 | W |  | load on LN2 | 14.4613 | W
[5 empty] | volume | 10.3737 | l | center fill
[5 empty] | tank area | 0.18292 | sq.m
[5 empty] | tank height | 0.056711 | m
[6 empty] | 2.2327 | inch
Nitrogen
[5 empty] | density | 0.807 | kg/l
[5 empty] | vaporization | 199 | kJ/kg
[6 empty] | 160.593 | kJ/l
[6 empty] | 44.6092 | W*hr/l
[6 empty] | 0.022417 | l/w/hr
sq.m | 0.0024159
Power | 0.92938

## Heat capacity
HEAT CAPACITY
|  |  | at 240K |  |  | From Hal's SolidWorks Model |  | Cross-check | Total
Component | Number | Material | Specific Heat | Density |  | Volume each | Volume each | Mass each | Heat Capacity
|  |  | (J/K/kg) | (kg/m^3) | (J/K/m^3) | (inch^3) | (m^3) | (lb) | (J/K)
Detector assembly at 300k (pessimistic)
FPMOLY | 1 | Moly | 235 | 10240 | 2406400 | 6.615 | 0.0001084 | 2.445 | 260.8564
MOLYFOOT | 12 | Moly | 235 | 10240 | 2406400 | 0.49 | 8.0297e-06 | 0.18111 | 231.8723
COLDCLAMP | 12 | Moly | 235 | 10240 | 2406400 | 0.106 | 1.737e-06 | 0.039179 | 50.1601
CCD20 | 12 | AlN | 580 | 3250 | 1885000 | 0.729 | 1.1946e-05 | 0.085518 | 270.224
Standoff Pad | 4 | Al ? | 885 | 2770 | 2451450 | 0.0984 | 1.6125e-06 | 0.0098384 | 15.8118
FPBRACKET | 4 | Al ? | 885 | 2770 | 2451450 | 0.535 | 8.7671e-06 | 0.053491 | 85.9687
Resistors | 12 | AlN | 580 | 3250 | 1885000 | 0.09 | 1.4748e-06 | 0.010558 | 33.361 |  | guess
Fasteners | 60 | SS 304 | 450 | 8030 | 3613500 | 0.003927 | 6.4352e-08 | 0.0011382 | 13.9522 |  | guess
TOTAL at Detector | [5 empty] | 8.6673 | 0.00014203 | 2.8258 | 962.2065 |  |  | 0.023095 | K/s | 1.3857
[9 empty] | 58.4458 | singlle CCD+spacer+R |  |  |  | 1.9011
Cryocooler Head, at 128K
Heat spreader |  | Pure Copper | 300.7 | 8860 | 2664202 |  | 0 | 0 | 0
Bolts for straps |  | SS 304 | 332 | 8030 | 2665960 |  | 0 | 0 | 0
getter mesh |  | SS | 332 | 8030 | 2665960 |  | 0 | 0 | 0
getter bolts |  | SS 304 | 332 | 8030 | 2665960 |  | 0 | 0 | 0
getter clamp |  | Copper | 300.7 | 8860 | 2664202 |  | 0 | 0 | 0
getter material |  | charcoal | neglect
head itself |  | Copper | 300.7 | 8860 | 2664202 |  | 0 | 0 | 0
[6 empty] | 0 | 0 | 0 | 300.7 |  | assuming 1 kg of Cu
CONDUCTANCES FROM FOCAL PLANE
|  |  | 151 | K |  |  | per strap
Component | Number | Material | Conductivity | Width | Thickness | Length | Thermal R |  |  | Base Load | Unheated T | Eqbm Load | Eqbm Td | Eqbm Tc
|  |  | W/m/K | (m) | (m) | (cm) | (K/W) |  |  | (W) | (K) | (W) | (K) | (K)
Cu straps old |  | Copper OFHC | 418 | 0.00764 | 0.00014 | 0.138 | 308.6608 | dimensions measured |  | 11.6361 | 193
Contact R per joint | [6 empty] | 5 | guess
predicted total link R | [6 empty] | 13.8992 | K/W
…based on load and T | [6 empty] | 9.969 | K/W | 1.3942
Cu straps new |  | Copper OFHC | 418 | 0.02 | 0.000254 | 0.08618 | 40.5852 |  |  |  |  | 21.9093 | 175 | 126.7995
total link R (new) | [6 empty] | 2.7099 | K/W
target | [6 empty] | 2.2 | K/W | 1.2318
| original | new
Tdet_min | 193 | 163 | K
Tcooler | 77 | 127 | K
Tdelta | 116 | 36 | K
Cu strap length | 138 |  | mm
| At 128K |  |  |  | Heat spreader dimenions (one prong)
| Al 1100 | Cu OFHC |  |  | l | h | w |  | Th Resistance | mass | delta T
(kg)
conductivity | 234 | 432 | W/m/K |  | 0.10478 | 0.02 | 0.01 |  | 1.2127 | 1.0788 | 1.7638
density | 2770 | 8580 | kg/m^3
density/conductivity | 0.084477 | 0.05035
strap # | power | area | delta T | T - T0
1 | 6 | 1.7143 | 3.5 | 3.5 |  | Strap pitch on spreader |  |  | 0.031242 | m
2 | 5 | 1.4286 | 3.5 | 7 |  | Spreader width between last 2 |  |  | 0.0127 | m
3 | 4 | 1.1429 | 3.5 | 10.5 |  | Spreader thickness |  |  | 0.01 | m
4 | 3 | 0.85714 | 3.5 | 14 |  | Th Resistenace between last 2 |  |  | 0.56944 | K/W
5 | 2 | 0.57143 | 3.5 | 17.5 |  | Total power to cooler |  |  | 16 | W
6 | 1 | 0.28571 | 3.5 | 21 |  | Power flow per strap |  |  | 1.3333 | W
[6 empty] | Delta T from cooler to tip |  |  | 2.2778 | K
6
[6 empty] | Approx mass of spreader |  |  | 0.97363 | Kg
HEATER
6 in parallel, in series with another 6 in parallel
each resistor | 150 | ohms
allowable P | 5 | W
Total resistance | 50 | ohm
Lakeshore specs:
I_max_heater | 2 | A
V max_heater | 50 | V
P_max_heater | 100 | W
P max_actual_total | 50 | W | …limit this in software and add thermal fuse !
P_max_actual_per_R | 4.1667 | W
I_max_actual | 1 | A
Link thermal R | 3.0938 | K/W
Temperature range | 154.6916 | K
T_det | 173
T_head | 127
P_base | 8.5
P_heater | 5
P_total | 13.5
R_th | 3.4074
Note on A8: Roger Smith: FPMOLY, "Focal plane"
Note on C8: Roger Smith: Confirmed by Luppino not to be TZM
Note on G35: Roger Smith: hole to nearest hole

## Cooling curves
Standard PT 30 Cryocooler
Heat inflow model:
Ambient temperature |  | Ta | 295 | K
Detector temperature setpoint |  | Tsp | 175 | K
detector tempertaure |  | Td | variable
Cold head temperature |  | Tc | variable
Detector to Cold head Resistance |  | Rd | 2.2 | K/W
Wiring thermal resistance |  | Rw | 32.6667 | K/W
Post thermal resistance |  | Rp | 1989.4368 | K/W
Stephan boltzman constant |  | sigma | 5.669e-08 | W/sq.m.K^4
Enclosure emissivity |  | Ea | 0.15
window emissivity |  | Ew | 0.97
Detector assy backside emsissivity |  | Eb | 0.3
Detector assy frontside emissivity |  | Ef | 0.97
Detector assy back+side area |  | Ab | 0.054155 | sq.m
Detector assy front surface area |  | Af | 0.027078 | sq.m
heat inflow to cold head = sum of…
conduction from ambient = |  | (Ta -Tc)*(1/Rw + 1/Rp)
conduction from detector = |  | (Td-Tc)/Rd
radiation: amb to det. = |  | sigma*(Af*Ef*Ew + Ab*Eb*Ec)(Ta^4 -Td^4)
Room temp heat capacity at detector |  | Cd | 962.2065 | J/K .. from heat Capacity sheet
heat capcity of cryotiger tip |  | Ct | 135.16 | J/K ..from cooling time on data sheet and simulation
heat capacity of spreader+getter |  | Cc | 300.7 | J/K .. from heat Capacity sheet
Thermal time constant of focal plane |  |  | 3278.6297 | s
Slew rate with heater Vmax=50V |  |  | 1.3857 | K/min
If CCD heaters rewired to 3 groups in series of four 150 ohm resistors in parallel = 112.5 ohm
Max slew rate for single CCD+spacer+R |  |  | 1.9011 | K/min
Notes:
The detector tempertaure servo has been assumed to be perfect once setpoint is reached passively
There is some unphysical oscillation at small amplitude in the simulation due to the use of lookup tables which quantize. | [9 empty] | 4.8611 | hours | Fastest rate=
[6 empty] | 15 | s increment |  |  |  |  | -0.60252
Cold head temperature, Tc | Cooling power | heat capacity at cooler head |  | Detector temperature, Td | heat capacity at detector | time since start cooling | net heat into detector | Td | Heat flow through heat straps, Rd | net heat into heat capacity of cooling head | Tc | CCD cooling rate | Radiation | conduction | Radiation plus conduction | Heater power
(K) | (W) | (J/K) |  | (K) | (J/K) | (s) | (W) | (K) | (W) | (W) | (K) | (K/min) | (W) | (W) | (W) | (W)
105 | 0 | 435.86 |  | 105 | 962.2065 | 0 | 0 | 295 | 0 | -14.5 | 295 |  | 0 | 0 | 0 | 0
106 | 0.83333 | 435.86 |  | 106 | 962.2065 | 15 | -0.22682 | 295 | 0.22682 | -14.2732 | 294.501 | 0 | 0 | 0 | 0 | 0
107 | 1.6667 | 435.86 |  | 107 | 962.2065 | 30 | -0.44781 | 294.9965 | 0.44849 | -14.0515 | 294.0098 | -0.014144 | 0.0005746 | 0.00011002 | 0.00068462 | 0
108 | 2.5 | 435.86 |  | 108 | 962.2065 | 45 | -0.66309 | 294.9895 | 0.66513 | -13.8349 | 293.5262 | -0.027924 | 0.0017089 | 0.00032723 | 0.0020362 | 0
109 | 3.3333 | 435.86 |  | 109 | 962.2065 | 60 | -0.87281 | 294.9791 | 0.87685 | -13.6232 | 293.0501 | -0.041348 | 0.0033885 | 0.00064887 | 0.0040373 | 0
110 | 4.1667 | 435.86 |  | 110 | 962.2065 | 75 | -1.0771 | 294.9655 | 1.0838 | -13.3662 | 292.5812 | -0.054426 | 0.0055989 | 0.0010722 | 0.0066711 | 0
111 | 5 | 435.86 |  | 111 | 962.2065 | 90 | -1.2753 | 294.9487 | 1.2852 | -13.1648 | 292.1212 | -0.067164 | 0.0083263 | 0.0015947 | 0.009921 | 0
112 | 5.8333 | 435.86 |  | 112 | 962.2065 | 105 | -1.4684 | 294.9289 | 1.4821 | -12.9179 | 291.6682 | -0.079524 | 0.011555 | 0.0022133 | 0.013768 | 0
113 | 6.6667 | 435.86 |  | 113 | 962.2065 | 120 | -1.6556 | 294.906 | 1.6738 | -12.7262 | 291.2236 | -0.091562 | 0.015272 | 0.0029255 | 0.018197 | 0
114 | 7.5 | 435.86 |  | 114 | 962.2065 | 135 | -1.838 | 294.8802 | 1.8611 | -12.4889 | 290.7856 | -0.10324 | 0.019461 | 0.0037286 | 0.02319 | 0
115 | 8.3333 | 435.86 |  | 115 | 962.2065 | 150 | -2.0148 | 294.8515 | 2.0435 | -12.3065 | 290.3558 | -0.11461 | 0.024111 | 0.0046201 | 0.028731 | 0
116 | 9.1667 | 435.86 |  | 116 | 962.2065 | 165 | -2.1869 | 294.8201 | 2.2217 | -12.0783 | 289.9323 | -0.12563 | 0.029206 | 0.0055974 | 0.034804 | 0
117 | 10 | 435.86 |  | 117 | 962.2065 | 180 | -2.3538 | 294.786 | 2.3952 | -11.9048 | 289.5167 | -0.13637 | 0.034735 | 0.0066581 | 0.041393 | 0
118 | 10.8333 | 435.86 |  | 118 | 962.2065 | 195 | -2.5162 | 294.7493 | 2.5647 | -11.7353 | 289.107 | -0.14677 | 0.040684 | 0.0077998 | 0.048484 | 0
119 | 11.6667 | 435.86 |  | 119 | 962.2065 | 210 | -2.6744 | 294.7101 | 2.7305 | -11.5195 | 288.7031 | -0.1569 | 0.047041 | 0.0090203 | 0.056061 | 0
120 | 12.5 | 435.86 |  | 120 | 962.2065 | 225 | -2.8276 | 294.6684 | 2.8917 | -11.3583 | 288.3066 | -0.16677 | 0.053794 | 0.010318 | 0.064112 | 0
121 | 13.3333 | 435.86 |  | 121 | 962.2065 | 240 | -2.9767 | 294.6243 | 3.0494 | -11.1506 | 287.9158 | -0.17632 | 0.060932 | 0.011689 | 0.072621 | 0
122 | 14.1667 | 435.86 |  | 122 | 962.2065 | 255 | -3.1211 | 294.5779 | 3.2027 | -10.9973 | 287.532 | -0.18562 | 0.068442 | 0.013133 | 0.081575 | 0
123 | 15 | 435.86 |  | 123 | 962.2065 | 270 | -3.2616 | 294.5293 | 3.3526 | -10.8474 | 287.1535 | -0.19462 | 0.076313 | 0.014647 | 0.09096 | 0
124 | 16.6667 | 435.86 |  | 124 | 962.2065 | 285 | -3.3984 | 294.4784 | 3.4992 | -10.6508 | 286.7802 | -0.20339 | 0.084534 | 0.016229 | 0.10076 | 0
125 | 18.3333 | 435.86 |  | 125 | 962.2065 | 300 | -3.5307 | 294.4254 | 3.6417 | -10.5083 | 286.4137 | -0.21191 | 0.093095 | 0.017877 | 0.11097 | 0
126 | 20 | 435.86 |  | 126 | 962.2065 | 315 | -3.6595 | 294.3704 | 3.7811 | -10.3689 | 286.052 | -0.22016 | 0.10198 | 0.01959 | 0.12157 | 0
126.2 | 20.5714 | 435.86 |  | 126.2 | 962.2065 | 330 | -3.7848 | 294.3133 | 3.9173 | -10.1827 | 285.6952 | -0.22819 | 0.11119 | 0.021365 | 0.13256 | 0
126.4 | 21.1429 | 435.86 |  | 126.4 | 962.2065 | 345 | -3.9059 | 294.2543 | 4.0498 | -10.0502 | 285.3448 | -0.23601 | 0.12071 | 0.023201 | 0.14391 | 0
126.6 | 21.7143 | 435.86 |  | 126.6 | 962.2065 | 360 | -4.0237 | 294.1935 | 4.1793 | -9.8707 | 284.9989 | -0.24356 | 0.13053 | 0.025095 | 0.15562 | 0
126.8 | 22.2857 | 435.86 |  | 126.8 | 962.2065 | 375 | -4.1376 | 294.1307 | 4.3052 | -9.7448 | 284.6592 | -0.25091 | 0.14063 | 0.027047 | 0.16768 | 0
127 | 22.8571 | 435.86 |  | 127 | 962.2065 | 390 | -4.2483 | 294.0662 | 4.4284 | -9.6216 | 284.3238 | -0.258 | 0.15102 | 0.029054 | 0.18007 | 0
127.2 | 23.4286 | 435.86 |  | 127.2 | 962.2065 | 405 | -4.356 | 294 | 4.5488 | -9.4512 | 283.9927 | -0.26491 | 0.16168 | 0.031115 | 0.19279 | 0
127.4 | 24 | 435.86 |  | 127.4 | 962.2065 | 420 | -4.4599 | 293.9321 | 4.6658 | -9.3342 | 283.6674 | -0.27162 | 0.1726 | 0.033228 | 0.20582 | 0
127.6 | 24.5714 | 435.86 |  | 127.6 | 962.2065 | 435 | -4.561 | 293.8626 | 4.7802 | -9.2198 | 283.3462 | -0.27811 | 0.18377 | 0.035391 | 0.21916 | 0
127.8 | 25.1429 | 435.86 |  | 127.8 | 962.2065 | 450 | -4.6593 | 293.7915 | 4.8921 | -9.1079 | 283.0289 | -0.28441 | 0.19519 | 0.037603 | 0.23279 | 0
128 | 25.7143 | 435.86 |  | 128 | 962.2065 | 465 | -4.7548 | 293.7188 | 5.0015 | -8.9485 | 282.7155 | -0.29054 | 0.20684 | 0.039863 | 0.2467 | 0
128.2 | 26.2857 | 435.86 |  | 128.2 | 962.2065 | 480 | -4.8469 | 293.6447 | 5.1078 | -8.8422 | 282.4075 | -0.2965 | 0.21872 | 0.04217 | 0.26089 | 0
128.4 | 26.8571 | 435.86 |  | 128.4 | 962.2065 | 495 | -4.9364 | 293.5692 | 5.2118 | -8.7382 | 282.1032 | -0.30224 | 0.23083 | 0.044521 | 0.27535 | 0
128.6 | 27.4286 | 435.86 |  | 128.6 | 962.2065 | 330 | -3.7848 | 294.3133 | 3.9173 | -10.1827 | 285.6952 | -0.27062 | 0.11119 | 0.021365 | 0.13256 | 0
128.8 | 28 | 435.86 |  | 128.8 | 962.2065 | 345 | -3.9059 | 294.2543 | 4.0498 | -10.0502 | 285.3448 | -0.23601 | 0.12071 | 0.023201 | 0.14391 | 0
129 | 29.5 | 435.86 |  | 129 | 962.2065 | 360 | -4.0237 | 294.1935 | 4.1793 | -9.8707 | 284.9989 | -0.24356 | 0.13053 | 0.025095 | 0.15562 | 0
130 | 30 | 435.86 |  | 130 | 962.2065 | 375 | -4.1376 | 294.1307 | 4.3052 | -9.7448 | 284.6592 | -0.25091 | 0.14063 | 0.027047 | 0.16768 | 0
131 | 30 | 435.86 |  | 131 | 962.2065 | 390 | -4.2483 | 294.0662 | 4.4284 | -9.6216 | 284.3238 | -0.258 | 0.15102 | 0.029054 | 0.18007 | 0
132 | 29.8 | 435.86 |  | 132 | 962.2065 | 405 | -4.356 | 294 | 4.5488 | -9.4512 | 283.9927 | -0.26491 | 0.16168 | 0.031115 | 0.19279 | 0
133 | 29.5 | 435.86 |  | 133 | 962.2065 | 420 | -4.4599 | 293.9321 | 4.6658 | -9.3342 | 283.6674 | -0.27162 | 0.1726 | 0.033228 | 0.20582 | 0
134 | 29.2 | 435.86 |  | 134 | 962.2065 | 435 | -4.561 | 293.8626 | 4.7802 | -9.2198 | 283.3462 | -0.27811 | 0.18377 | 0.035391 | 0.21916 | 0
135 | 28.9 | 435.86 |  | 135 | 962.2065 | 450 | -4.6593 | 293.7915 | 4.8921 | -9.1079 | 283.0289 | -0.28441 | 0.19519 | 0.037603 | 0.23279 | 0
136 | 28.6 | 435.86 |  | 136 | 962.2065 | 465 | -4.7548 | 293.7188 | 5.0015 | -8.9485 | 282.7155 | -0.29054 | 0.20684 | 0.039863 | 0.2467 | 0
137 | 28.3 | 435.86 |  | 137 | 962.2065 | 480 | -4.8469 | 293.6447 | 5.1078 | -8.8422 | 282.4075 | -0.2965 | 0.21872 | 0.04217 | 0.26089 | 0
138 | 28 | 435.86 |  | 138 | 962.2065 | 495 | -4.9364 | 293.5692 | 5.2118 | -8.7382 | 282.1032 | -0.30224 | 0.23083 | 0.044521 | 0.27535 | 0
139 | 27.7 | 435.86 |  | 139 | 962.2065 | 510 | -5.0234 | 293.4922 | 5.3135 | -8.5865 | 281.8025 | -0.30782 | 0.24315 | 0.046915 | 0.29006 | 0
140 | 27.4 | 435.86 |  | 140 | 962.2065 | 525 | -5.1072 | 293.4139 | 5.4122 | -8.4878 | 281.507 | -0.31325 | 0.25568 | 0.049352 | 0.30503 | 0
141 | 27.1 | 435.86 |  | 141 | 962.2065 | 540 | -5.1886 | 293.3343 | 5.5088 | -8.3912 | 281.2149 | -0.31847 | 0.2684 | 0.051829 | 0.32023 | 0
142 | 26.8 | 435.86 |  | 142 | 962.2065 | 555 | -5.2676 | 293.2534 | 5.6033 | -8.2467 | 280.9261 | -0.32354 | 0.28132 | 0.054346 | 0.33566 | 0
143 | 26.5 | 435.86 |  | 143 | 962.2065 | 570 | -5.3437 | 293.1713 | 5.695 | -8.155 | 280.6423 | -0.32847 | 0.29442 | 0.056901 | 0.35132 | 0
144 | 26.2 | 435.86 |  | 144 | 962.2065 | 585 | -5.4175 | 293.088 | 5.7847 | -8.0653 | 280.3616 | -0.33321 | 0.3077 | 0.059493 | 0.3672 | 0
145 | 25.9 | 435.86 |  | 145 | 962.2065 | 600 | -5.4892 | 293.0035 | 5.8725 | -7.9775 | 280.0841 | -0.33782 | 0.32116 | 0.062121 | 0.38328 | 0
146 | 25.6 | 435.86 |  | 146 | 962.2065 | 615 | -5.5588 | 292.9179 | 5.9584 | -7.8416 | 279.8095 | -0.34229 | 0.33478 | 0.064783 | 0.39956 | 0
147 | 25.3 | 435.86 |  | 147 | 962.2065 | 630 | -5.6256 | 292.8313 | 6.0416 | -7.7584 | 279.5397 | -0.34663 | 0.34856 | 0.06748 | 0.41603 | 0
148 | 25 | 435.86 |  | 148 | 962.2065 | 645 | -5.6905 | 292.7436 | 6.1231 | -7.6769 | 279.2727 | -0.35079 | 0.36249 | 0.070208 | 0.4327 | 0
149 | 24.8333 | 435.86 |  | 149 | 962.2065 | 660 | -5.7534 | 292.6549 | 6.2029 | -7.5971 | 279.0085 | -0.35484 | 0.37657 | 0.072969 | 0.44954 | 0
150 | 24.6667 | 435.86 |  | 150 | 962.2065 | 675 | -5.8144 | 292.5652 | 6.281 | -7.469 | 278.747 | -0.35876 | 0.39079 | 0.075759 | 0.46655 | 0
151 | 24.5 | 435.86 |  | 151 | 962.2065 | 690 | -5.8729 | 292.4745 | 6.3566 | -7.3934 | 278.49 | -0.36257 | 0.40515 | 0.07858 | 0.48373 | 0
152 | 24.3333 | 435.86 |  | 152 | 962.2065 | 705 | -5.9296 | 292.383 | 6.4307 | -7.3193 | 278.2355 | -0.36621 | 0.41965 | 0.081428 | 0.50107 | 0
153 | 24.1667 | 435.86 |  | 153 | 962.2065 | 720 | -5.9846 | 292.2905 | 6.5031 | -7.1969 | 277.9836 | -0.36975 | 0.43426 | 0.084304 | 0.51857 | 0
154 | 24 | 435.86 |  | 154 | 962.2065 | 735 | -6.0371 | 292.1972 | 6.5733 | -7.1267 | 277.736 | -0.37318 | 0.449 | 0.087207 | 0.53621 | 0
155 | 23.8333 | 435.86 |  | 155 | 962.2065 | 750 | -6.088 | 292.1031 | 6.642 | -7.058 | 277.4907 | -0.37645 | 0.46386 | 0.090136 | 0.55399 | 0
156 | 23.6667 | 435.86 |  | 156 | 962.2065 | 765 | -6.1374 | 292.0082 | 6.7093 | -6.9907 | 277.2478 | -0.37963 | 0.47882 | 0.093089 | 0.57191 | 0
157 | 23.5 | 435.86 |  | 157 | 962.2065 | 780 | -6.1852 | 291.9126 | 6.7752 | -6.9248 | 277.0072 | -0.38271 | 0.4939 | 0.096066 | 0.58996 | 0
158 | 23.3333 | 435.86 |  | 158 | 962.2065 | 795 | -6.2315 | 291.8161 | 6.8397 | -6.8103 | 276.7689 | -0.38569 | 0.50907 | 0.099066 | 0.60814 | 0
159 | 23.1667 | 435.86 |  | 159 | 962.2065 | 810 | -6.2756 | 291.719 | 6.902 | -6.748 | 276.5345 | -0.38858 | 0.52434 | 0.10209 | 0.62643 | 0
160 | 23 | 435.86 |  | 160 | 962.2065 | 825 | -6.3183 | 291.6212 | 6.9631 | -6.6869 | 276.3023 | -0.39133 | 0.53971 | 0.10513 | 0.64484 | 0
161 | 22.8333 | 435.86 |  | 161 | 962.2065 | 840 | -6.3596 | 291.5227 | 7.023 | -6.627 | 276.0722 | -0.39399 | 0.55516 | 0.1082 | 0.66336 | 0
162 | 22.6667 | 435.86 |  | 162 | 962.2065 | 855 | -6.3996 | 291.4235 | 7.0816 | -6.5184 | 275.8441 | -0.39656 | 0.5707 | 0.11128 | 0.68198 | 0
163 | 22.5 | 435.86 |  | 163 | 962.2065 | 870 | -6.4375 | 291.3238 | 7.1382 | -6.4618 | 275.6198 | -0.39906 | 0.58632 | 0.11439 | 0.70071 | 0
164 | 22.45 | 435.86 |  | 164 | 962.2065 | 885 | -6.4741 | 291.2234 | 7.1936 | -6.4064 | 275.3974 | -0.40142 | 0.60202 | 0.11751 | 0.71953 | 0
165 | 22.4 | 435.86 |  | 165 | 962.2065 | 900 | -6.5095 | 291.1225 | 7.248 | -6.352 | 275.1769 | -0.4037 | 0.61779 | 0.12065 | 0.73844 | 0
166 | 22.35 | 435.86 |  | 166 | 962.2065 | 915 | -6.5438 | 291.021 | 7.3012 | -6.2488 | 274.9583 | -0.40591 | 0.63363 | 0.12381 | 0.75744 | 0
167 | 22.3 | 435.86 |  | 167 | 962.2065 | 930 | -6.5761 | 290.919 | 7.3526 | -6.1974 | 274.7433 | -0.40805 | 0.64954 | 0.12698 | 0.77652 | 0
168 | 22.25 | 435.86 |  | 168 | 962.2065 | 945 | -6.6073 | 290.8165 | 7.403 | -6.147 | 274.53 | -0.41006 | 0.66551 | 0.13017 | 0.79568 | 0
169 | 22.2 | 435.86 |  | 169 | 962.2065 | 960 | -6.6374 | 290.7135 | 7.4523 | -6.0977 | 274.3184 | -0.41201 | 0.68154 | 0.13338 | 0.81491 | 0
170 | 22.15 | 435.86 |  | 170 | 962.2065 | 975 | -6.6664 | 290.61 | 7.5006 | -6.0494 | 274.1086 | -0.41389 | 0.69762 | 0.13659 | 0.83421 | 0
171 | 22.1 | 435.86 |  | 171 | 962.2065 | 990 | -6.6945 | 290.5061 | 7.548 | -5.952 | 273.9004 | -0.4157 | 0.71376 | 0.13983 | 0.85358 | 0
172 | 22.05 | 435.86 |  | 172 | 962.2065 | 1005 | -6.7207 | 290.4017 | 7.5937 | -5.9063 | 273.6955 | -0.41744 | 0.72994 | 0.14308 | 0.87302 | 0
173 | 22 | 435.86 |  | 173 | 962.2065 | 1020 | -6.746 | 290.2969 | 7.6385 | -5.8615 | 273.4923 | -0.41908 | 0.74618 | 0.14634 | 0.89251 | 0
174 | 21.875 | 435.86 |  | 174 | 962.2065 | 1035 | -6.7703 | 290.1918 | 7.6824 | -5.8176 | 273.2906 | -0.42066 | 0.76245 | 0.14961 | 0.91206 | 0
175 | 21.75 | 435.86 |  | 175 | 962.2065 | 1050 | -6.7937 | 290.0862 | 7.7254 | -5.7746 | 273.0903 | -0.42217 | 0.77877 | 0.15289 | 0.93166 | 0
176 | 21.625 | 435.86 |  | 176 | 962.2065 | 1065 | -6.8163 | 289.9803 | 7.7676 | -5.7199 | 272.8916 | -0.42364 | 0.79512 | 0.15619 | 0.95131 | 0
177 | 21.5 | 435.86 |  | 177 | 962.2065 | 1080 | -6.8378 | 289.8741 | 7.8088 | -5.6787 | 272.6948 | -0.42504 | 0.81152 | 0.15949 | 0.97101 | 0
178 | 21.375 | 435.86 |  | 178 | 962.2065 | 1095 | -6.8584 | 289.7675 | 7.8491 | -5.6384 | 272.4993 | -0.42638 | 0.82794 | 0.16281 | 0.99075 | 0
179 | 21.25 | 435.86 |  | 179 | 962.2065 | 1110 | -6.8782 | 289.6605 | 7.8888 | -5.5987 | 272.3053 | -0.42767 | 0.8444 | 0.16614 | 1.0105 | 0
180 | 21.125 | 435.86 |  | 180 | 962.2065 | 1125 | -6.8972 | 289.5533 | 7.9276 | -5.5599 | 272.1126 | -0.4289 | 0.86088 | 0.16947 | 1.0304 | 0
181 | 21 | 435.86 |  | 181 | 962.2065 | 1140 | -6.9155 | 289.4458 | 7.9657 | -5.5093 | 271.9213 | -0.43009 | 0.8774 | 0.17282 | 1.0502 | 0
182 | 20.875 | 435.86 |  | 182 | 962.2065 | 1155 | -6.9328 | 289.338 | 8.0029 | -5.4721 | 271.7317 | -0.43123 | 0.89394 | 0.17617 | 1.0701 | 0
183 | 20.75 | 435.86 |  | 183 | 962.2065 | 1170 | -6.9493 | 289.2299 | 8.0393 | -5.4357 | 271.5434 | -0.4323 | 0.9105 | 0.17954 | 1.09 | 0
184 | 20.625 | 435.86 |  | 184 | 962.2065 | 1185 | -6.9652 | 289.1216 | 8.0751 | -5.3999 | 271.3563 | -0.43334 | 0.92708 | 0.18291 | 1.11 | 0
185 | 20.5 | 435.86 |  | 185 | 962.2065 | 1200 | -6.9803 | 289.013 | 8.1103 | -5.3647 | 271.1704 | -0.43432 | 0.94368 | 0.18628 | 1.13 | 0
186 | 20.4375 | 435.86 |  | 186 | 962.2065 | 1215 | -6.9947 | 288.9042 | 8.1447 | -5.3178 | 270.9858 | -0.43527 | 0.9603 | 0.18967 | 1.15 | 0
187 | 20.375 | 435.86 |  | 187 | 962.2065 | 1230 | -7.0083 | 288.7951 | 8.1783 | -5.2842 | 270.8028 | -0.43617 | 0.97693 | 0.19306 | 1.17 | 0
188 | 20.3125 | 435.86 |  | 188 | 962.2065 | 1245 | -7.0213 | 288.6859 | 8.2113 | -5.2512 | 270.621 | -0.43702 | 0.99358 | 0.19646 | 1.19 | 0
189 | 20.25 | 435.86 |  | 189 | 962.2065 | 1260 | -7.0336 | 288.5764 | 8.2437 | -5.2188 | 270.4402 | -0.43782 | 1.0102 | 0.19987 | 1.2101 | 0
190 | 20.1875 | 435.86 |  | 190 | 962.2065 | 1275 | -7.0453 | 288.4668 | 8.2755 | -5.187 | 270.2606 | -0.43859 | 1.0269 | 0.20328 | 1.2302 | 0
191 | 20.125 | 435.86 |  | 191 | 962.2065 | 1290 | -7.0565 | 288.357 | 8.3067 | -5.1558 | 270.0821 | -0.43932 | 1.0436 | 0.2067 | 1.2503 | 0
192 | 20.0625 | 435.86 |  | 192 | 962.2065 | 1305 | -7.067 | 288.247 | 8.3374 | -5.1126 | 269.9047 | -0.44002 | 1.0603 | 0.21012 | 1.2704 | 0
193 | 20 | 435.86 |  | 193 | 962.2065 | 1320 | -7.0768 | 288.1368 | 8.3673 | -5.0827 | 269.7287 | -0.44067 | 1.077 | 0.21355 | 1.2905 | 0
194 | 19.8387 | 435.86 |  | 194 | 962.2065 | 1335 | -7.086 | 288.0265 | 8.3967 | -5.0533 | 269.5538 | -0.44128 | 1.0937 | 0.21698 | 1.3106 | 0
195 | 19.6774 | 435.86 |  | 195 | 962.2065 | 1350 | -7.0947 | 287.916 | 8.4255 | -5.0245 | 269.3799 | -0.44186 | 1.1104 | 0.22042 | 1.3308 | 0
196 | 19.5161 | 435.86 |  | 196 | 962.2065 | 1365 | -7.1029 | 287.8054 | 8.4538 | -4.9962 | 269.207 | -0.4424 | 1.1271 | 0.22386 | 1.3509 | 0
197 | 19.3548 | 435.86 |  | 197 | 962.2065 | 1380 | -7.1106 | 287.6947 | 8.4816 | -4.9684 | 269.0351 | -0.44291 | 1.1438 | 0.2273 | 1.3711 | 0
198 | 19.1935 | 435.86 |  | 198 | 962.2065 | 1395 | -7.1178 | 287.5838 | 8.509 | -4.9285 | 268.8641 | -0.44339 | 1.1605 | 0.23075 | 1.3912 | 0
199 | 19.0323 | 435.86 |  | 199 | 962.2065 | 1410 | -7.1243 | 287.4729 | 8.5356 | -4.9019 | 268.6945 | -0.44384 | 1.1772 | 0.23421 | 1.4114 | 0
200 | 18.871 | 435.86 |  | 200 | 962.2065 | 1425 | -7.1303 | 287.3618 | 8.5618 | -4.8757 | 268.5258 | -0.44425 | 1.1938 | 0.23766 | 1.4315 | 0
201 | 18.7097 | 435.86 |  | 201 | 962.2065 | 1440 | -7.1359 | 287.2506 | 8.5876 | -4.8499 | 268.358 | -0.44462 | 1.2105 | 0.24112 | 1.4517 | 0
202 | 18.5484 | 435.86 |  | 202 | 962.2065 | 1455 | -7.1411 | 287.1394 | 8.6129 | -4.8246 | 268.1911 | -0.44497 | 1.2272 | 0.24458 | 1.4718 | 0
203 | 18.3871 | 435.86 |  | 203 | 962.2065 | 1470 | -7.1458 | 287.0281 | 8.6378 | -4.7997 | 268.025 | -0.44529 | 1.2439 | 0.24805 | 1.4919 | 0
204 | 18.2258 | 435.86 |  | 204 | 962.2065 | 1485 | -7.1501 | 286.9167 | 8.6622 | -4.7628 | 267.8598 | -0.44559 | 1.2605 | 0.25151 | 1.5121 | 0
205 | 18.0645 | 435.86 |  | 205 | 962.2065 | 1500 | -7.1539 | 286.8052 | 8.686 | -4.739 | 267.6959 | -0.44586 | 1.2772 | 0.25498 | 1.5322 | 0
206 | 17.9032 | 435.86 |  | 206 | 962.2065 | 1515 | -7.1572 | 286.6937 | 8.7095 | -4.7155 | 267.5328 | -0.44609 | 1.2938 | 0.25845 | 1.5523 | 0
207 | 17.7419 | 435.86 |  | 207 | 962.2065 | 1530 | -7.1601 | 286.5821 | 8.7325 | -4.6925 | 267.3706 | -0.4463 | 1.3105 | 0.26192 | 1.5724 | 0
208 | 17.5806 | 435.86 |  | 208 | 962.2065 | 1545 | -7.1627 | 286.4705 | 8.7552 | -4.6698 | 267.2091 | -0.44648 | 1.3271 | 0.26539 | 1.5925 | 0
209 | 17.4194 | 435.86 |  | 209 | 962.2065 | 1560 | -7.1649 | 286.3588 | 8.7775 | -4.6475 | 267.0484 | -0.44664 | 1.3437 | 0.26887 | 1.6126 | 0
210 | 17.2581 | 435.86 |  | 210 | 962.2065 | 1575 | -7.1668 | 286.2471 | 8.7994 | -4.6131 | 266.8884 | -0.44678 | 1.3603 | 0.27234 | 1.6326 | 0
211 | 17.0968 | 435.86 |  | 211 | 962.2065 | 1590 | -7.1681 | 286.1354 | 8.8208 | -4.5917 | 266.7297 | -0.4469 | 1.3769 | 0.27582 | 1.6527 | 0
212 | 16.9355 | 435.86 |  | 212 | 962.2065 | 1605 | -7.1691 | 286.0237 | 8.8418 | -4.5707 | 266.5716 | -0.44698 | 1.3934 | 0.2793 | 1.6727 | 0
213 | 16.7742 | 435.86 |  | 213 | 962.2065 | 1620 | -7.1698 | 285.9119 | 8.8625 | -4.55 | 266.4143 | -0.44704 | 1.41 | 0.28277 | 1.6928 | 0
214 | 16.6129 | 435.86 |  | 214 | 962.2065 | 1635 | -7.1701 | 285.8001 | 8.8829 | -4.5296 | 266.2578 | -0.44708 | 1.4265 | 0.28625 | 1.7128 | 0
215 | 16.4516 | 435.86 |  | 215 | 962.2065 | 1650 | -7.1702 | 285.6884 | 8.903 | -4.5095 | 266.1019 | -0.44711 | 1.443 | 0.28973 | 1.7327 | 0
216 | 16.2903 | 435.86 |  | 216 | 962.2065 | 1665 | -7.17 | 285.5766 | 8.9227 | -4.4773 | 265.9467 | -0.44711 | 1.4595 | 0.29321 | 1.7527 | 0
217 | 16.129 | 435.86 |  | 217 | 962.2065 | 1680 | -7.1693 | 285.4648 | 8.9419 | -4.4581 | 265.7926 | -0.4471 | 1.476 | 0.29669 | 1.7727 | 0
218 | 15.9677 | 435.86 |  | 218 | 962.2065 | 1695 | -7.1683 | 285.353 | 8.9609 | -4.4391 | 265.6392 | -0.44705 | 1.4924 | 0.30016 | 1.7926 | 0
219 | 15.8065 | 435.86 |  | 219 | 962.2065 | 1710 | -7.167 | 285.2413 | 8.9795 | -4.4205 | 265.4864 | -0.44699 | 1.5088 | 0.30364 | 1.8125 | 0
220 | 15.6452 | 435.86 |  | 220 | 962.2065 | 1725 | -7.1655 | 285.1296 | 8.9979 | -4.4021 | 265.3343 | -0.44691 | 1.5252 | 0.30712 | 1.8324 | 0
221 | 15.4839 | 435.86 |  | 221 | 962.2065 | 1740 | -7.1637 | 285.0179 | 9.016 | -4.384 | 265.1828 | -0.44682 | 1.5416 | 0.31059 | 1.8522 | 0
222 | 15.3226 | 435.86 |  | 222 | 962.2065 | 1755 | -7.1617 | 284.9062 | 9.0338 | -4.3662 | 265.0319 | -0.44671 | 1.558 | 0.31407 | 1.8721 | 0
223 | 15.1613 | 435.86 |  | 223 | 962.2065 | 1770 | -7.1595 | 284.7945 | 9.0513 | -4.3362 | 264.8816 | -0.44658 | 1.5743 | 0.31754 | 1.8919 | 0
224 | 15 | 435.86 |  | 224 | 962.2065 | 1785 | -7.1568 | 284.6829 | 9.0684 | -4.3191 | 264.7324 | -0.44644 | 1.5906 | 0.32101 | 1.9116 | 0
225 | 14.7778 | 435.86 |  | 225 | 962.2065 | 1800 | -7.1539 | 284.5714 | 9.0853 | -4.3022 | 264.5838 | -0.44627 | 1.6069 | 0.32449 | 1.9314 | 0
226 | 14.5556 | 435.86 |  | 226 | 962.2065 | 1815 | -7.1508 | 284.4598 | 9.1019 | -4.2856 | 264.4357 | -0.44609 | 1.6232 | 0.32796 | 1.9511 | 0
227 | 14.3333 | 435.86 |  | 227 | 962.2065 | 1830 | -7.1474 | 284.3484 | 9.1183 | -4.2692 | 264.2882 | -0.4459 | 1.6394 | 0.33142 | 1.9708 | 0
228 | 14.1111 | 435.86 |  | 228 | 962.2065 | 1845 | -7.1439 | 284.2369 | 9.1344 | -4.2531 | 264.1413 | -0.44569 | 1.6556 | 0.33489 | 1.9905 | 0
229 | 13.8889 | 435.86 |  | 229 | 962.2065 | 1860 | -7.1402 | 284.1256 | 9.1503 | -4.2247 | 263.9949 | -0.44547 | 1.6718 | 0.33836 | 2.0101 | 0
230 | 13.6667 | 435.86 |  | 230 | 962.2065 | 1875 | -7.136 | 284.0143 | 9.1658 | -4.2092 | 263.8495 | -0.44524 | 1.6879 | 0.34182 | 2.0298 | 0
231 | 13.4444 | 435.86 |  | 231 | 962.2065 | 1890 | -7.1317 | 283.903 | 9.1811 | -4.1939 | 263.7047 | -0.44498 | 1.7041 | 0.34528 | 2.0493 | 0
232 | 13.2222 | 435.86 |  | 232 | 962.2065 | 1905 | -7.1273 | 283.7918 | 9.1961 | -4.1789 | 263.5603 | -0.44471 | 1.7202 | 0.34874 | 2.0689 | 0
233 | 13 | 435.86 |  | 233 | 962.2065 | 1920 | -7.1226 | 283.6807 | 9.211 | -4.164 | 263.4165 | -0.44443 | 1.7362 | 0.3522 | 2.0884 | 0
234 | 13.0125 | 435.86 |  | 234 | 962.2065 | 1935 | -7.1178 | 283.5697 | 9.2257 | -4.1493 | 263.2732 | -0.44414 | 1.7523 | 0.35565 | 2.1079 | 0
235 | 13.025 | 435.86 |  | 235 | 962.2065 | 1950 | -7.1128 | 283.4587 | 9.2401 | -4.1349 | 263.1304 | -0.44384 | 1.7683 | 0.3591 | 2.1274 | 0
236 | 13.0375 | 435.86 |  | 236 | 962.2065 | 1965 | -7.1076 | 283.3479 | 9.2544 | -4.1081 | 262.9881 | -0.44353 | 1.7842 | 0.36256 | 2.1468 | 0
237 | 13.05 | 435.86 |  | 237 | 962.2065 | 1980 | -7.1021 | 283.2371 | 9.2683 | -4.0942 | 262.8467 | -0.44321 | 1.8002 | 0.366 | 2.1662 | 0
238 | 13.0625 | 435.86 |  | 238 | 962.2065 | 1995 | -7.0965 | 283.1263 | 9.282 | -4.0805 | 262.7058 | -0.44287 | 1.8161 | 0.36945 | 2.1855 | 0
239 | 13.075 | 435.86 |  | 239 | 962.2065 | 2010 | -7.0907 | 283.0157 | 9.2956 | -4.0669 | 262.5654 | -0.44251 | 1.832 | 0.37289 | 2.2049 | 0
240 | 13.0875 | 435.86 |  | 240 | 962.2065 | 2025 | -7.0848 | 282.9052 | 9.309 | -4.0535 | 262.4254 | -0.44215 | 1.8478 | 0.37633 | 2.2242 | 0
241 | 13.1 | 435.86 |  | 241 | 962.2065 | 2040 | -7.0788 | 282.7947 | 9.3222 | -4.0403 | 262.2859 | -0.44178 | 1.8637 | 0.37977 | 2.2434 | 0
242 | 13.1125 | 435.86 |  | 242 | 962.2065 | 2055 | -7.0726 | 282.6844 | 9.3352 | -4.0273 | 262.1469 | -0.44141 | 1.8794 | 0.3832 | 2.2626 | 0
243 | 13.125 | 435.86 |  | 243 | 962.2065 | 2070 | -7.0663 | 282.5741 | 9.3481 | -4.0144 | 262.0083 | -0.44102 | 1.8952 | 0.38663 | 2.2818 | 0
244 | 13.1375 | 435.86 |  | 244 | 962.2065 | 2085 | -7.0598 | 282.464 | 9.3608 | -3.9892 | 261.8701 | -0.44063 | 1.9109 | 0.39006 | 2.301 | 0
245 | 13.15 | 435.86 |  | 245 | 962.2065 | 2100 | -7.0531 | 282.3539 | 9.3732 | -3.9768 | 261.7329 | -0.44023 | 1.9266 | 0.39348 | 2.3201 | 0
246 | 13.1625 | 435.86 |  | 246 | 962.2065 | 2115 | -7.0463 | 282.244 | 9.3854 | -3.9646 | 261.596 | -0.43981 | 1.9423 | 0.3969 | 2.3392 | 0
247 | 13.175 | 435.86 |  | 247 | 962.2065 | 2130 | -7.0393 | 282.1341 | 9.3975 | -3.9525 | 261.4596 | -0.43938 | 1.9579 | 0.40032 | 2.3582 | 0
248 | 13.1875 | 435.86 |  | 248 | 962.2065 | 2145 | -7.0323 | 282.0244 | 9.4095 | -3.9405 | 261.3235 | -0.43895 | 1.9735 | 0.40374 | 2.3772 | 0
249 | 13.2 | 435.86 |  | 249 | 962.2065 | 2160 | -7.0251 | 281.9147 | 9.4213 | -3.9287 | 261.1879 | -0.43851 | 1.989 | 0.40715 | 2.3962 | 0
250 | 13.2125 | 435.86 |  | 250 | 962.2065 | 2175 | -7.0178 | 281.8052 | 9.433 | -3.917 | 261.0527 | -0.43806 | 2.0046 | 0.41055 | 2.4151 | 0
251 | 13.225 | 435.86 |  | 251 | 962.2065 | 2190 | -7.0105 | 281.6958 | 9.4445 | -3.893 | 260.9179 | -0.43761 | 2.0201 | 0.41396 | 2.434 | 0
252 | 13.2375 | 435.86 |  | 252 | 962.2065 | 2205 | -7.0029 | 281.5865 | 9.4557 | -3.8818 | 260.7839 | -0.43715 | 2.0355 | 0.41736 | 2.4529 | 0
253 | 13.25 | 435.86 |  | 253 | 962.2065 | 2220 | -6.9951 | 281.4774 | 9.4668 | -3.8707 | 260.6504 | -0.43668 | 2.0509 | 0.42076 | 2.4717 | 0
254 | 13.2625 | 435.86 |  | 254 | 962.2065 | 2235 | -6.9874 | 281.3683 | 9.4778 | -3.8597 | 260.5171 | -0.43619 | 2.0663 | 0.42415 | 2.4905 | 0
255 | 13.275 | 435.86 |  | 255 | 962.2065 | 2250 | -6.9795 | 281.2594 | 9.4887 | -3.8488 | 260.3843 | -0.43571 | 2.0817 | 0.42754 | 2.5092 | 0
256 | 13.2875 | 435.86 |  | 256 | 962.2065 | 2265 | -6.9715 | 281.1506 | 9.4994 | -3.8381 | 260.2519 | -0.43522 | 2.097 | 0.43092 | 2.5279 | 0
257 | 13.3 | 435.86 |  | 257 | 962.2065 | 2280 | -6.9635 | 281.0419 | 9.5101 | -3.8274 | 260.1198 | -0.43472 | 2.1123 | 0.4343 | 2.5466 | 0
258 | 13.3125 | 435.86 |  | 258 | 962.2065 | 2295 | -6.9554 | 280.9334 | 9.5206 | -3.8044 | 259.9881 | -0.43422 | 2.1275 | 0.43768 | 2.5652 | 0
259 | 13.325 | 435.86 |  | 259 | 962.2065 | 2310 | -6.9471 | 280.8249 | 9.5308 | -3.7942 | 259.8571 | -0.43372 | 2.1427 | 0.44106 | 2.5838 | 0
260 | 13.3375 | 435.86 |  | 260 | 962.2065 | 2325 | -6.9386 | 280.7166 | 9.5409 | -3.7841 | 259.7265 | -0.4332 | 2.1579 | 0.44443 | 2.6023 | 0
261 | 13.35 | 435.86 |  | 261 | 962.2065 | 2340 | -6.9302 | 280.6085 | 9.551 | -3.774 | 259.5963 | -0.43267 | 2.173 | 0.44779 | 2.6208 | 0
262 | 13.3625 | 435.86 |  | 262 | 962.2065 | 2355 | -6.9216 | 280.5004 | 9.5609 | -3.7641 | 259.4664 | -0.43214 | 2.1881 | 0.45115 | 2.6393 | 0
263 | 13.375 | 435.86 |  | 263 | 962.2065 | 2370 | -6.913 | 280.3925 | 9.5707 | -3.7543 | 259.3369 | -0.43161 | 2.2032 | 0.45451 | 2.6577 | 0
264 | 13.3875 | 435.86 |  | 264 | 962.2065 | 2385 | -6.9044 | 280.2848 | 9.5805 | -3.7445 | 259.2077 | -0.43107 | 2.2182 | 0.45786 | 2.6761 | 0
265 | 13.4 | 435.86 |  | 265 | 962.2065 | 2400 | -6.8957 | 280.1771 | 9.5901 | -3.7349 | 259.0788 | -0.43054 | 2.2332 | 0.46121 | 2.6944 | 0
266 | 13.4125 | 435.86 |  | 266 | 962.2065 | 2415 | -6.887 | 280.0696 | 9.5997 | -3.7128 | 258.9503 | -0.42999 | 2.2482 | 0.46456 | 2.7127 | 0
267 | 13.425 | 435.86 |  | 267 | 962.2065 | 2430 | -6.878 | 279.9623 | 9.609 | -3.7035 | 258.8225 | -0.42945 | 2.2631 | 0.4679 | 2.731 | 0
268 | 13.4375 | 435.86 |  | 268 | 962.2065 | 2445 | -6.869 | 279.855 | 9.6182 | -3.6943 | 258.6951 | -0.42889 | 2.278 | 0.47123 | 2.7492 | 0
269 | 13.45 | 435.86 |  | 269 | 962.2065 | 2460 | -6.8599 | 279.748 | 9.6273 | -3.6852 | 258.5679 | -0.42833 | 2.2928 | 0.47457 | 2.7674 | 0
270 | 13.4625 | 435.86 |  | 270 | 962.2065 | 2475 | -6.8508 | 279.641 | 9.6363 | -3.6762 | 258.4411 | -0.42776 | 2.3076 | 0.47789 | 2.7855 | 0
271 | 13.475 | 435.86 |  | 271 | 962.2065 | 2490 | -6.8417 | 279.5342 | 9.6453 | -3.6672 | 258.3146 | -0.42719 | 2.3224 | 0.48122 | 2.8036 | 0
272 | 13.4875 | 435.86 |  | 272 | 962.2065 | 2505 | -6.8325 | 279.4276 | 9.6542 | -3.6583 | 258.1884 | -0.42662 | 2.3371 | 0.48453 | 2.8217 | 0
273 | 13.5 | 435.86 |  | 273 | 962.2065 | 2520 | -6.8233 | 279.3211 | 9.663 | -3.6495 | 258.0625 | -0.42605 | 2.3518 | 0.48785 | 2.8397 | 0
274 | 13.55 | 435.86 |  | 274 | 962.2065 | 2535 | -6.8141 | 279.2147 | 9.6717 | -3.6283 | 257.9369 | -0.42548 | 2.3665 | 0.49116 | 2.8577 | 0
275 | 13.6 | 435.86 |  | 275 | 962.2065 | 2550 | -6.8046 | 279.1085 | 9.6802 | -3.6198 | 257.812 | -0.4249 | 2.3811 | 0.49446 | 2.8756 | 0
276 | 13.65 | 435.86 |  | 276 | 962.2065 | 2565 | -6.7951 | 279.0024 | 9.6886 | -3.6114 | 257.6874 | -0.42431 | 2.3957 | 0.49776 | 2.8935 | 0
277 | 13.7 | 435.86 |  | 277 | 962.2065 | 2580 | -6.7856 | 278.8964 | 9.697 | -3.603 | 257.5632 | -0.42372 | 2.4103 | 0.50106 | 2.9113 | 0
278 | 13.75 | 435.86 |  | 278 | 962.2065 | 2595 | -6.7761 | 278.7907 | 9.7052 | -3.5948 | 257.4392 | -0.42313 | 2.4248 | 0.50435 | 2.9291 | 0
279 | 13.8 | 435.86 |  | 279 | 962.2065 | 2610 | -6.7665 | 278.685 | 9.7134 | -3.5866 | 257.3154 | -0.42253 | 2.4393 | 0.50764 | 2.9469 | 0
280 | 13.85 | 435.86 |  | 280 | 962.2065 | 2625 | -6.757 | 278.5795 | 9.7216 | -3.5784 | 257.192 | -0.42194 | 2.4537 | 0.51092 | 2.9646 | 0
281 | 13.9 | 435.86 |  | 281 | 962.2065 | 2640 | -6.7474 | 278.4742 | 9.7297 | -3.5703 | 257.0689 | -0.42134 | 2.4681 | 0.5142 | 2.9823 | 0
282 | 13.95 | 435.86 |  | 282 | 962.2065 | 2655 | -6.7378 | 278.369 | 9.7377 | -3.5498 | 256.946 | -0.42074 | 2.4825 | 0.51747 | 3 | 0
283 | 14 | 435.86 |  | 283 | 962.2065 | 2670 | -6.728 | 278.264 | 9.7455 | -3.542 | 256.8238 | -0.42015 | 2.4968 | 0.52074 | 3.0176 | 0
284 | 14.05 | 435.86 |  | 284 | 962.2065 | 2685 | -6.7181 | 278.1591 | 9.7533 | -3.5342 | 256.7019 | -0.41953 | 2.5111 | 0.524 | 3.0351 | 0
285 | 14.1 | 435.86 |  | 285 | 962.2065 | 2700 | -6.7083 | 278.0544 | 9.7609 | -3.5266 | 256.5803 | -0.41892 | 2.5254 | 0.52726 | 3.0526 | 0
286 | 14.15 | 435.86 |  | 286 | 962.2065 | 2715 | -6.6984 | 277.9498 | 9.7686 | -3.5189 | 256.4589 | -0.41831 | 2.5396 | 0.53052 | 3.0701 | 0
287 | 14.2 | 435.86 |  | 287 | 962.2065 | 2730 | -6.6886 | 277.8454 | 9.7762 | -3.5113 | 256.3378 | -0.41769 | 2.5538 | 0.53376 | 3.0876 | 0
288 | 14.25 | 435.86 |  | 288 | 962.2065 | 2745 | -6.6787 | 277.7411 | 9.7837 | -3.5038 | 256.217 | -0.41708 | 2.5679 | 0.53701 | 3.105 | 0
289 | 14.3 | 435.86 |  | 289 | 962.2065 | 2760 | -6.6689 | 277.637 | 9.7912 | -3.4963 | 256.0964 | -0.41646 | 2.5821 | 0.54025 | 3.1223 | 0
290 | 14.35 | 435.86 |  | 290 | 962.2065 | 2775 | -6.659 | 277.533 | 9.7986 | -3.4764 | 255.9761 | -0.41585 | 2.5961 | 0.54348 | 3.1396 | 0
291 | 14.4 | 435.86 |  | 291 | 962.2065 | 2790 | -6.6489 | 277.4292 | 9.8058 | -3.4692 | 255.8564 | -0.41523 | 2.6102 | 0.54671 | 3.1569 | 0
292 | 14.45 | 435.86 |  | 292 | 962.2065 | 2805 | -6.6389 | 277.3256 | 9.813 | -3.462 | 255.7371 | -0.4146 | 2.6242 | 0.54994 | 3.1741 | 0
293 | 14.5 | 435.86 |  | 293 | 962.2065 | 2820 | -6.6288 | 277.2221 | 9.8201 | -3.4549 | 255.6179 | -0.41398 | 2.6381 | 0.55316 | 3.1913 | 0
294 | 14.5 | 435.86 |  | 294 | 962.2065 | 2835 | -6.6187 | 277.1187 | 9.8271 | -3.4479 | 255.499 | -0.41335 | 2.6521 | 0.55637 | 3.2084 | 0
295 | 14.5 | 435.86 |  | 295 | 962.2065 | 2850 | -6.6086 | 277.0156 | 9.8342 | -3.4408 | 255.3804 | -0.41272 | 2.666 | 0.55958 | 3.2255 | 0
296 | 14.5 | 435.86 |  | 296 | 962.2065 | 2865 | -6.5986 | 276.9125 | 9.8412 | -3.4338 | 255.2619 | -0.41209 | 2.6798 | 0.56279 | 3.2426 | 0
297 | 14.5 | 435.86 |  | 297 | 962.2065 | 2880 | -6.5885 | 276.8097 | 9.8481 | -3.4269 | 255.1438 | -0.41147 | 2.6936 | 0.56599 | 3.2596 | 0
298 | 14.5 | 435.86 |  | 298 | 962.2065 | 2895 | -6.5785 | 276.707 | 9.8551 | -3.4199 | 255.0258 | -0.41084 | 2.7074 | 0.56919 | 3.2766 | 0
299 | 14.5 | 435.86 |  | 299 | 962.2065 | 2910 | -6.5684 | 276.6044 | 9.8619 | -3.4006 | 254.9081 | -0.41021 | 2.7212 | 0.57238 | 3.2935 | 0
300 | 14.5 | 435.86 |  | 300 | 962.2065 | 2925 | -6.5582 | 276.502 | 9.8686 | -3.3939 | 254.7911 | -0.40958 | 2.7349 | 0.57556 | 3.3104 | 0
[6 empty] | 2940 | -6.5479 | 276.3998 | 9.8752 | -3.3873 | 254.6743 | -0.40895 | 2.7485 | 0.57874 | 3.3273 | 0
[6 empty] | 2955 | -6.5377 | 276.2977 | 9.8818 | -3.3807 | 254.5577 | -0.40831 | 2.7622 | 0.58192 | 3.3441 | 0
[6 empty] | 2970 | -6.5275 | 276.1958 | 9.8884 | -3.3741 | 254.4414 | -0.40767 | 2.7758 | 0.58509 | 3.3609 | 0
[6 empty] | 2985 | -6.5173 | 276.094 | 9.8949 | -3.3676 | 254.3253 | -0.40703 | 2.7893 | 0.58826 | 3.3776 | 0
[6 empty] | 3000 | -6.5071 | 275.9924 | 9.9014 | -3.3611 | 254.2094 | -0.4064 | 2.8029 | 0.59142 | 3.3943 | 0
[6 empty] | 3015 | -6.4969 | 275.891 | 9.9079 | -3.3546 | 254.0937 | -0.40576 | 2.8164 | 0.59458 | 3.4109 | 0
[6 empty] | 3030 | -6.4868 | 275.7897 | 9.9143 | -3.3357 | 253.9783 | -0.40513 | 2.8298 | 0.59773 | 3.4275 | 0
[6 empty] | 3045 | -6.4764 | 275.6886 | 9.9205 | -3.3295 | 253.8635 | -0.40449 | 2.8432 | 0.60087 | 3.4441 | 0
[6 empty] | 3060 | -6.4661 | 275.5876 | 9.9267 | -3.3233 | 253.7489 | -0.40385 | 2.8566 | 0.60401 | 3.4606 | 0
[6 empty] | 3075 | -6.4558 | 275.4868 | 9.9329 | -3.3171 | 253.6345 | -0.4032 | 2.87 | 0.60715 | 3.4771 | 0
[6 empty] | 3090 | -6.4455 | 275.3862 | 9.939 | -3.311 | 253.5203 | -0.40256 | 2.8833 | 0.61028 | 3.4936 | 0
[6 empty] | 3105 | -6.4352 | 275.2857 | 9.9451 | -3.3049 | 253.4064 | -0.40192 | 2.8965 | 0.61341 | 3.51 | 0
[6 empty] | 3120 | -6.4249 | 275.1854 | 9.9512 | -3.2988 | 253.2927 | -0.40128 | 2.9098 | 0.61653 | 3.5263 | 0
[6 empty] | 3135 | -6.4147 | 275.0852 | 9.9573 | -3.2927 | 253.1791 | -0.40064 | 2.923 | 0.61965 | 3.5426 | 0
[6 empty] | 3150 | -6.4044 | 274.9852 | 9.9634 | -3.2866 | 253.0658 | -0.4 | 2.9362 | 0.62276 | 3.5589 | 0
[6 empty] | 3165 | -6.3942 | 274.8854 | 9.9694 | -3.2681 | 252.9527 | -0.39936 | 2.9493 | 0.62586 | 3.5752 | 0
[6 empty] | 3180 | -6.3838 | 274.7857 | 9.9752 | -3.2623 | 252.8402 | -0.39872 | 2.9624 | 0.62897 | 3.5914 | 0
[6 empty] | 3195 | -6.3735 | 274.6862 | 9.981 | -3.2565 | 252.728 | -0.39808 | 2.9755 | 0.63206 | 3.6075 | 0
[6 empty] | 3210 | -6.3631 | 274.5868 | 9.9868 | -3.2507 | 252.6159 | -0.39743 | 2.9885 | 0.63515 | 3.6236 | 0
[6 empty] | 3225 | -6.3528 | 274.4876 | 9.9925 | -3.245 | 252.504 | -0.39678 | 3.0015 | 0.63824 | 3.6397 | 0
[6 empty] | 3240 | -6.3425 | 274.3886 | 9.9983 | -3.2392 | 252.3923 | -0.39614 | 3.0144 | 0.64132 | 3.6558 | 0
[6 empty] | 3255 | -6.3323 | 274.2897 | 10.004 | -3.2335 | 252.2809 | -0.3955 | 3.0274 | 0.6444 | 3.6718 | 0
[6 empty] | 3270 | -6.322 | 274.191 | 10.0097 | -3.2278 | 252.1696 | -0.39486 | 3.0402 | 0.64747 | 3.6877 | 0
[6 empty] | 3285 | -6.3118 | 274.0924 | 10.0154 | -3.2221 | 252.0585 | -0.39422 | 3.0531 | 0.65054 | 3.7036 | 0
[6 empty] | 3300 | -6.3016 | 273.994 | 10.0211 | -3.2039 | 251.9476 | -0.39358 | 3.0659 | 0.6536 | 3.7195 | 0
[6 empty] | 3315 | -6.2912 | 273.8958 | 10.0266 | -3.1984 | 251.8374 | -0.39295 | 3.0787 | 0.65665 | 3.7354 | 0
[6 empty] | 3330 | -6.2809 | 273.7977 | 10.032 | -3.193 | 251.7273 | -0.3923 | 3.0914 | 0.65971 | 3.7512 | 0
[6 empty] | 3345 | -6.2705 | 273.6998 | 10.0375 | -3.1875 | 251.6174 | -0.39165 | 3.1042 | 0.66275 | 3.7669 | 0
[6 empty] | 3360 | -6.2603 | 273.6021 | 10.0429 | -3.1821 | 251.5077 | -0.39101 | 3.1168 | 0.66579 | 3.7826 | 0
[6 empty] | 3375 | -6.25 | 273.5045 | 10.0483 | -3.1767 | 251.3982 | -0.39037 | 3.1295 | 0.66883 | 3.7983 | 0
[6 empty] | 3390 | -6.2398 | 273.407 | 10.0537 | -3.1713 | 251.2889 | -0.38973 | 3.1421 | 0.67186 | 3.814 | 0
Ambient | 295 | K |  |  |  | 3405 | -6.2295 | 273.3098 | 10.0591 | -3.1659 | 251.1797 | -0.38909 | 3.1547 | 0.67489 | 3.8296 | 0
window | 280 | K |  |  |  | 3420 | -6.2194 | 273.2127 | 10.0645 | -3.1605 | 251.0708 | -0.38845 | 3.1672 | 0.67791 | 3.8451 | 0
the radiation reduction | 18.84% |  |  |  |  | 3435 | -6.2092 | 273.1157 | 10.0699 | -3.1426 | 250.962 | -0.38782 | 3.1797 | 0.68093 | 3.8607 | 0
[6 empty] | 3450 | -6.1989 | 273.0189 | 10.075 | -3.1375 | 250.8539 | -0.38719 | 3.1922 | 0.68394 | 3.8761 | 0
[6 empty] | 3465 | -6.1886 | 272.9223 | 10.0802 | -3.1323 | 250.7459 | -0.38654 | 3.2046 | 0.68695 | 3.8916 | 0
[6 empty] | 3480 | -6.1783 | 272.8258 | 10.0853 | -3.1272 | 250.6381 | -0.3859 | 3.2171 | 0.68995 | 3.907 | 0
[6 empty] | 3495 | -6.1681 | 272.7295 | 10.0905 | -3.122 | 250.5305 | -0.38526 | 3.2294 | 0.69294 | 3.9224 | 0
[6 empty] | 3510 | -6.1579 | 272.6333 | 10.0956 | -3.1169 | 250.423 | -0.38462 | 3.2418 | 0.69594 | 3.9377 | 0
[6 empty] | 3525 | -6.1477 | 272.5373 | 10.1007 | -3.1118 | 250.3157 | -0.38399 | 3.2541 | 0.69892 | 3.953 | 0
[6 empty] | 3540 | -6.1376 | 272.4415 | 10.1058 | -3.1067 | 250.2087 | -0.38335 | 3.2663 | 0.70191 | 3.9683 | 0
[6 empty] | 3555 | -6.1275 | 272.3458 | 10.1109 | -3.1016 | 250.1017 | -0.38272 | 3.2786 | 0.70488 | 3.9835 | 0
[6 empty] | 3570 | -6.1174 | 272.2503 | 10.116 | -3.084 | 249.995 | -0.38209 | 3.2908 | 0.70785 | 3.9987 | 0
[6 empty] | 3585 | -6.1071 | 272.1549 | 10.1209 | -3.0791 | 249.8889 | -0.38146 | 3.303 | 0.71082 | 4.0138 | 0
[6 empty] | 3600 | -6.0969 | 272.0597 | 10.1258 | -3.0742 | 249.7829 | -0.38082 | 3.3151 | 0.71378 | 4.0289 | 0
[6 empty] | 3615 | -6.0867 | 271.9647 | 10.1307 | -3.0693 | 249.6771 | -0.38018 | 3.3272 | 0.71674 | 4.044 | 0
[6 empty] | 3630 | -6.0766 | 271.8698 | 10.1356 | -3.0644 | 249.5715 | -0.37955 | 3.3393 | 0.71969 | 4.059 | 0
[6 empty] | 3645 | -6.0665 | 271.7751 | 10.1405 | -3.0595 | 249.466 | -0.37892 | 3.3513 | 0.72264 | 4.074 | 0
[6 empty] | 3660 | -6.0564 | 271.6805 | 10.1453 | -3.0547 | 249.3607 | -0.37829 | 3.3634 | 0.72558 | 4.0889 | 0
[6 empty] | 3675 | -6.0464 | 271.5861 | 10.1502 | -3.0498 | 249.2556 | -0.37766 | 3.3753 | 0.72852 | 4.1039 | 0
[6 empty] | 3690 | -6.0363 | 271.4918 | 10.1551 | -3.0449 | 249.1506 | -0.37703 | 3.3873 | 0.73145 | 4.1187 | 0
[6 empty] | 3705 | -6.0264 | 271.3977 | 10.1599 | -3.0401 | 249.0459 | -0.37641 | 3.3992 | 0.73438 | 4.1336 | 0
[6 empty] | 3720 | -6.0164 | 271.3038 | 10.1648 | -3.0227 | 248.9412 | -0.37578 | 3.4111 | 0.73731 | 4.1484 | 0
[6 empty] | 3735 | -6.0063 | 271.21 | 10.1694 | -3.0181 | 248.8372 | -0.37516 | 3.4229 | 0.74022 | 4.1631 | 0
[6 empty] | 3750 | -5.9962 | 271.1163 | 10.1741 | -3.0134 | 248.7333 | -0.37453 | 3.4347 | 0.74314 | 4.1779 | 0
[6 empty] | 3765 | -5.9862 | 271.0229 | 10.1787 | -3.0088 | 248.6296 | -0.3739 | 3.4465 | 0.74605 | 4.1926 | 0
[6 empty] | 3780 | -5.9762 | 270.9295 | 10.1834 | -3.0041 | 248.5261 | -0.37328 | 3.4583 | 0.74895 | 4.2072 | 0
[6 empty] | 3795 | -5.9662 | 270.8364 | 10.188 | -2.9995 | 248.4227 | -0.37265 | 3.47 | 0.75185 | 4.2218 | 0
[6 empty] | 3810 | -5.9563 | 270.7434 | 10.1927 | -2.9948 | 248.3195 | -0.37203 | 3.4817 | 0.75474 | 4.2364 | 0
[6 empty] | 3825 | -5.9463 | 270.6505 | 10.1973 | -2.9902 | 248.2164 | -0.37141 | 3.4933 | 0.75763 | 4.251 | 0
[6 empty] | 3840 | -5.9365 | 270.5578 | 10.202 | -2.9855 | 248.1135 | -0.37079 | 3.505 | 0.76052 | 4.2655 | 0
[6 empty] | 3855 | -5.9266 | 270.4653 | 10.2066 | -2.9809 | 248.0108 | -0.37018 | 3.5166 | 0.7634 | 4.28 | 0
[6 empty] | 3870 | -5.9168 | 270.3729 | 10.2112 | -2.9638 | 247.9082 | -0.36957 | 3.5281 | 0.76627 | 4.2944 | 0
[6 empty] | 3885 | -5.9069 | 270.2807 | 10.2157 | -2.9593 | 247.8062 | -0.36895 | 3.5397 | 0.76914 | 4.3088 | 0
[6 empty] | 3900 | -5.8969 | 270.1886 | 10.2201 | -2.9549 | 247.7043 | -0.36833 | 3.5512 | 0.77201 | 4.3232 | 0
[6 empty] | 3915 | -5.887 | 270.0966 | 10.2245 | -2.9505 | 247.6026 | -0.36771 | 3.5626 | 0.77487 | 4.3375 | 0
[6 empty] | 3930 | -5.8772 | 270.0049 | 10.229 | -2.946 | 247.5011 | -0.3671 | 3.5741 | 0.77772 | 4.3518 | 0
[6 empty] | 3945 | -5.8674 | 269.9132 | 10.2334 | -2.9416 | 247.3997 | -0.36648 | 3.5855 | 0.78057 | 4.3661 | 0
[6 empty] | 3960 | -5.8576 | 269.8218 | 10.2379 | -2.9371 | 247.2985 | -0.36587 | 3.5969 | 0.78342 | 4.3803 | 0
[6 empty] | 3975 | -5.8478 | 269.7305 | 10.2423 | -2.9327 | 247.1974 | -0.36526 | 3.6082 | 0.78626 | 4.3945 | 0
[6 empty] | 3990 | -5.8381 | 269.6393 | 10.2467 | -2.9283 | 247.0965 | -0.36465 | 3.6195 | 0.7891 | 4.4086 | 0
[6 empty] | 4005 | -5.8284 | 269.5483 | 10.2512 | -2.9113 | 246.9957 | -0.36405 | 3.6308 | 0.79193 | 4.4228 | 0
[6 empty] | 4020 | -5.8186 | 269.4574 | 10.2554 | -2.9071 | 246.8955 | -0.36344 | 3.6421 | 0.79475 | 4.4368 | 0
[6 empty] | 4035 | -5.8088 | 269.3667 | 10.2597 | -2.9028 | 246.7955 | -0.36283 | 3.6533 | 0.79758 | 4.4509 | 0
[6 empty] | 4050 | -5.799 | 269.2762 | 10.2639 | -2.8986 | 246.6956 | -0.36222 | 3.6645 | 0.80039 | 4.4649 | 0
[6 empty] | 4065 | -5.7893 | 269.1858 | 10.2682 | -2.8943 | 246.5958 | -0.36161 | 3.6757 | 0.80321 | 4.4789 | 0
[6 empty] | 4080 | -5.7796 | 269.0955 | 10.2724 | -2.8901 | 246.4962 | -0.361 | 3.6868 | 0.80602 | 4.4928 | 0
[6 empty] | 4095 | -5.7699 | 269.0054 | 10.2767 | -2.8858 | 246.3967 | -0.3604 | 3.6979 | 0.80882 | 4.5067 | 0
[6 empty] | 4110 | -5.7603 | 268.9155 | 10.2809 | -2.8816 | 246.2974 | -0.35979 | 3.709 | 0.81162 | 4.5206 | 0
[6 empty] | 4125 | -5.7507 | 268.8257 | 10.2852 | -2.8773 | 246.1983 | -0.35919 | 3.72 | 0.81441 | 4.5345 | 0
[6 empty] | 4140 | -5.7412 | 268.736 | 10.2895 | -2.873 | 246.0992 | -0.3586 | 3.7311 | 0.8172 | 4.5483 | 0
[6 empty] | 4155 | -5.7317 | 268.6465 | 10.2937 | -2.8688 | 246.0004 | -0.358 | 3.7421 | 0.81999 | 4.562 | 0
[6 empty] | 4170 | -5.7222 | 268.5572 | 10.298 | -2.852 | 245.9016 | -0.35741 | 3.753 | 0.82277 | 4.5758 | 0
[6 empty] | 4185 | -5.7126 | 268.468 | 10.302 | -2.848 | 245.8035 | -0.35682 | 3.7639 | 0.82554 | 4.5895 | 0
[6 empty] | 4200 | -5.703 | 268.3789 | 10.3061 | -2.8439 | 245.7055 | -0.35622 | 3.7748 | 0.82831 | 4.6032 | 0
[6 empty] | 4215 | -5.6934 | 268.29 | 10.3102 | -2.8398 | 245.6076 | -0.35562 | 3.7857 | 0.83108 | 4.6168 | 0
[6 empty] | 4230 | -5.6839 | 268.2013 | 10.3143 | -2.8357 | 245.5099 | -0.35502 | 3.7966 | 0.83384 | 4.6304 | 0
[6 empty] | 4245 | -5.6744 | 268.1126 | 10.3184 | -2.8316 | 245.4123 | -0.35443 | 3.8074 | 0.8366 | 4.644 | 0
[6 empty] | 4260 | -5.6649 | 268.0242 | 10.3224 | -2.8276 | 245.3148 | -0.35384 | 3.8182 | 0.83935 | 4.6575 | 0
[6 empty] | 4275 | -5.6555 | 267.9359 | 10.3265 | -2.8235 | 245.2175 | -0.35325 | 3.8289 | 0.8421 | 4.671 | 0
[6 empty] | 4290 | -5.6461 | 267.8477 | 10.3306 | -2.8194 | 245.1203 | -0.35266 | 3.8397 | 0.84484 | 4.6845 | 0
[6 empty] | 4305 | -5.6368 | 267.7597 | 10.3347 | -2.8153 | 245.0233 | -0.35207 | 3.8504 | 0.84758 | 4.6979 | 0
[6 empty] | 4320 | -5.6275 | 267.6718 | 10.3388 | -2.7987 | 244.9264 | -0.35149 | 3.861 | 0.85031 | 4.7113 | 0
[6 empty] | 4335 | -5.618 | 267.5841 | 10.3427 | -2.7948 | 244.8301 | -0.35091 | 3.8717 | 0.85304 | 4.7247 | 0
[6 empty] | 4350 | -5.6086 | 267.4965 | 10.3466 | -2.7909 | 244.7339 | -0.35032 | 3.8823 | 0.85577 | 4.7381 | 0
[6 empty] | 4365 | -5.5992 | 267.4091 | 10.3505 | -2.787 | 244.6379 | -0.34973 | 3.8929 | 0.85849 | 4.7514 | 0
[6 empty] | 4380 | -5.5898 | 267.3218 | 10.3545 | -2.783 | 244.542 | -0.34915 | 3.9034 | 0.8612 | 4.7646 | 0
[6 empty] | 4395 | -5.5805 | 267.2346 | 10.3584 | -2.7791 | 244.4462 | -0.34856 | 3.914 | 0.86392 | 4.7779 | 0
[6 empty] | 4410 | -5.5712 | 267.1477 | 10.3623 | -2.7752 | 244.3505 | -0.34798 | 3.9245 | 0.86662 | 4.7911 | 0
[6 empty] | 4425 | -5.562 | 267.0608 | 10.3663 | -2.7712 | 244.255 | -0.3474 | 3.9349 | 0.86933 | 4.8043 | 0
[6 empty] | 4440 | -5.5528 | 266.9741 | 10.3702 | -2.7673 | 244.1597 | -0.34683 | 3.9454 | 0.87202 | 4.8174 | 0
[6 empty] | 4455 | -5.5436 | 266.8875 | 10.3741 | -2.7634 | 244.0644 | -0.34625 | 3.9558 | 0.87472 | 4.8305 | 0
[6 empty] | 4470 | -5.5345 | 266.8011 | 10.3781 | -2.7469 | 243.9693 | -0.34568 | 3.9662 | 0.87741 | 4.8436 | 0
[6 empty] | 4485 | -5.5252 | 266.7148 | 10.3818 | -2.7432 | 243.8748 | -0.34511 | 3.9766 | 0.88009 | 4.8567 | 0
[6 empty] | 4500 | -5.5159 | 266.6287 | 10.3856 | -2.7394 | 243.7804 | -0.34453 | 3.9869 | 0.88277 | 4.8697 | 0
[6 empty] | 4515 | -5.5067 | 266.5427 | 10.3894 | -2.7356 | 243.6861 | -0.34395 | 3.9972 | 0.88545 | 4.8827 | 0
[6 empty] | 4530 | -5.4975 | 266.4569 | 10.3931 | -2.7319 | 243.592 | -0.34338 | 4.0075 | 0.88812 | 4.8956 | 0
[6 empty] | 4545 | -5.4884 | 266.3712 | 10.3969 | -2.7281 | 243.498 | -0.34281 | 4.0178 | 0.89078 | 4.9086 | 0
[6 empty] | 4560 | -5.4793 | 266.2856 | 10.4007 | -2.7243 | 243.4041 | -0.34224 | 4.028 | 0.89345 | 4.9214 | 0
[6 empty] | 4575 | -5.4702 | 266.2002 | 10.4045 | -2.7205 | 243.3103 | -0.34167 | 4.0382 | 0.8961 | 4.9343 | 0
[6 empty] | 4590 | -5.4611 | 266.1149 | 10.4083 | -2.7167 | 243.2167 | -0.3411 | 4.0484 | 0.89876 | 4.9471 | 0
[6 empty] | 4605 | -5.4521 | 266.0298 | 10.4121 | -2.7129 | 243.1232 | -0.34054 | 4.0585 | 0.90141 | 4.9599 | 0
[6 empty] | 4620 | -5.4432 | 265.9448 | 10.4159 | -2.7091 | 243.0298 | -0.33998 | 4.0687 | 0.90405 | 4.9727 | 0
[6 empty] | 4635 | -5.4343 | 265.8599 | 10.4197 | -2.6928 | 242.9366 | -0.33942 | 4.0788 | 0.90669 | 4.9854 | 0
[6 empty] | 4650 | -5.4252 | 265.7752 | 10.4233 | -2.6892 | 242.8439 | -0.33886 | 4.0888 | 0.90933 | 4.9982 | 0
[6 empty] | 4665 | -5.4161 | 265.6906 | 10.4269 | -2.6856 | 242.7514 | -0.3383 | 4.0989 | 0.91196 | 5.0108 | 0
[6 empty] | 4680 | -5.4071 | 265.6062 | 10.4306 | -2.6819 | 242.659 | -0.33773 | 4.1089 | 0.91458 | 5.0235 | 0
[6 empty] | 4695 | -5.3981 | 265.5219 | 10.4342 | -2.6783 | 242.5667 | -0.33717 | 4.1189 | 0.91721 | 5.0361 | 0
[6 empty] | 4710 | -5.3892 | 265.4378 | 10.4379 | -2.6746 | 242.4745 | -0.33661 | 4.1289 | 0.91983 | 5.0487 | 0
[6 empty] | 4725 | -5.3803 | 265.3538 | 10.4415 | -2.671 | 242.3824 | -0.33605 | 4.1388 | 0.92244 | 5.0612 | 0
[6 empty] | 4740 | -5.3714 | 265.2699 | 10.4452 | -2.6673 | 242.2905 | -0.3355 | 4.1487 | 0.92505 | 5.0738 | 0
[6 empty] | 4755 | -5.3626 | 265.1861 | 10.4488 | -2.6637 | 242.1987 | -0.33494 | 4.1586 | 0.92766 | 5.0863 | 0
[6 empty] | 4770 | -5.3538 | 265.1025 | 10.4525 | -2.66 | 242.107 | -0.33439 | 4.1685 | 0.93026 | 5.0987 | 0
[6 empty] | 4785 | -5.345 | 265.0191 | 10.4562 | -2.6563 | 242.0155 | -0.33384 | 4.1783 | 0.93285 | 5.1112 | 0
[6 empty] | 4800 | -5.3363 | 264.9358 | 10.4599 | -2.6401 | 241.9241 | -0.3333 | 4.1881 | 0.93545 | 5.1236 | 0
[6 empty] | 4815 | -5.3274 | 264.8526 | 10.4633 | -2.6367 | 241.8332 | -0.33275 | 4.1979 | 0.93803 | 5.1359 | 0
[6 empty] | 4830 | -5.3186 | 264.7695 | 10.4668 | -2.6332 | 241.7425 | -0.3322 | 4.2077 | 0.94062 | 5.1483 | 0
[6 empty] | 4845 | -5.3097 | 264.6866 | 10.4703 | -2.6297 | 241.6519 | -0.33165 | 4.2174 | 0.9432 | 5.1606 | 0
[6 empty] | 4860 | -5.301 | 264.6038 | 10.4738 | -2.6262 | 241.5614 | -0.3311 | 4.2271 | 0.94577 | 5.1729 | 0
[6 empty] | 4875 | -5.2922 | 264.5212 | 10.4774 | -2.6226 | 241.471 | -0.33055 | 4.2368 | 0.94834 | 5.1851 | 0
[6 empty] | 4890 | -5.2835 | 264.4387 | 10.4809 | -2.6191 | 241.3807 | -0.33001 | 4.2465 | 0.95091 | 5.1974 | 0
[6 empty] | 4905 | -5.2749 | 264.3563 | 10.4844 | -2.6156 | 241.2906 | -0.32946 | 4.2561 | 0.95347 | 5.2096 | 0
[6 empty] | 4920 | -5.2662 | 264.2741 | 10.488 | -2.612 | 241.2006 | -0.32892 | 4.2657 | 0.95603 | 5.2217 | 0
[6 empty] | 4935 | -5.2576 | 264.192 | 10.4915 | -2.6085 | 241.1107 | -0.32838 | 4.2753 | 0.95859 | 5.2339 | 0
[6 empty] | 4950 | -5.2491 | 264.11 | 10.4951 | -2.6049 | 241.0209 | -0.32785 | 4.2849 | 0.96114 | 5.246 | 0
[6 empty] | 4965 | -5.2405 | 264.0282 | 10.4986 | -2.5889 | 240.9313 | -0.32731 | 4.2944 | 0.96368 | 5.2581 | 0
[6 empty] | 4980 | -5.2318 | 263.9465 | 10.502 | -2.5855 | 240.8422 | -0.32678 | 4.3039 | 0.96623 | 5.2701 | 0
[6 empty] | 4995 | -5.2232 | 263.865 | 10.5053 | -2.5822 | 240.7532 | -0.32624 | 4.3134 | 0.96876 | 5.2822 | 0
[6 empty] | 5010 | -5.2146 | 263.7835 | 10.5087 | -2.5788 | 240.6643 | -0.3257 | 4.3229 | 0.9713 | 5.2942 | 0
[6 empty] | 5025 | -5.206 | 263.7022 | 10.5121 | -2.5754 | 240.5756 | -0.32516 | 4.3323 | 0.97383 | 5.3061 | 0
[6 empty] | 5040 | -5.1974 | 263.6211 | 10.5155 | -2.572 | 240.487 | -0.32463 | 4.3417 | 0.97635 | 5.3181 | 0
[6 empty] | 5055 | -5.1889 | 263.5401 | 10.5189 | -2.5686 | 240.3984 | -0.3241 | 4.3511 | 0.97887 | 5.33 | 0
[6 empty] | 5070 | -5.1805 | 263.4592 | 10.5223 | -2.5652 | 240.31 | -0.32357 | 4.3605 | 0.98139 | 5.3419 | 0
[6 empty] | 5085 | -5.172 | 263.3784 | 10.5257 | -2.5618 | 240.2218 | -0.32304 | 4.3698 | 0.9839 | 5.3537 | 0
[6 empty] | 5100 | -5.1636 | 263.2978 | 10.5292 | -2.5583 | 240.1336 | -0.32251 | 4.3791 | 0.98641 | 5.3655 | 0
[6 empty] | 5115 | -5.1553 | 263.2173 | 10.5326 | -2.5549 | 240.0456 | -0.32199 | 4.3884 | 0.98892 | 5.3774 | 0
[6 empty] | 5130 | -5.1469 | 263.1369 | 10.536 | -2.539 | 239.9576 | -0.32146 | 4.3977 | 0.99142 | 5.3891 | 0
[6 empty] | 5145 | -5.1384 | 263.0567 | 10.5393 | -2.5357 | 239.8703 | -0.32094 | 4.407 | 0.99391 | 5.4009 | 0
[6 empty] | 5160 | -5.13 | 262.9766 | 10.5425 | -2.5325 | 239.783 | -0.32041 | 4.4162 | 0.99641 | 5.4126 | 0
[6 empty] | 5175 | -5.1215 | 262.8966 | 10.5458 | -2.5292 | 239.6958 | -0.31989 | 4.4254 | 0.99889 | 5.4243 | 0
[6 empty] | 5190 | -5.1131 | 262.8168 | 10.5491 | -2.5259 | 239.6088 | -0.31936 | 4.4346 | 1.0014 | 5.4359 | 0
[6 empty] | 5205 | -5.1048 | 262.7371 | 10.5524 | -2.5226 | 239.5219 | -0.31884 | 4.4437 | 1.0039 | 5.4476 | 0
[6 empty] | 5220 | -5.0965 | 262.6575 | 10.5557 | -2.5193 | 239.435 | -0.31832 | 4.4528 | 1.0063 | 5.4592 | 0
[6 empty] | 5235 | -5.0882 | 262.578 | 10.5589 | -2.5161 | 239.3483 | -0.3178 | 4.462 | 1.0088 | 5.4708 | 0
[6 empty] | 5250 | -5.0799 | 262.4987 | 10.5623 | -2.5127 | 239.2618 | -0.31728 | 4.471 | 1.0113 | 5.4823 | 0
[6 empty] | 5265 | -5.0717 | 262.4195 | 10.5656 | -2.5094 | 239.1753 | -0.31677 | 4.4801 | 1.0137 | 5.4938 | 0
[6 empty] | 5280 | -5.0635 | 262.3405 | 10.5689 | -2.5061 | 239.0889 | -0.31626 | 4.4891 | 1.0162 | 5.5053 | 0
[6 empty] | 5295 | -5.0554 | 262.2615 | 10.5722 | -2.5028 | 239.0027 | -0.31575 | 4.4982 | 1.0187 | 5.5168 | 0
[6 empty] | 5310 | -5.0473 | 262.1827 | 10.5755 | -2.487 | 238.9165 | -0.31524 | 4.5072 | 1.0211 | 5.5283 | 0
[6 empty] | 5325 | -5.039 | 262.104 | 10.5787 | -2.4838 | 238.8309 | -0.31473 | 4.5161 | 1.0236 | 5.5397 | 0
[6 empty] | 5340 | -5.0307 | 262.0255 | 10.5818 | -2.4807 | 238.7455 | -0.31421 | 4.5251 | 1.026 | 5.5511 | 0
[6 empty] | 5355 | -5.0225 | 261.947 | 10.585 | -2.4775 | 238.6601 | -0.3137 | 4.534 | 1.0284 | 5.5624 | 0
[6 empty] | 5370 | -5.0144 | 261.8687 | 10.5881 | -2.4744 | 238.5748 | -0.31319 | 4.5429 | 1.0309 | 5.5738 | 0
[6 empty] | 5385 | -5.0062 | 261.7906 | 10.5913 | -2.4712 | 238.4897 | -0.31268 | 4.5518 | 1.0333 | 5.5851 | 0
[6 empty] | 5400 | -4.9981 | 261.7125 | 10.5945 | -2.468 | 238.4046 | -0.31217 | 4.5607 | 1.0357 | 5.5964 | 0
[6 empty] | 5415 | -4.99 | 261.6346 | 10.5977 | -2.4648 | 238.3197 | -0.31167 | 4.5695 | 1.0382 | 5.6076 | 0
[6 empty] | 5430 | -4.982 | 261.5568 | 10.6009 | -2.4616 | 238.2349 | -0.31116 | 4.5783 | 1.0406 | 5.6189 | 0
[6 empty] | 5445 | -4.974 | 261.4792 | 10.6041 | -2.4584 | 238.1502 | -0.31066 | 4.5871 | 1.043 | 5.6301 | 0
[6 empty] | 5460 | -4.966 | 261.4016 | 10.6073 | -2.4552 | 238.0656 | -0.31016 | 4.5959 | 1.0454 | 5.6413 | 0
[6 empty] | 5475 | -4.9581 | 261.3242 | 10.6105 | -2.4395 | 237.9811 | -0.30967 | 4.6046 | 1.0478 | 5.6524 | 0
[6 empty] | 5490 | -4.95 | 261.2469 | 10.6136 | -2.4364 | 237.8971 | -0.30917 | 4.6133 | 1.0502 | 5.6636 | 0
[6 empty] | 5505 | -4.9419 | 261.1697 | 10.6166 | -2.4334 | 237.8133 | -0.30866 | 4.622 | 1.0526 | 5.6747 | 0
[6 empty] | 5520 | -4.9339 | 261.0927 | 10.6196 | -2.4304 | 237.7295 | -0.30816 | 4.6307 | 1.055 | 5.6858 | 0
[6 empty] | 5535 | -4.9259 | 261.0158 | 10.6227 | -2.4273 | 237.6459 | -0.30766 | 4.6394 | 1.0574 | 5.6968 | 0
[6 empty] | 5550 | -4.9179 | 260.939 | 10.6258 | -2.4242 | 237.5623 | -0.30716 | 4.648 | 1.0598 | 5.7078 | 0
[6 empty] | 5565 | -4.91 | 260.8623 | 10.6288 | -2.4212 | 237.4789 | -0.30667 | 4.6567 | 1.0622 | 5.7188 | 0
[6 empty] | 5580 | -4.9021 | 260.7858 | 10.6319 | -2.4181 | 237.3956 | -0.30617 | 4.6652 | 1.0646 | 5.7298 | 0
[6 empty] | 5595 | -4.8942 | 260.7094 | 10.635 | -2.415 | 237.3124 | -0.30568 | 4.6738 | 1.0669 | 5.7408 | 0
[6 empty] | 5610 | -4.8864 | 260.6331 | 10.6381 | -2.4119 | 237.2293 | -0.30519 | 4.6824 | 1.0693 | 5.7517 | 0
[6 empty] | 5625 | -4.8786 | 260.5569 | 10.6412 | -2.4088 | 237.1462 | -0.3047 | 4.6909 | 1.0717 | 5.7626 | 0
[6 empty] | 5640 | -4.8708 | 260.4808 | 10.6443 | -2.4057 | 237.0633 | -0.30421 | 4.6994 | 1.0741 | 5.7735 | 0
[6 empty] | 5655 | -4.8631 | 260.4049 | 10.6474 | -2.3901 | 236.9806 | -0.30373 | 4.7079 | 1.0764 | 5.7843 | 0
[6 empty] | 5670 | -4.8552 | 260.3291 | 10.6504 | -2.3871 | 236.8983 | -0.30325 | 4.7164 | 1.0788 | 5.7952 | 0
[6 empty] | 5685 | -4.8473 | 260.2534 | 10.6533 | -2.3842 | 236.8162 | -0.30275 | 4.7248 | 1.0811 | 5.806 | 0
[6 empty] | 5700 | -4.8395 | 260.1779 | 10.6563 | -2.3812 | 236.7341 | -0.30226 | 4.7333 | 1.0835 | 5.8167 | 0
[6 empty] | 5715 | -4.8317 | 260.1024 | 10.6592 | -2.3783 | 236.6521 | -0.30178 | 4.7417 | 1.0858 | 5.8275 | 0
[6 empty] | 5730 | -4.8239 | 260.0271 | 10.6622 | -2.3753 | 236.5703 | -0.30129 | 4.7501 | 1.0882 | 5.8382 | 0
[6 empty] | 5745 | -4.8162 | 259.9519 | 10.6651 | -2.3724 | 236.4886 | -0.3008 | 4.7584 | 1.0905 | 5.8489 | 0
[6 empty] | 5760 | -4.8085 | 259.8768 | 10.6681 | -2.3694 | 236.4069 | -0.30032 | 4.7668 | 1.0929 | 5.8596 | 0
[6 empty] | 5775 | -4.8008 | 259.8018 | 10.6711 | -2.3664 | 236.3254 | -0.29984 | 4.7751 | 1.0952 | 5.8703 | 0
[6 empty] | 5790 | -4.7932 | 259.727 | 10.6741 | -2.3634 | 236.2439 | -0.29936 | 4.7834 | 1.0975 | 5.8809 | 0
[6 empty] | 5805 | -4.7856 | 259.6523 | 10.6771 | -2.3604 | 236.1626 | -0.29889 | 4.7917 | 1.0998 | 5.8915 | 0
[6 empty] | 5820 | -4.778 | 259.5777 | 10.6801 | -2.3574 | 236.0814 | -0.29841 | 4.7999 | 1.1022 | 5.9021 | 0
[6 empty] | 5835 | -4.7705 | 259.5032 | 10.6832 | -2.3543 | 236.0002 | -0.29794 | 4.8082 | 1.1045 | 5.9127 | 0
[6 empty] | 5850 | -4.763 | 259.4288 | 10.6862 | -2.3388 | 235.9192 | -0.29747 | 4.8164 | 1.1068 | 5.9232 | 0
[6 empty] | 5865 | -4.7553 | 259.3546 | 10.689 | -2.336 | 235.8387 | -0.297 | 4.8246 | 1.1091 | 5.9337 | 0
[6 empty] | 5880 | -4.7477 | 259.2804 | 10.6919 | -2.3331 | 235.7583 | -0.29653 | 4.8328 | 1.1114 | 5.9442 | 0
[6 empty] | 5895 | -4.7401 | 259.2064 | 10.6947 | -2.3303 | 235.678 | -0.29605 | 4.841 | 1.1137 | 5.9547 | 0
[6 empty] | 5910 | -4.7325 | 259.1325 | 10.6976 | -2.3274 | 235.5978 | -0.29557 | 4.8491 | 1.116 | 5.9651 | 0
[6 empty] | 5925 | -4.7249 | 259.0588 | 10.7005 | -2.3245 | 235.5177 | -0.2951 | 4.8572 | 1.1183 | 5.9755 | 0
[6 empty] | 5940 | -4.7174 | 258.9851 | 10.7033 | -2.3217 | 235.4377 | -0.29463 | 4.8653 | 1.1206 | 5.9859 | 0
[6 empty] | 5955 | -4.7099 | 258.9116 | 10.7062 | -2.3188 | 235.3578 | -0.29416 | 4.8734 | 1.1229 | 5.9963 | 0
[6 empty] | 5970 | -4.7025 | 258.8381 | 10.7091 | -2.3159 | 235.278 | -0.2937 | 4.8815 | 1.1252 | 6.0066 | 0
[6 empty] | 5985 | -4.6951 | 258.7648 | 10.712 | -2.313 | 235.1983 | -0.29323 | 4.8895 | 1.1275 | 6.017 | 0
[6 empty] | 6000 | -4.6877 | 258.6916 | 10.7149 | -2.3101 | 235.1187 | -0.29277 | 4.8975 | 1.1297 | 6.0273 | 0
[6 empty] | 6015 | -4.6803 | 258.6186 | 10.7179 | -2.3071 | 235.0392 | -0.29231 | 4.9055 | 1.132 | 6.0375 | 0
[6 empty] | 6030 | -4.673 | 258.5456 | 10.7208 | -2.2917 | 234.9598 | -0.29185 | 4.9135 | 1.1343 | 6.0478 | 0
[6 empty] | 6045 | -4.6655 | 258.4727 | 10.7235 | -2.289 | 234.881 | -0.29139 | 4.9215 | 1.1365 | 6.058 | 0
[6 empty] | 6060 | -4.658 | 258.4 | 10.7263 | -2.2862 | 234.8022 | -0.29093 | 4.9294 | 1.1388 | 6.0682 | 0
[6 empty] | 6075 | -4.6506 | 258.3274 | 10.729 | -2.2835 | 234.7235 | -0.29046 | 4.9374 | 1.1411 | 6.0784 | 0
[6 empty] | 6090 | -4.6432 | 258.2549 | 10.7318 | -2.2807 | 234.6449 | -0.29 | 4.9453 | 1.1433 | 6.0886 | 0
[6 empty] | 6105 | -4.6358 | 258.1825 | 10.7346 | -2.2779 | 234.5665 | -0.28954 | 4.9532 | 1.1456 | 6.0987 | 0
[6 empty] | 6120 | -4.6285 | 258.1102 | 10.7374 | -2.2751 | 234.4881 | -0.28908 | 4.961 | 1.1478 | 6.1088 | 0
[6 empty] | 6135 | -4.6212 | 258.0381 | 10.7402 | -2.2723 | 234.4098 | -0.28862 | 4.9689 | 1.1501 | 6.1189 | 0
[6 empty] | 6150 | -4.6139 | 257.9661 | 10.743 | -2.2695 | 234.3316 | -0.28816 | 4.9767 | 1.1523 | 6.129 | 0
[6 empty] | 6165 | -4.6067 | 257.8941 | 10.7458 | -2.2667 | 234.2535 | -0.28771 | 4.9845 | 1.1545 | 6.1391 | 0
[6 empty] | 6180 | -4.5995 | 257.8223 | 10.7486 | -2.2639 | 234.1754 | -0.28726 | 4.9923 | 1.1568 | 6.1491 | 0
[6 empty] | 6195 | -4.5923 | 257.7506 | 10.7514 | -2.2611 | 234.0975 | -0.28681 | 5.0001 | 1.159 | 6.1591 | 0
[6 empty] | 6210 | -4.5851 | 257.679 | 10.7542 | -2.2583 | 234.0197 | -0.28636 | 5.0078 | 1.1612 | 6.1691 | 0
[6 empty] | 6225 | -4.578 | 257.6075 | 10.7571 | -2.2429 | 233.942 | -0.28591 | 5.0156 | 1.1635 | 6.179 | 0
[6 empty] | 6240 | -4.5707 | 257.5362 | 10.7597 | -2.2403 | 233.8648 | -0.28547 | 5.0233 | 1.1657 | 6.189 | 0
[6 empty] | 6255 | -4.5635 | 257.4649 | 10.7624 | -2.2376 | 233.7877 | -0.28502 | 5.031 | 1.1679 | 6.1989 | 0
[6 empty] | 6270 | -4.5562 | 257.3938 | 10.765 | -2.235 | 233.7107 | -0.28456 | 5.0387 | 1.1701 | 6.2088 | 0
[6 empty] | 6285 | -4.549 | 257.3228 | 10.7677 | -2.2323 | 233.6338 | -0.28411 | 5.0464 | 1.1723 | 6.2187 | 0
[6 empty] | 6300 | -4.5419 | 257.2518 | 10.7704 | -2.2296 | 233.557 | -0.28366 | 5.054 | 1.1745 | 6.2285 | 0
[6 empty] | 6315 | -4.5347 | 257.181 | 10.7731 | -2.2269 | 233.4802 | -0.28322 | 5.0616 | 1.1767 | 6.2384 | 0
[6 empty] | 6330 | -4.5276 | 257.1103 | 10.7758 | -2.2242 | 233.4036 | -0.28277 | 5.0692 | 1.1789 | 6.2482 | 0
[6 empty] | 6345 | -4.5206 | 257.0398 | 10.7785 | -2.2215 | 233.327 | -0.28233 | 5.0768 | 1.1811 | 6.258 | 0
[6 empty] | 6360 | -4.5135 | 256.9693 | 10.7812 | -2.2188 | 233.2506 | -0.28189 | 5.0844 | 1.1833 | 6.2677 | 0
[6 empty] | 6375 | -4.5065 | 256.8989 | 10.7839 | -2.2161 | 233.1742 | -0.28145 | 5.0919 | 1.1855 | 6.2775 | 0
[6 empty] | 6390 | -4.4995 | 256.8287 | 10.7867 | -2.2133 | 233.098 | -0.28101 | 5.0995 | 1.1877 | 6.2872 | 0
[6 empty] | 6405 | -4.4925 | 256.7585 | 10.7894 | -2.2106 | 233.0218 | -0.28057 | 5.107 | 1.1899 | 6.2969 | 0
[6 empty] | 6420 | -4.4856 | 256.6885 | 10.7922 | -2.4301 | 232.9457 | -0.28014 | 5.1145 | 1.1921 | 6.3066 | 0
[6 empty] | 6435 | -4.4822 | 256.6186 | 10.7984 | -2.4238 | 232.8621 | -0.27971 | 5.122 | 1.1942 | 6.3162 | 0
[6 empty] | 6450 | -4.4787 | 256.5487 | 10.8046 | -2.4177 | 232.7787 | -0.27949 | 5.1295 | 1.1964 | 6.3259 | 0
[6 empty] | 6465 | -4.4751 | 256.4789 | 10.8106 | -2.4116 | 232.6955 | -0.27928 | 5.1369 | 1.1986 | 6.3355 | 0
[6 empty] | 6480 | -4.4715 | 256.4091 | 10.8166 | -2.4056 | 232.6125 | -0.27905 | 5.1444 | 1.2008 | 6.3451 | 0
[6 empty] | 6495 | -4.4679 | 256.3394 | 10.8226 | -2.3996 | 232.5297 | -0.27883 | 5.1518 | 1.2029 | 6.3547 | 0
[6 empty] | 6510 | -4.4642 | 256.2698 | 10.8285 | -2.3937 | 232.4471 | -0.2786 | 5.1592 | 1.2051 | 6.3643 | 0
[6 empty] | 6525 | -4.4604 | 256.2002 | 10.8343 | -2.3879 | 232.3647 | -0.27837 | 5.1666 | 1.2073 | 6.3739 | 0
[6 empty] | 6540 | -4.4566 | 256.1306 | 10.84 | -2.3822 | 232.2826 | -0.27814 | 5.174 | 1.2094 | 6.3834 | 0
[6 empty] | 6555 | -4.4527 | 256.0612 | 10.8457 | -2.3765 | 232.2006 | -0.2779 | 5.1814 | 1.2116 | 6.393 | 0
[6 empty] | 6570 | -4.4488 | 255.9917 | 10.8513 | -2.3709 | 232.1188 | -0.27766 | 5.1888 | 1.2137 | 6.4025 | 0
[6 empty] | 6585 | -4.4449 | 255.9224 | 10.8569 | -2.3653 | 232.0372 | -0.27741 | 5.1962 | 1.2159 | 6.4121 | 0
[6 empty] | 6600 | -4.4409 | 255.8531 | 10.8624 | -2.582 | 231.9558 | -0.27717 | 5.2035 | 1.2181 | 6.4216 | 0
[6 empty] | 6615 | -4.4403 | 255.7839 | 10.8713 | -2.5731 | 231.8669 | -0.27692 | 5.2108 | 1.2202 | 6.431 | 0
[6 empty] | 6630 | -4.4396 | 255.7146 | 10.8801 | -2.5643 | 231.7784 | -0.27688 | 5.2182 | 1.2224 | 6.4405 | 0
[6 empty] | 6645 | -4.4388 | 255.6454 | 10.8888 | -2.5557 | 231.6901 | -0.27684 | 5.2255 | 1.2245 | 6.45 | 0
[6 empty] | 6660 | -4.4378 | 255.5762 | 10.8973 | -2.5471 | 231.6022 | -0.27679 | 5.2328 | 1.2267 | 6.4595 | 0
[6 empty] | 6675 | -4.4368 | 255.5071 | 10.9057 | -2.5387 | 231.5145 | -0.27673 | 5.2401 | 1.2288 | 6.4689 | 0
[6 empty] | 6690 | -4.4356 | 255.4379 | 10.914 | -2.5305 | 231.4271 | -0.27666 | 5.2474 | 1.231 | 6.4784 | 0
[6 empty] | 6705 | -4.4343 | 255.3687 | 10.9221 | -2.5223 | 231.3401 | -0.27659 | 5.2547 | 1.2331 | 6.4878 | 0
[6 empty] | 6720 | -4.4329 | 255.2996 | 10.9302 | -2.5143 | 231.2533 | -0.27651 | 5.262 | 1.2353 | 6.4973 | 0
[6 empty] | 6735 | -4.4314 | 255.2305 | 10.9381 | -2.5064 | 231.1667 | -0.27642 | 5.2693 | 1.2374 | 6.5067 | 0
[6 empty] | 6750 | -4.4298 | 255.1614 | 10.9459 | -2.4986 | 231.0805 | -0.27633 | 5.2765 | 1.2396 | 6.5161 | 0
[6 empty] | 6765 | -4.4281 | 255.0924 | 10.9536 | -2.7131 | 230.9945 | -0.27623 | 5.2838 | 1.2417 | 6.5255 | 0
[6 empty] | 6780 | -4.4297 | 255.0233 | 10.9646 | -2.702 | 230.9011 | -0.27612 | 5.291 | 1.2439 | 6.5349 | 0
[6 empty] | 6795 | -4.4312 | 254.9543 | 10.9755 | -2.6911 | 230.8081 | -0.27622 | 5.2983 | 1.246 | 6.5443 | 0
[6 empty] | 6810 | -4.4325 | 254.8852 | 10.9862 | -2.6804 | 230.7155 | -0.27632 | 5.3055 | 1.2482 | 6.5537 | 0
[6 empty] | 6825 | -4.4337 | 254.8161 | 10.9967 | -2.6699 | 230.6233 | -0.2764 | 5.3128 | 1.2503 | 6.5631 | 0
[6 empty] | 6840 | -4.4346 | 254.747 | 11.0071 | -2.6596 | 230.5314 | -0.27647 | 5.32 | 1.2525 | 6.5725 | 0
[6 empty] | 6855 | -4.4354 | 254.6779 | 11.0173 | -2.6494 | 230.4398 | -0.27653 | 5.3272 | 1.2546 | 6.5819 | 0
[6 empty] | 6870 | -4.4361 | 254.6087 | 11.0273 | -2.6394 | 230.3487 | -0.27658 | 5.3345 | 1.2568 | 6.5912 | 0
[6 empty] | 6885 | -4.4365 | 254.5396 | 11.0371 | -2.6295 | 230.2578 | -0.27662 | 5.3417 | 1.2589 | 6.6006 | 0
[6 empty] | 6900 | -4.4369 | 254.4704 | 11.0468 | -2.6198 | 230.1673 | -0.27665 | 5.3489 | 1.2611 | 6.61 | 0
[6 empty] | 6915 | -4.437 | 254.4012 | 11.0564 | -2.6103 | 230.0772 | -0.27667 | 5.3561 | 1.2632 | 6.6193 | 0
[6 empty] | 6930 | -4.4371 | 254.3321 | 11.0658 | -2.8231 | 229.9873 | -0.27668 | 5.3633 | 1.2654 | 6.6287 | 0
[6 empty] | 6945 | -4.4404 | 254.2629 | 11.0785 | -2.8104 | 229.8902 | -0.27668 | 5.3705 | 1.2675 | 6.6381 | 0
[6 empty] | 6960 | -4.4436 | 254.1937 | 11.091 | -2.7979 | 229.7935 | -0.27689 | 5.3777 | 1.2697 | 6.6474 | 0
[6 empty] | 6975 | -4.4465 | 254.1244 | 11.1033 | -2.7856 | 229.6972 | -0.27709 | 5.3849 | 1.2718 | 6.6568 | 0
[6 empty] | 6990 | -4.4492 | 254.0551 | 11.1153 | -2.7735 | 229.6013 | -0.27727 | 5.3921 | 1.274 | 6.6661 | 0
[6 empty] | 7005 | -4.4517 | 253.9857 | 11.1272 | -2.7617 | 229.5059 | -0.27744 | 5.3993 | 1.2762 | 6.6755 | 0
[6 empty] | 7020 | -4.454 | 253.9163 | 11.1389 | -2.75 | 229.4108 | -0.2776 | 5.4065 | 1.2783 | 6.6848 | 0
[6 empty] | 7035 | -4.4561 | 253.8469 | 11.1503 | -2.7386 | 229.3162 | -0.27774 | 5.4137 | 1.2805 | 6.6942 | 0
[6 empty] | 7050 | -4.4581 | 253.7774 | 11.1616 | -2.7273 | 229.2219 | -0.27787 | 5.4209 | 1.2826 | 6.7035 | 0
[6 empty] | 7065 | -4.4598 | 253.7079 | 11.1727 | -2.7162 | 229.1281 | -0.27799 | 5.4281 | 1.2848 | 6.7129 | 0
[6 empty] | 7080 | -4.4613 | 253.6384 | 11.1835 | -2.7053 | 229.0346 | -0.2781 | 5.4353 | 1.287 | 6.7222 | 0
[6 empty] | 7095 | -4.4627 | 253.5688 | 11.1943 | -2.9169 | 228.9415 | -0.27819 | 5.4424 | 1.2891 | 6.7316 | 0
[6 empty] | 7110 | -4.4673 | 253.4993 | 11.2083 | -2.9029 | 228.8411 | -0.27828 | 5.4496 | 1.2913 | 6.7409 | 0
[6 empty] | 7125 | -4.4718 | 253.4296 | 11.222 | -2.8891 | 228.7412 | -0.27857 | 5.4568 | 1.2935 | 6.7503 | 0
[6 empty] | 7140 | -4.4759 | 253.3599 | 11.2355 | -2.8756 | 228.6418 | -0.27884 | 5.464 | 1.2956 | 6.7596 | 0
[6 empty] | 7155 | -4.4798 | 253.2901 | 11.2488 | -2.8623 | 228.5428 | -0.2791 | 5.4712 | 1.2978 | 6.769 | 0
[6 empty] | 7170 | -4.4835 | 253.2203 | 11.2618 | -2.8493 | 228.4443 | -0.27935 | 5.4783 | 1.3 | 6.7783 | 0
[6 empty] | 7185 | -4.487 | 253.1504 | 11.2746 | -2.8365 | 228.3463 | -0.27958 | 5.4855 | 1.3021 | 6.7877 | 0
[6 empty] | 7200 | -4.4902 | 253.0805 | 11.2872 | -2.8239 | 228.2486 | -0.27979 | 5.4927 | 1.3043 | 6.797 | 0
[6 empty] | 7215 | -4.4932 | 253.0105 | 11.2996 | -2.8116 | 228.1515 | -0.27999 | 5.4999 | 1.3065 | 6.8064 | 0
[6 empty] | 7230 | -4.496 | 252.9404 | 11.3117 | -2.7994 | 228.0547 | -0.28018 | 5.5071 | 1.3087 | 6.8157 | 0
[6 empty] | 7245 | -4.4985 | 252.8703 | 11.3236 | -3.0097 | 227.9584 | -0.28035 | 5.5142 | 1.3109 | 6.8251 | 0
[6 empty] | 7260 | -4.5044 | 252.8002 | 11.3388 | -2.9945 | 227.8548 | -0.28051 | 5.5214 | 1.313 | 6.8344 | 0
[6 empty] | 7275 | -4.5099 | 252.73 | 11.3538 | -2.9796 | 227.7517 | -0.28088 | 5.5286 | 1.3152 | 6.8438 | 0
[6 empty] | 7290 | -4.5152 | 252.6597 | 11.3684 | -2.9649 | 227.6492 | -0.28123 | 5.5358 | 1.3174 | 6.8532 | 0
[6 empty] | 7305 | -4.5202 | 252.5893 | 11.3828 | -2.9505 | 227.5471 | -0.28155 | 5.5429 | 1.3196 | 6.8626 | 0
[6 empty] | 7320 | -4.525 | 252.5188 | 11.3969 | -2.9364 | 227.4456 | -0.28187 | 5.5501 | 1.3218 | 6.8719 | 0
[6 empty] | 7335 | -4.5295 | 252.4483 | 11.4108 | -2.9225 | 227.3445 | -0.28216 | 5.5573 | 1.324 | 6.8813 | 0
[6 empty] | 7350 | -4.5337 | 252.3777 | 11.4244 | -2.9089 | 227.244 | -0.28244 | 5.5645 | 1.3262 | 6.8907 | 0
[6 empty] | 7365 | -4.5377 | 252.307 | 11.4378 | -2.8955 | 227.1439 | -0.28271 | 5.5717 | 1.3284 | 6.9001 | 0
[6 empty] | 7380 | -4.5415 | 252.2363 | 11.4509 | -2.8824 | 227.0442 | -0.28296 | 5.5789 | 1.3306 | 6.9095 | 0
[6 empty] | 7395 | -4.545 | 252.1655 | 11.4638 | -3.0917 | 226.945 | -0.28319 | 5.5861 | 1.3328 | 6.9189 | 0
[6 empty] | 7410 | -4.5517 | 252.0946 | 11.48 | -3.0756 | 226.8386 | -0.28341 | 5.5933 | 1.335 | 6.9283 | 0
[6 empty] | 7425 | -4.5582 | 252.0237 | 11.4959 | -3.0597 | 226.7328 | -0.28383 | 5.6005 | 1.3372 | 6.9377 | 0
[6 empty] | 7440 | -4.5644 | 251.9526 | 11.5114 | -3.0441 | 226.6275 | -0.28423 | 5.6077 | 1.3394 | 6.9471 | 0
[6 empty] | 7455 | -4.5702 | 251.8814 | 11.5267 | -3.0289 | 226.5227 | -0.28462 | 5.6149 | 1.3416 | 6.9565 | 0
[6 empty] | 7470 | -4.5758 | 251.8102 | 11.5417 | -3.0139 | 226.4185 | -0.28498 | 5.6221 | 1.3438 | 6.9659 | 0
[6 empty] | 7485 | -4.5811 | 251.7389 | 11.5564 | -2.9991 | 226.3147 | -0.28533 | 5.6293 | 1.3461 | 6.9753 | 0
[6 empty] | 7500 | -4.5861 | 251.6674 | 11.5709 | -2.9847 | 226.2115 | -0.28566 | 5.6365 | 1.3483 | 6.9848 | 0
[6 empty] | 7515 | -4.5909 | 251.596 | 11.5851 | -2.9705 | 226.1088 | -0.28597 | 5.6437 | 1.3505 | 6.9942 | 0
[6 empty] | 7530 | -4.5954 | 251.5244 | 11.599 | -2.9566 | 226.0066 | -0.28627 | 5.6509 | 1.3527 | 7.0036 | 0
[6 empty] | 7545 | -4.5996 | 251.4527 | 11.6127 | -3.1651 | 225.9048 | -0.28655 | 5.6581 | 1.355 | 7.0131 | 0
[6 empty] | 7560 | -4.6071 | 251.381 | 11.6296 | -3.1482 | 225.7959 | -0.28682 | 5.6653 | 1.3572 | 7.0225 | 0
[6 empty] | 7575 | -4.6142 | 251.3092 | 11.6462 | -3.1316 | 225.6876 | -0.28728 | 5.6725 | 1.3594 | 7.032 | 0
[6 empty] | 7590 | -4.6211 | 251.2373 | 11.6625 | -3.1153 | 225.5798 | -0.28773 | 5.6798 | 1.3617 | 7.0414 | 0
[6 empty] | 7605 | -4.6276 | 251.1653 | 11.6785 | -3.0993 | 225.4726 | -0.28815 | 5.687 | 1.3639 | 7.0509 | 0
[6 empty] | 7620 | -4.6338 | 251.0931 | 11.6942 | -3.0836 | 225.3659 | -0.28856 | 5.6942 | 1.3662 | 7.0604 | 0
[6 empty] | 7635 | -4.6397 | 251.0209 | 11.7096 | -3.0682 | 225.2598 | -0.28895 | 5.7015 | 1.3684 | 7.0699 | 0
[6 empty] | 7650 | -4.6453 | 250.9485 | 11.7247 | -3.0531 | 225.1542 | -0.28932 | 5.7087 | 1.3707 | 7.0794 | 0
[6 empty] | 7665 | -4.6507 | 250.8761 | 11.7395 | -3.0382 | 225.0491 | -0.28967 | 5.7159 | 1.3729 | 7.0888 | 0
[6 empty] | 7680 | -4.6558 | 250.8036 | 11.7541 | -3.2459 | 224.9446 | -0.29 | 5.7232 | 1.3752 | 7.0983 | 0
[6 empty] | 7695 | -4.664 | 250.731 | 11.7719 | -3.2281 | 224.8329 | -0.29032 | 5.7304 | 1.3774 | 7.1078 | 0
[6 empty] | 7710 | -4.672 | 250.6583 | 11.7893 | -3.2107 | 224.7218 | -0.29083 | 5.7377 | 1.3797 | 7.1174 | 0
[6 empty] | 7725 | -4.6796 | 250.5855 | 11.8065 | -3.1935 | 224.6113 | -0.29133 | 5.7449 | 1.382 | 7.1269 | 0
[6 empty] | 7740 | -4.6868 | 250.5126 | 11.8233 | -3.1767 | 224.5014 | -0.2918 | 5.7522 | 1.3842 | 7.1364 | 0
[6 empty] | 7755 | -4.6938 | 250.4395 | 11.8397 | -3.1603 | 224.3921 | -0.29226 | 5.7595 | 1.3865 | 7.146 | 0
[6 empty] | 7770 | -4.7004 | 250.3663 | 11.8559 | -3.1441 | 224.2833 | -0.29269 | 5.7667 | 1.3888 | 7.1555 | 0
[6 empty] | 7785 | -4.7067 | 250.293 | 11.8718 | -3.1282 | 224.1751 | -0.2931 | 5.774 | 1.3911 | 7.1651 | 0
[6 empty] | 7800 | -4.7128 | 250.2197 | 11.8874 | -3.1126 | 224.0674 | -0.2935 | 5.7813 | 1.3933 | 7.1746 | 0
[6 empty] | 7815 | -4.7185 | 250.1462 | 11.9027 | -3.2586 | 223.9603 | -0.29387 | 5.7886 | 1.3956 | 7.1842 | 0
[6 empty] | 7830 | -4.7264 | 250.0726 | 11.9202 | -3.2411 | 223.8482 | -0.29423 | 5.7959 | 1.3979 | 7.1938 | 0
[6 empty] | 7845 | -4.7341 | 249.999 | 11.9374 | -3.2239 | 223.7366 | -0.29473 | 5.8031 | 1.4002 | 7.2034 | 0
[6 empty] | 7860 | -4.7414 | 249.9252 | 11.9543 | -3.207 | 223.6257 | -0.2952 | 5.8104 | 1.4025 | 7.2129 | 0
[6 empty] | 7875 | -4.7483 | 249.8512 | 11.9709 | -3.1904 | 223.5153 | -0.29566 | 5.8177 | 1.4048 | 7.2225 | 0
[6 empty] | 7890 | -4.755 | 249.7772 | 11.9871 | -3.1742 | 223.4055 | -0.29609 | 5.825 | 1.4071 | 7.2322 | 0
[6 empty] | 7905 | -4.7613 | 249.7031 | 12.0031 | -3.1582 | 223.2963 | -0.29651 | 5.8324 | 1.4094 | 7.2418 | 0
[6 empty] | 7920 | -4.7674 | 249.6289 | 12.0188 | -3.1425 | 223.1876 | -0.2969 | 5.8397 | 1.4117 | 7.2514 | 0
[6 empty] | 7935 | -4.7731 | 249.5546 | 12.0341 | -3.1271 | 223.0794 | -0.29728 | 5.847 | 1.414 | 7.261 | 0
[6 empty] | 7950 | -4.7786 | 249.4801 | 12.0492 | -3.2733 | 222.9718 | -0.29764 | 5.8543 | 1.4163 | 7.2706 | 0
[6 empty] | 7965 | -4.7863 | 249.4057 | 12.0666 | -3.256 | 222.8592 | -0.29798 | 5.8616 | 1.4187 | 7.2803 | 0
[6 empty] | 7980 | -4.7937 | 249.331 | 12.0836 | -3.239 | 222.7471 | -0.29846 | 5.8689 | 1.421 | 7.2899 | 0
[6 empty] | 7995 | -4.8007 | 249.2563 | 12.1003 | -3.2223 | 222.6356 | -0.29892 | 5.8763 | 1.4233 | 7.2996 | 0
[6 empty] | 8010 | -4.8074 | 249.1815 | 12.1167 | -3.2059 | 222.5248 | -0.29936 | 5.8836 | 1.4256 | 7.3092 | 0
[6 empty] | 8025 | -4.8139 | 249.1065 | 12.1328 | -3.1898 | 222.4144 | -0.29978 | 5.8909 | 1.428 | 7.3189 | 0
[6 empty] | 8040 | -4.82 | 249.0315 | 12.1486 | -3.174 | 222.3046 | -0.30018 | 5.8983 | 1.4303 | 7.3286 | 0
[6 empty] | 8055 | -4.8258 | 248.9563 | 12.1641 | -3.1585 | 222.1954 | -0.30056 | 5.9056 | 1.4326 | 7.3383 | 0
[6 empty] | 8070 | -4.8313 | 248.8811 | 12.1793 | -3.1433 | 222.0867 | -0.30092 | 5.913 | 1.435 | 7.348 | 0
[6 empty] | 8085 | -4.8366 | 248.8058 | 12.1942 | -3.2897 | 221.9785 | -0.30126 | 5.9203 | 1.4373 | 7.3576 | 0
[6 empty] | 8100 | -4.8441 | 248.7304 | 12.2114 | -3.2725 | 221.8653 | -0.30159 | 5.9277 | 1.4397 | 7.3673 | 0
[6 empty] | 8115 | -4.8512 | 248.6549 | 12.2283 | -3.2556 | 221.7527 | -0.30206 | 5.935 | 1.442 | 7.377 | 0
[6 empty] | 8130 | -4.8581 | 248.5793 | 12.2448 | -3.2391 | 221.6407 | -0.30251 | 5.9424 | 1.4444 | 7.3867 | 0
[6 empty] | 8145 | -4.8646 | 248.5035 | 12.2611 | -3.2228 | 221.5292 | -0.30293 | 5.9497 | 1.4467 | 7.3965 | 0
[6 empty] | 8160 | -4.8708 | 248.4277 | 12.277 | -3.2069 | 221.4183 | -0.30334 | 5.9571 | 1.4491 | 7.4062 | 0
[6 empty] | 8175 | -4.8767 | 248.3518 | 12.2927 | -3.1912 | 221.3079 | -0.30373 | 5.9645 | 1.4515 | 7.4159 | 0
[6 empty] | 8190 | -4.8824 | 248.2757 | 12.308 | -3.1759 | 221.1981 | -0.3041 | 5.9718 | 1.4538 | 7.4256 | 0
[6 empty] | 8205 | -4.8877 | 248.1996 | 12.3231 | -3.1608 | 221.0888 | -0.30445 | 5.9792 | 1.4562 | 7.4354 | 0
[6 empty] | 8220 | -4.8928 | 248.1234 | 12.3379 | -3.3073 | 220.98 | -0.30478 | 5.9866 | 1.4586 | 7.4451 | 0
[6 empty] | 8235 | -4.9001 | 248.0471 | 12.355 | -3.2902 | 220.8662 | -0.3051 | 5.9939 | 1.4609 | 7.4549 | 0
[6 empty] | 8250 | -4.9071 | 247.9708 | 12.3717 | -3.2734 | 220.753 | -0.30555 | 6.0013 | 1.4633 | 7.4646 | 0
[6 empty] | 8265 | -4.9138 | 247.8943 | 12.3882 | -3.257 | 220.6403 | -0.30599 | 6.0087 | 1.4657 | 7.4744 | 0
[6 empty] | 8280 | -4.9201 | 247.8177 | 12.4043 | -3.2409 | 220.5282 | -0.30641 | 6.0161 | 1.4681 | 7.4842 | 0
[6 empty] | 8295 | -4.9262 | 247.741 | 12.4201 | -3.225 | 220.4167 | -0.3068 | 6.0235 | 1.4705 | 7.4939 | 0
[6 empty] | 8310 | -4.932 | 247.6642 | 12.4357 | -3.2095 | 220.3057 | -0.30718 | 6.0309 | 1.4728 | 7.5037 | 0
[6 empty] | 8325 | -4.9374 | 247.5873 | 12.4509 | -3.1942 | 220.1953 | -0.30754 | 6.0382 | 1.4752 | 7.5135 | 0
[6 empty] | 8340 | -4.9426 | 247.5103 | 12.4659 | -3.1793 | 220.0853 | -0.30788 | 6.0456 | 1.4776 | 7.5233 | 0
[6 empty] | 8355 | -4.9476 | 247.4333 | 12.4806 | -3.3258 | 219.9759 | -0.30821 | 6.053 | 1.48 | 7.5331 | 0
[6 empty] | 8370 | -4.9547 | 247.3561 | 12.4976 | -3.3089 | 219.8615 | -0.30851 | 6.0604 | 1.4824 | 7.5429 | 0
[6 empty] | 8385 | -4.9616 | 247.2789 | 12.5142 | -3.2922 | 219.7476 | -0.30896 | 6.0678 | 1.4848 | 7.5526 | 0
[6 empty] | 8400 | -4.9681 | 247.2015 | 12.5306 | -3.2759 | 219.6343 | -0.30939 | 6.0752 | 1.4872 | 7.5625 | 0
[6 empty] | 8415 | -4.9743 | 247.1241 | 12.5466 | -3.2598 | 219.5215 | -0.3098 | 6.0826 | 1.4897 | 7.5723 | 0
[6 empty] | 8430 | -4.9803 | 247.0465 | 12.5624 | -3.2441 | 219.4094 | -0.31018 | 6.09 | 1.4921 | 7.5821 | 0
[6 empty] | 8445 | -4.9859 | 246.9689 | 12.5778 | -3.2286 | 219.2977 | -0.31055 | 6.0974 | 1.4945 | 7.5919 | 0
[6 empty] | 8460 | -4.9913 | 246.8912 | 12.593 | -3.2135 | 219.1866 | -0.3109 | 6.1048 | 1.4969 | 7.6017 | 0
[6 empty] | 8475 | -4.9963 | 246.8134 | 12.6079 | -3.1986 | 219.076 | -0.31124 | 6.1122 | 1.4993 | 7.6116 | 0
[6 empty] | 8490 | -5.0011 | 246.7355 | 12.6225 | -3.3452 | 218.9659 | -0.31155 | 6.1196 | 1.5017 | 7.6214 | 0
[6 empty] | 8505 | -5.0082 | 246.6575 | 12.6394 | -3.3283 | 218.8508 | -0.31185 | 6.1271 | 1.5042 | 7.6312 | 0
[6 empty] | 8520 | -5.0149 | 246.5794 | 12.656 | -3.3117 | 218.7363 | -0.31229 | 6.1345 | 1.5066 | 7.6411 | 0
[6 empty] | 8535 | -5.0213 | 246.5013 | 12.6723 | -3.2955 | 218.6223 | -0.31271 | 6.1419 | 1.509 | 7.6509 | 0
[6 empty] | 8550 | -5.0275 | 246.423 | 12.6882 | -3.2795 | 218.5089 | -0.31311 | 6.1493 | 1.5115 | 7.6608 | 0
[6 empty] | 8565 | -5.0333 | 246.3446 | 12.7039 | -3.2638 | 218.396 | -0.3135 | 6.1567 | 1.5139 | 7.6706 | 0
[6 empty] | 8580 | -5.0388 | 246.2662 | 12.7193 | -3.2484 | 218.2837 | -0.31386 | 6.1641 | 1.5163 | 7.6805 | 0
[6 empty] | 8595 | -5.0441 | 246.1876 | 12.7344 | -3.2333 | 218.1719 | -0.3142 | 6.1716 | 1.5188 | 7.6904 | 0
[6 empty] | 8610 | -5.049 | 246.109 | 12.7493 | -3.2185 | 218.0606 | -0.31453 | 6.179 | 1.5212 | 7.7002 | 0
[6 empty] | 8625 | -5.0537 | 246.0303 | 12.7638 | -3.3652 | 217.9499 | -0.31484 | 6.1864 | 1.5237 | 7.7101 | 0
[6 empty] | 8640 | -5.0607 | 245.9515 | 12.7807 | -3.3484 | 217.834 | -0.31513 | 6.1938 | 1.5261 | 7.72 | 0
[6 empty] | 8655 | -5.0673 | 245.8726 | 12.7972 | -3.3319 | 217.7188 | -0.31557 | 6.2013 | 1.5286 | 7.7299 | 0
[6 empty] | 8670 | -5.0736 | 245.7936 | 12.8134 | -3.3156 | 217.6041 | -0.31598 | 6.2087 | 1.5311 | 7.7397 | 0
[6 empty] | 8685 | -5.0797 | 245.7145 | 12.8293 | -3.2997 | 217.49 | -0.31638 | 6.2161 | 1.5335 | 7.7496 | 0
[6 empty] | 8700 | -5.0854 | 245.6353 | 12.8449 | -3.2841 | 217.3765 | -0.31675 | 6.2236 | 1.536 | 7.7595 | 0
[6 empty] | 8715 | -5.0908 | 245.556 | 12.8603 | -3.2688 | 217.2635 | -0.31711 | 6.231 | 1.5384 | 7.7694 | 0
[6 empty] | 8730 | -5.096 | 245.4767 | 12.8753 | -3.2537 | 217.151 | -0.31745 | 6.2384 | 1.5409 | 7.7793 | 0
[6 empty] | 8745 | -5.1009 | 245.3972 | 12.8901 | -3.2389 | 217.039 | -0.31777 | 6.2459 | 1.5434 | 7.7892 | 0
[6 empty] | 8760 | -5.1055 | 245.3177 | 12.9046 | -3.3857 | 216.9275 | -0.31807 | 6.2533 | 1.5459 | 7.7992 | 0
[6 empty] | 8775 | -5.1124 | 245.2381 | 12.9214 | -3.3689 | 216.811 | -0.31836 | 6.2607 | 1.5483 | 7.8091 | 0
[6 empty] | 8790 | -5.1189 | 245.1584 | 12.9379 | -3.3524 | 216.6951 | -0.31879 | 6.2682 | 1.5508 | 7.819 | 0
[6 empty] | 8805 | -5.1252 | 245.0786 | 12.9541 | -3.3363 | 216.5797 | -0.3192 | 6.2756 | 1.5533 | 7.8289 | 0
[6 empty] | 8820 | -5.1311 | 244.9987 | 12.9699 | -3.3204 | 216.4649 | -0.31959 | 6.283 | 1.5558 | 7.8388 | 0
[6 empty] | 8835 | -5.1368 | 244.9187 | 12.9855 | -3.3048 | 216.3506 | -0.31996 | 6.2905 | 1.5583 | 7.8488 | 0
[6 empty] | 8850 | -5.1421 | 244.8387 | 13.0008 | -3.2895 | 216.2369 | -0.32031 | 6.2979 | 1.5608 | 7.8587 | 0
[6 empty] | 8865 | -5.1472 | 244.7585 | 13.0158 | -3.2745 | 216.1237 | -0.32065 | 6.3054 | 1.5633 | 7.8686 | 0
[6 empty] | 8880 | -5.152 | 244.6783 | 13.0306 | -3.2597 | 216.011 | -0.32096 | 6.3128 | 1.5658 | 7.8786 | 0
[6 empty] | 8895 | -5.1566 | 244.5979 | 13.0451 | -3.4065 | 215.8988 | -0.32126 | 6.3203 | 1.5683 | 7.8885 | 0
[6 empty] | 8910 | -5.1634 | 244.5176 | 13.0618 | -3.3898 | 215.7816 | -0.32155 | 6.3277 | 1.5708 | 7.8985 | 0
[6 empty] | 8925 | -5.1698 | 244.4371 | 13.0783 | -3.3734 | 215.6649 | -0.32197 | 6.3351 | 1.5733 | 7.9084 | 0
[6 empty] | 8940 | -5.176 | 244.3565 | 13.0944 | -3.3572 | 215.5488 | -0.32237 | 6.3426 | 1.5758 | 7.9184 | 0
[6 empty] | 8955 | -5.1819 | 244.2758 | 13.1102 | -3.3414 | 215.4333 | -0.32276 | 6.35 | 1.5783 | 7.9283 | 0
[6 empty] | 8970 | -5.1875 | 244.195 | 13.1258 | -3.3258 | 215.3183 | -0.32313 | 6.3575 | 1.5808 | 7.9383 | 0
[6 empty] | 8985 | -5.1928 | 244.1141 | 13.1411 | -3.3106 | 215.2038 | -0.32348 | 6.3649 | 1.5833 | 7.9483 | 0
[6 empty] | 9000 | -5.1978 | 244.0332 | 13.156 | -3.2956 | 215.0899 | -0.32381 | 6.3724 | 1.5858 | 7.9582 | 0
[6 empty] | 9015 | -5.2026 | 243.9521 | 13.1708 | -3.4421 | 214.9765 | -0.32412 | 6.3798 | 1.5883 | 7.9682 | 0
[6 empty] | 9030 | -5.2096 | 243.871 | 13.1877 | -3.4252 | 214.858 | -0.32442 | 6.3873 | 1.5909 | 7.9782 | 0
[6 empty] | 9045 | -5.2163 | 243.7898 | 13.2044 | -3.4085 | 214.7401 | -0.32485 | 6.3947 | 1.5934 | 7.9881 | 0
[6 empty] | 9060 | -5.2226 | 243.7085 | 13.2208 | -3.3921 | 214.6228 | -0.32527 | 6.4022 | 1.5959 | 7.9981 | 0
[6 empty] | 9075 | -5.2287 | 243.6271 | 13.2368 | -3.3761 | 214.5061 | -0.32567 | 6.4097 | 1.5985 | 8.0081 | 0
[6 empty] | 9090 | -5.2345 | 243.5456 | 13.2526 | -3.3603 | 214.3899 | -0.32604 | 6.4171 | 1.601 | 8.0181 | 0
[6 empty] | 9105 | -5.2399 | 243.464 | 13.2681 | -3.3448 | 214.2743 | -0.3264 | 6.4246 | 1.6035 | 8.0281 | 0
[6 empty] | 9120 | -5.2451 | 243.3823 | 13.2833 | -3.3297 | 214.1591 | -0.32675 | 6.432 | 1.6061 | 8.0381 | 0
[6 empty] | 9135 | -5.2501 | 243.3005 | 13.2982 | -3.3147 | 214.0446 | -0.32707 | 6.4395 | 1.6086 | 8.0481 | 0
[6 empty] | 9150 | -5.2547 | 243.2187 | 13.3128 | -3.4614 | 213.9305 | -0.32738 | 6.4469 | 1.6112 | 8.0581 | 0
[6 empty] | 9165 | -5.2616 | 243.1368 | 13.3297 | -3.4445 | 213.8114 | -0.32767 | 6.4544 | 1.6137 | 8.0681 | 0
[6 empty] | 9180 | -5.2682 | 243.0547 | 13.3463 | -3.4279 | 213.6928 | -0.3281 | 6.4619 | 1.6163 | 8.0781 | 0
[6 empty] | 9195 | -5.2745 | 242.9726 | 13.3626 | -3.4116 | 213.5748 | -0.32851 | 6.4693 | 1.6188 | 8.0881 | 0
[6 empty] | 9210 | -5.2804 | 242.8904 | 13.3786 | -3.3956 | 213.4574 | -0.3289 | 6.4768 | 1.6214 | 8.0982 | 0
[6 empty] | 9225 | -5.2861 | 242.8081 | 13.3943 | -3.3799 | 213.3406 | -0.32927 | 6.4842 | 1.6239 | 8.1082 | 0
[6 empty] | 9240 | -5.2915 | 242.7257 | 13.4097 | -3.3645 | 213.2243 | -0.32963 | 6.4917 | 1.6265 | 8.1182 | 0
[6 empty] | 9255 | -5.2966 | 242.6432 | 13.4249 | -3.3493 | 213.1085 | -0.32996 | 6.4992 | 1.6291 | 8.1282 | 0
[6 empty] | 9270 | -5.3014 | 242.5606 | 13.4397 | -3.4958 | 212.9932 | -0.33028 | 6.5066 | 1.6316 | 8.1383 | 0
[6 empty] | 9285 | -5.3085 | 242.478 | 13.4568 | -3.4786 | 212.8729 | -0.33058 | 6.5141 | 1.6342 | 8.1483 | 0
[6 empty] | 9300 | -5.3153 | 242.3952 | 13.4736 | -3.4618 | 212.7532 | -0.33102 | 6.5216 | 1.6368 | 8.1584 | 0
[6 empty] | 9315 | -5.3217 | 242.3123 | 13.4901 | -3.4453 | 212.634 | -0.33144 | 6.529 | 1.6394 | 8.1684 | 0
[6 empty] | 9330 | -5.3279 | 242.2294 | 13.5063 | -3.4292 | 212.5155 | -0.33185 | 6.5365 | 1.642 | 8.1784 | 0
[6 empty] | 9345 | -5.3337 | 242.1463 | 13.5222 | -3.4133 | 212.3975 | -0.33223 | 6.544 | 1.6445 | 8.1885 | 0
[6 empty] | 9360 | -5.3393 | 242.0632 | 13.5378 | -3.3977 | 212.28 | -0.33259 | 6.5514 | 1.6471 | 8.1986 | 0
[6 empty] | 9375 | -5.3445 | 241.9799 | 13.5531 | -3.3824 | 212.1631 | -0.33294 | 6.5589 | 1.6497 | 8.2086 | 0
[6 empty] | 9390 | -5.3495 | 241.8966 | 13.5682 | -3.3673 | 212.0467 | -0.33327 | 6.5664 | 1.6523 | 8.2187 | 0
[6 empty] | 9405 | -5.3542 | 241.8132 | 13.5829 | -3.5138 | 211.9308 | -0.33358 | 6.5738 | 1.6549 | 8.2287 | 0
[6 empty] | 9420 | -5.3612 | 241.7298 | 13.6 | -3.4968 | 211.8099 | -0.33387 | 6.5813 | 1.6575 | 8.2388 | 0
[6 empty] | 9435 | -5.3678 | 241.6462 | 13.6167 | -3.4801 | 211.6895 | -0.3343 | 6.5888 | 1.6601 | 8.2489 | 0
[6 empty] | 9450 | -5.3741 | 241.5625 | 13.6331 | -3.4637 | 211.5697 | -0.33472 | 6.5962 | 1.6627 | 8.2589 | 0
[6 empty] | 9465 | -5.3802 | 241.4787 | 13.6492 | -3.4476 | 211.4505 | -0.33511 | 6.6037 | 1.6653 | 8.269 | 0
[6 empty] | 9480 | -5.3859 | 241.3949 | 13.665 | -3.4318 | 211.3319 | -0.33549 | 6.6112 | 1.6679 | 8.2791 | 0
[6 empty] | 9495 | -5.3913 | 241.3109 | 13.6805 | -3.4163 | 211.2138 | -0.33585 | 6.6187 | 1.6705 | 8.2892 | 0
[6 empty] | 9510 | -5.3965 | 241.2269 | 13.6957 | -3.401 | 211.0962 | -0.33618 | 6.6261 | 1.6731 | 8.2993 | 0
[6 empty] | 9525 | -5.4013 | 241.1427 | 13.7107 | -3.5474 | 210.9792 | -0.33651 | 6.6336 | 1.6758 | 8.3094 | 0
[6 empty] | 9540 | -5.4085 | 241.0585 | 13.7279 | -3.5301 | 210.8571 | -0.33681 | 6.6411 | 1.6784 | 8.3195 | 0
[6 empty] | 9555 | -5.4153 | 240.9742 | 13.7448 | -3.5132 | 210.7356 | -0.33725 | 6.6485 | 1.681 | 8.3295 | 0
[6 empty] | 9570 | -5.4218 | 240.8898 | 13.7614 | -3.4967 | 210.6147 | -0.33768 | 6.656 | 1.6836 | 8.3396 | 0
[6 empty] | 9585 | -5.4279 | 240.8053 | 13.7777 | -3.4804 | 210.4944 | -0.33808 | 6.6635 | 1.6863 | 8.3498 | 0
[6 empty] | 9600 | -5.4338 | 240.7207 | 13.7937 | -3.4644 | 210.3746 | -0.33847 | 6.671 | 1.6889 | 8.3599 | 0
[6 empty] | 9615 | -5.4394 | 240.6359 | 13.8094 | -3.4487 | 210.2554 | -0.33883 | 6.6784 | 1.6915 | 8.37 | 0
[6 empty] | 9630 | -5.4447 | 240.5512 | 13.8248 | -3.4333 | 210.1367 | -0.33918 | 6.6859 | 1.6942 | 8.3801 | 0
[6 empty] | 9645 | -5.4497 | 240.4663 | 13.8399 | -3.4182 | 210.0185 | -0.33951 | 6.6934 | 1.6968 | 8.3902 | 0
[6 empty] | 9660 | -5.4544 | 240.3813 | 13.8547 | -3.5646 | 209.9009 | -0.33982 | 6.7009 | 1.6995 | 8.4003 | 0
[6 empty] | 9675 | -5.4614 | 240.2963 | 13.8719 | -3.5475 | 209.7782 | -0.34012 | 6.7083 | 1.7021 | 8.4104 | 0
[6 empty] | 9690 | -5.4681 | 240.2111 | 13.8886 | -3.5307 | 209.6561 | -0.34056 | 6.7158 | 1.7047 | 8.4206 | 0
[6 empty] | 9705 | -5.4744 | 240.1259 | 13.9051 | -3.5142 | 209.5346 | -0.34097 | 6.7233 | 1.7074 | 8.4307 | 0
[6 empty] | 9720 | -5.4805 | 240.0406 | 13.9213 | -3.498 | 209.4137 | -0.34137 | 6.7308 | 1.7101 | 8.4408 | 0
[6 empty] | 9735 | -5.4862 | 239.9551 | 13.9372 | -3.4822 | 209.2933 | -0.34175 | 6.7382 | 1.7127 | 8.451 | 0
[6 empty] | 9750 | -5.4917 | 239.8696 | 13.9528 | -3.4666 | 209.1734 | -0.3421 | 6.7457 | 1.7154 | 8.4611 | 0
[6 empty] | 9765 | -5.4969 | 239.784 | 13.9681 | -3.4512 | 209.0541 | -0.34244 | 6.7532 | 1.718 | 8.4712 | 0
[6 empty] | 9780 | -5.5018 | 239.6983 | 13.9831 | -3.5975 | 208.9354 | -0.34277 | 6.7607 | 1.7207 | 8.4814 | 0
[6 empty] | 9795 | -5.5089 | 239.6125 | 14.0004 | -3.5802 | 208.8116 | -0.34307 | 6.7681 | 1.7234 | 8.4915 | 0
[6 empty] | 9810 | -5.5158 | 239.5266 | 14.0174 | -3.5632 | 208.6884 | -0.34352 | 6.7756 | 1.726 | 8.5017 | 0
[6 empty] | 9825 | -5.5223 | 239.4407 | 14.0341 | -3.5466 | 208.5657 | -0.34394 | 6.7831 | 1.7287 | 8.5118 | 0
[6 empty] | 9840 | -5.5285 | 239.3546 | 14.0504 | -3.5302 | 208.4437 | -0.34435 | 6.7906 | 1.7314 | 8.522 | 0
[6 empty] | 9855 | -5.5343 | 239.2684 | 14.0665 | -3.5142 | 208.3222 | -0.34474 | 6.798 | 1.7341 | 8.5321 | 0
[6 empty] | 9870 | -5.5399 | 239.1821 | 14.0822 | -3.4984 | 208.2012 | -0.3451 | 6.8055 | 1.7368 | 8.5423 | 0
[6 empty] | 9885 | -5.5452 | 239.0958 | 14.0977 | -3.483 | 208.0808 | -0.34545 | 6.813 | 1.7395 | 8.5524 | 0
[6 empty] | 9900 | -5.5503 | 239.0093 | 14.1129 | -3.6291 | 207.961 | -0.34578 | 6.8205 | 1.7421 | 8.5626 | 0
[6 empty] | 9915 | -5.5575 | 238.9228 | 14.1303 | -3.6116 | 207.8361 | -0.3461 | 6.8279 | 1.7448 | 8.5728 | 0
[6 empty] | 9930 | -5.5645 | 238.8361 | 14.1474 | -3.5945 | 207.7118 | -0.34655 | 6.8354 | 1.7475 | 8.5829 | 0
[6 empty] | 9945 | -5.5711 | 238.7494 | 14.1642 | -3.5777 | 207.5881 | -0.34698 | 6.8429 | 1.7502 | 8.5931 | 0
[6 empty] | 9960 | -5.5774 | 238.6626 | 14.1807 | -3.5612 | 207.465 | -0.3474 | 6.8504 | 1.7529 | 8.6033 | 0
[6 empty] | 9975 | -5.5834 | 238.5756 | 14.1969 | -3.545 | 207.3424 | -0.34779 | 6.8578 | 1.7556 | 8.6135 | 0
[6 empty] | 9990 | -5.5891 | 238.4886 | 14.2128 | -3.5291 | 207.2204 | -0.34816 | 6.8653 | 1.7583 | 8.6237 | 0
[6 empty] | 10005 | -5.5945 | 238.4014 | 14.2284 | -3.5135 | 207.0989 | -0.34852 | 6.8728 | 1.7611 | 8.6339 | 0
[6 empty] | 10020 | -5.5997 | 238.3142 | 14.2437 | -3.6595 | 206.978 | -0.34886 | 6.8803 | 1.7638 | 8.644 | 0
[6 empty] | 10035 | -5.6071 | 238.2269 | 14.2613 | -3.6419 | 206.8521 | -0.34918 | 6.8877 | 1.7665 | 8.6542 | 0
[6 empty] | 10050 | -5.6141 | 238.1395 | 14.2785 | -3.6247 | 206.7268 | -0.34964 | 6.8952 | 1.7692 | 8.6644 | 0
[6 empty] | 10065 | -5.6208 | 238.052 | 14.2954 | -3.6078 | 206.602 | -0.35008 | 6.9027 | 1.7719 | 8.6746 | 0
[6 empty] | 10080 | -5.6272 | 237.9644 | 14.3121 | -3.5912 | 206.4779 | -0.3505 | 6.9102 | 1.7747 | 8.6848 | 0
[6 empty] | 10095 | -5.6333 | 237.8766 | 14.3284 | -3.5749 | 206.3543 | -0.35089 | 6.9177 | 1.7774 | 8.695 | 0
[6 empty] | 10110 | -5.6391 | 237.7888 | 14.3444 | -3.5589 | 206.2312 | -0.35127 | 6.9251 | 1.7801 | 8.7053 | 0
[6 empty] | 10125 | -5.6446 | 237.7009 | 14.3601 | -3.5432 | 206.1088 | -0.35164 | 6.9326 | 1.7829 | 8.7155 | 0
[6 empty] | 10140 | -5.6498 | 237.6129 | 14.3755 | -3.689 | 205.9868 | -0.35198 | 6.9401 | 1.7856 | 8.7257 | 0
[6 empty] | 10155 | -5.6573 | 237.5248 | 14.3932 | -3.6713 | 205.8599 | -0.3523 | 6.9476 | 1.7883 | 8.7359 | 0
[6 empty] | 10170 | -5.6644 | 237.4367 | 14.4105 | -3.654 | 205.7335 | -0.35277 | 6.955 | 1.7911 | 8.7461 | 0
[6 empty] | 10185 | -5.6712 | 237.3484 | 14.4275 | -3.637 | 205.6078 | -0.35321 | 6.9625 | 1.7938 | 8.7563 | 0
[6 empty] | 10200 | -5.6777 | 237.2599 | 14.4442 | -3.6203 | 205.4826 | -0.35364 | 6.97 | 1.7966 | 8.7666 | 0
[6 empty] | 10215 | -5.6838 | 237.1714 | 14.4606 | -3.6039 | 205.358 | -0.35404 | 6.9775 | 1.7993 | 8.7768 | 0
[6 empty] | 10230 | -5.6897 | 237.0828 | 14.4767 | -3.5878 | 205.234 | -0.35443 | 6.985 | 1.8021 | 8.787 | 0
[6 empty] | 10245 | -5.6953 | 236.9941 | 14.4926 | -3.572 | 205.1105 | -0.35479 | 6.9924 | 1.8048 | 8.7973 | 0
[6 empty] | 10260 | -5.7006 | 236.9053 | 14.5081 | -3.7177 | 204.9876 | -0.35514 | 6.9999 | 1.8076 | 8.8075 | 0
[6 empty] | 10275 | -5.7081 | 236.8165 | 14.5258 | -3.7 | 204.8596 | -0.35547 | 7.0074 | 1.8104 | 8.8178 | 0
[6 empty] | 10290 | -5.7153 | 236.7275 | 14.5433 | -3.6825 | 204.7323 | -0.35594 | 7.0149 | 1.8131 | 8.828 | 0
[6 empty] | 10305 | -5.7221 | 236.6384 | 14.5604 | -3.6654 | 204.6056 | -0.35638 | 7.0223 | 1.8159 | 8.8383 | 0
[6 empty] | 10320 | -5.7287 | 236.5492 | 14.5772 | -3.6486 | 204.4794 | -0.35681 | 7.0298 | 1.8187 | 8.8485 | 0
[6 empty] | 10335 | -5.7349 | 236.4599 | 14.5937 | -3.6322 | 204.3539 | -0.35722 | 7.0373 | 1.8215 | 8.8588 | 0
[6 empty] | 10350 | -5.7408 | 236.3705 | 14.6098 | -3.616 | 204.2289 | -0.35761 | 7.0448 | 1.8243 | 8.869 | 0
[6 empty] | 10365 | -5.7464 | 236.281 | 14.6257 | -3.6001 | 204.1044 | -0.35798 | 7.0522 | 1.827 | 8.8793 | 0
[6 empty] | 10380 | -5.7518 | 236.1914 | 14.6413 | -3.7458 | 203.9805 | -0.35833 | 7.0597 | 1.8298 | 8.8895 | 0
[6 empty] | 10395 | -5.7593 | 236.1017 | 14.6592 | -3.7279 | 203.8516 | -0.35866 | 7.0672 | 1.8326 | 8.8998 | 0
[6 empty] | 10410 | -5.7666 | 236.012 | 14.6767 | -3.7104 | 203.7233 | -0.35913 | 7.0747 | 1.8354 | 8.9101 | 0
[6 empty] | 10425 | -5.7735 | 235.9221 | 14.6938 | -3.6933 | 203.5956 | -0.35958 | 7.0822 | 1.8382 | 8.9204 | 0
[6 empty] | 10440 | -5.7801 | 235.8321 | 14.7107 | -3.6764 | 203.4685 | -0.36002 | 7.0896 | 1.841 | 8.9306 | 0
[6 empty] | 10455 | -5.7863 | 235.742 | 14.7273 | -3.6598 | 203.342 | -0.36043 | 7.0971 | 1.8438 | 8.9409 | 0
[6 empty] | 10470 | -5.7923 | 235.6518 | 14.7435 | -3.6436 | 203.216 | -0.36082 | 7.1046 | 1.8466 | 8.9512 | 0
[6 empty] | 10485 | -5.798 | 235.5615 | 14.7595 | -3.6276 | 203.0907 | -0.36119 | 7.1121 | 1.8494 | 8.9615 | 0
[6 empty] | 10500 | -5.8034 | 235.4711 | 14.7751 | -3.7733 | 202.9658 | -0.36154 | 7.1195 | 1.8522 | 8.9718 | 0
[6 empty] | 10515 | -5.811 | 235.3806 | 14.793 | -3.7554 | 202.836 | -0.36188 | 7.127 | 1.8551 | 8.9821 | 0
[6 empty] | 10530 | -5.8182 | 235.29 | 14.8106 | -3.7378 | 202.7067 | -0.36235 | 7.1345 | 1.8579 | 8.9923 | 0
[6 empty] | 10545 | -5.8252 | 235.1993 | 14.8278 | -3.7206 | 202.5781 | -0.36281 | 7.142 | 1.8607 | 9.0026 | 0
[6 empty] | 10560 | -5.8318 | 235.1085 | 14.8448 | -3.7036 | 202.45 | -0.36324 | 7.1494 | 1.8635 | 9.0129 | 0
[6 empty] | 10575 | -5.8381 | 235.0176 | 14.8614 | -3.687 | 202.3226 | -0.36365 | 7.1569 | 1.8663 | 9.0232 | 0
[6 empty] | 10590 | -5.8441 | 234.9266 | 14.8777 | -3.6707 | 202.1957 | -0.36405 | 7.1644 | 1.8692 | 9.0336 | 0
[6 empty] | 10605 | -5.8498 | 234.8355 | 14.8937 | -3.6547 | 202.0694 | -0.36442 | 7.1718 | 1.872 | 9.0439 | 0
[6 empty] | 10620 | -5.8552 | 234.7443 | 14.9094 | -3.8003 | 201.9436 | -0.36478 | 7.1793 | 1.8749 | 9.0542 | 0
[6 empty] | 10635 | -5.8629 | 234.653 | 14.9274 | -3.7823 | 201.8128 | -0.36511 | 7.1868 | 1.8777 | 9.0645 | 0
[6 empty] | 10650 | -5.8702 | 234.5616 | 14.945 | -3.7647 | 201.6826 | -0.36559 | 7.1943 | 1.8805 | 9.0748 | 0
[6 empty] | 10665 | -5.8772 | 234.4701 | 14.9623 | -3.7474 | 201.5531 | -0.36605 | 7.2017 | 1.8834 | 9.0851 | 0
[6 empty] | 10680 | -5.8838 | 234.3785 | 14.9793 | -3.7304 | 201.4241 | -0.36648 | 7.2092 | 1.8862 | 9.0954 | 0
[6 empty] | 10695 | -5.8902 | 234.2867 | 14.9959 | -3.7138 | 201.2957 | -0.3669 | 7.2167 | 1.8891 | 9.1058 | 0
[6 empty] | 10710 | -5.8962 | 234.1949 | 15.0123 | -3.6974 | 201.1679 | -0.36729 | 7.2241 | 1.8919 | 9.1161 | 0
[6 empty] | 10725 | -5.9019 | 234.103 | 15.0283 | -3.6813 | 201.0407 | -0.36767 | 7.2316 | 1.8948 | 9.1264 | 0
[6 empty] | 10740 | -5.9074 | 234.011 | 15.0441 | -3.8269 | 200.914 | -0.36802 | 7.2391 | 1.8977 | 9.1367 | 0
[6 empty] | 10755 | -5.915 | 233.9189 | 15.0621 | -3.8089 | 200.7823 | -0.36836 | 7.2465 | 1.9005 | 9.1471 | 0
[6 empty] | 10770 | -5.9224 | 233.8267 | 15.0798 | -3.7912 | 200.6512 | -0.36884 | 7.254 | 1.9034 | 9.1574 | 0
[6 empty] | 10785 | -5.9294 | 233.7344 | 15.0971 | -3.7739 | 200.5207 | -0.3693 | 7.2615 | 1.9063 | 9.1677 | 0
[6 empty] | 10800 | -5.936 | 233.6419 | 15.1141 | -3.7568 | 200.3908 | -0.36974 | 7.2689 | 1.9091 | 9.1781 | 0
[6 empty] | 10815 | -5.9424 | 233.5494 | 15.1308 | -3.7401 | 200.2616 | -0.37015 | 7.2764 | 1.912 | 9.1884 | 0
[6 empty] | 10830 | -5.9485 | 233.4568 | 15.1472 | -3.7237 | 200.1328 | -0.37055 | 7.2839 | 1.9149 | 9.1988 | 0
[6 empty] | 10845 | -5.9542 | 233.364 | 15.1633 | -3.7076 | 200.0047 | -0.37093 | 7.2913 | 1.9178 | 9.2091 | 0
[6 empty] | 10860 | -5.9597 | 233.2712 | 15.1791 | -3.8531 | 199.8771 | -0.37128 | 7.2988 | 1.9207 | 9.2195 | 0
[6 empty] | 10875 | -5.9674 | 233.1783 | 15.1972 | -3.8351 | 199.7445 | -0.37162 | 7.3063 | 1.9236 | 9.2298 | 0
[6 empty] | 10890 | -5.9747 | 233.0853 | 15.2149 | -3.8174 | 199.6125 | -0.3721 | 7.3137 | 1.9265 | 9.2402 | 0
[6 empty] | 10905 | -5.9817 | 232.9921 | 15.2323 | -3.8 | 199.4811 | -0.37256 | 7.3212 | 1.9294 | 9.2506 | 0
[6 empty] | 10920 | -5.9884 | 232.8989 | 15.2493 | -3.7829 | 199.3504 | -0.373 | 7.3286 | 1.9323 | 9.2609 | 0
[6 empty] | 10935 | -5.9948 | 232.8055 | 15.2661 | -3.7662 | 199.2202 | -0.37342 | 7.3361 | 1.9352 | 9.2713 | 0
[6 empty] | 10950 | -6.0009 | 232.7121 | 15.2825 | -3.7497 | 199.0906 | -0.37382 | 7.3436 | 1.9381 | 9.2816 | 0
[6 empty] | 10965 | -6.0066 | 232.6185 | 15.2986 | -3.8949 | 198.9615 | -0.37419 | 7.351 | 1.941 | 9.292 | 0
[6 empty] | 10980 | -6.0146 | 232.5249 | 15.317 | -3.8765 | 198.8275 | -0.37455 | 7.3585 | 1.9439 | 9.3024 | 0
[6 empty] | 10995 | -6.0223 | 232.4311 | 15.335 | -3.8585 | 198.6941 | -0.37505 | 7.3659 | 1.9468 | 9.3128 | 0
[6 empty] | 11010 | -6.0296 | 232.3372 | 15.3527 | -3.8408 | 198.5613 | -0.37553 | 7.3734 | 1.9497 | 9.3231 | 0
[6 empty] | 11025 | -6.0366 | 232.2433 | 15.3701 | -3.8235 | 198.4291 | -0.37598 | 7.3809 | 1.9527 | 9.3335 | 0
[6 empty] | 11040 | -6.0432 | 232.1491 | 15.3871 | -3.8064 | 198.2975 | -0.37642 | 7.3883 | 1.9556 | 9.3439 | 0
[6 empty] | 11055 | -6.0495 | 232.0549 | 15.4038 | -3.7897 | 198.1665 | -0.37683 | 7.3958 | 1.9585 | 9.3543 | 0
[6 empty] | 11070 | -6.0556 | 231.9606 | 15.4202 | -3.7733 | 198.0361 | -0.37723 | 7.4032 | 1.9615 | 9.3647 | 0
[6 empty] | 11085 | -6.0613 | 231.8662 | 15.4364 | -3.9185 | 197.9062 | -0.3776 | 7.4107 | 1.9644 | 9.3751 | 0
[6 empty] | 11100 | -6.0692 | 231.7717 | 15.4547 | -3.9001 | 197.7714 | -0.37796 | 7.4181 | 1.9673 | 9.3855 | 0
[6 empty] | 11115 | -6.0769 | 231.6771 | 15.4727 | -3.8821 | 197.6372 | -0.37846 | 7.4256 | 1.9703 | 9.3959 | 0
[6 empty] | 11130 | -6.0841 | 231.5824 | 15.4904 | -3.8645 | 197.5036 | -0.37893 | 7.433 | 1.9732 | 9.4063 | 0
[6 empty] | 11145 | -6.0911 | 231.4875 | 15.5077 | -3.8471 | 197.3706 | -0.37939 | 7.4405 | 1.9762 | 9.4167 | 0
[6 empty] | 11160 | -6.0977 | 231.3926 | 15.5247 | -3.8301 | 197.2382 | -0.37982 | 7.4479 | 1.9791 | 9.4271 | 0
[6 empty] | 11175 | -6.104 | 231.2975 | 15.5414 | -3.8134 | 197.1063 | -0.38023 | 7.4554 | 1.9821 | 9.4375 | 0
[6 empty] | 11190 | -6.11 | 231.2024 | 15.5578 | -3.9583 | 196.9751 | -0.38062 | 7.4628 | 1.9851 | 9.4479 | 0
[6 empty] | 11205 | -6.1182 | 231.1071 | 15.5765 | -3.9397 | 196.8389 | -0.381 | 7.4703 | 1.988 | 9.4583 | 0
[6 empty] | 11220 | -6.126 | 231.0118 | 15.5947 | -3.9214 | 196.7033 | -0.38151 | 7.4777 | 1.991 | 9.4687 | 0
[6 empty] | 11235 | -6.1336 | 230.9163 | 15.6127 | -3.9034 | 196.5684 | -0.382 | 7.4852 | 1.994 | 9.4791 | 0
[6 empty] | 11250 | -6.1407 | 230.8206 | 15.6303 | -3.8858 | 196.434 | -0.38247 | 7.4926 | 1.9969 | 9.4896 | 0
[6 empty] | 11265 | -6.1476 | 230.7249 | 15.6476 | -3.8686 | 196.3003 | -0.38292 | 7.5001 | 1.9999 | 9.5 | 0
[6 empty] | 11280 | -6.1541 | 230.6291 | 15.6645 | -3.8516 | 196.1672 | -0.38334 | 7.5075 | 2.0029 | 9.5104 | 0
[6 empty] | 11295 | -6.1603 | 230.5331 | 15.6812 | -3.835 | 196.0346 | -0.38375 | 7.515 | 2.0059 | 9.5208 | 0
[6 empty] | 11310 | -6.1662 | 230.4371 | 15.6975 | -3.9799 | 195.9026 | -0.38414 | 7.5224 | 2.0089 | 9.5313 | 0
[6 empty] | 11325 | -6.1744 | 230.341 | 15.7161 | -3.9614 | 195.7657 | -0.38451 | 7.5298 | 2.0119 | 9.5417 | 0
[6 empty] | 11340 | -6.1821 | 230.2447 | 15.7343 | -3.9431 | 195.6293 | -0.38501 | 7.5373 | 2.0149 | 9.5521 | 0
[6 empty] | 11355 | -6.1896 | 230.1483 | 15.7521 | -3.9253 | 195.4936 | -0.3855 | 7.5447 | 2.0179 | 9.5626 | 0
[6 empty] | 11370 | -6.1967 | 230.0519 | 15.7697 | -3.9077 | 195.3585 | -0.38596 | 7.5522 | 2.0209 | 9.573 | 0
[6 empty] | 11385 | -6.2034 | 229.9553 | 15.7869 | -3.8905 | 195.224 | -0.3864 | 7.5596 | 2.0239 | 9.5835 | 0
[6 empty] | 11400 | -6.2099 | 229.8585 | 15.8038 | -3.8736 | 195.0902 | -0.38683 | 7.567 | 2.0269 | 9.5939 | 0
[6 empty] | 11415 | -6.216 | 229.7617 | 15.8204 | -4.0183 | 194.9568 | -0.38723 | 7.5745 | 2.0299 | 9.6044 | 0
[6 empty] | 11430 | -6.2244 | 229.6648 | 15.8392 | -3.9995 | 194.8186 | -0.38761 | 7.5819 | 2.0329 | 9.6148 | 0
[6 empty] | 11445 | -6.2324 | 229.5678 | 15.8577 | -3.981 | 194.6809 | -0.38813 | 7.5894 | 2.0359 | 9.6253 | 0
[6 empty] | 11460 | -6.2401 | 229.4706 | 15.8758 | -3.9629 | 194.5439 | -0.38863 | 7.5968 | 2.0389 | 9.6357 | 0
[6 empty] | 11475 | -6.2474 | 229.3734 | 15.8936 | -3.9451 | 194.4075 | -0.38911 | 7.6042 | 2.042 | 9.6462 | 0
[6 empty] | 11490 | -6.2544 | 229.276 | 15.911 | -3.9277 | 194.2718 | -0.38957 | 7.6117 | 2.045 | 9.6567 | 0
[6 empty] | 11505 | -6.261 | 229.1785 | 15.9281 | -3.9106 | 194.1366 | -0.39 | 7.6191 | 2.048 | 9.6671 | 0
[6 empty] | 11520 | -6.2673 | 229.0809 | 15.9449 | -3.8938 | 194.002 | -0.39042 | 7.6265 | 2.0511 | 9.6776 | 0
[6 empty] | 11535 | -6.2734 | 228.9832 | 15.9614 | -4.0386 | 193.868 | -0.39081 | 7.634 | 2.0541 | 9.6881 | 0
[6 empty] | 11550 | -6.2816 | 228.8854 | 15.9802 | -4.0198 | 193.729 | -0.39119 | 7.6414 | 2.0571 | 9.6985 | 0
[6 empty] | 11565 | -6.2895 | 228.7874 | 15.9985 | -4.0015 | 193.5907 | -0.3917 | 7.6488 | 2.0602 | 9.709 | 0
[6 empty] | 11580 | -6.2971 | 228.6894 | 16.0166 | -3.9834 | 193.453 | -0.39219 | 7.6562 | 2.0632 | 9.7195 | 0
[6 empty] | 11595 | -6.3043 | 228.5912 | 16.0343 | -3.9657 | 193.3159 | -0.39266 | 7.6637 | 2.0663 | 9.73 | 0
[6 empty] | 11610 | -6.3112 | 228.493 | 16.0516 | -3.9484 | 193.1794 | -0.39311 | 7.6711 | 2.0694 | 9.7404 | 0
[6 empty] | 11625 | -6.3177 | 228.3946 | 16.0687 | -3.9313 | 193.0435 | -0.39354 | 7.6785 | 2.0724 | 9.7509 | 0
[6 empty] | 11640 | -6.324 | 228.2961 | 16.0854 | -3.9771 | 192.9082 | -0.39395 | 7.6859 | 2.0755 | 9.7614 | 0
[6 empty] | 11655 | -6.3309 | 228.1975 | 16.1028 | -3.9597 | 192.7713 | -0.39434 | 7.6934 | 2.0786 | 9.7719 | 0
[6 empty] | 11670 | -6.3375 | 228.0988 | 16.1199 | -3.9426 | 192.6351 | -0.39477 | 7.7008 | 2.0816 | 9.7824 | 0
[6 empty] | 11685 | -6.3438 | 228 | 16.1366 | -3.9259 | 192.4994 | -0.39518 | 7.7082 | 2.0847 | 9.7929 | 0
[6 empty] | 11700 | -6.3497 | 227.9011 | 16.1531 | -3.9094 | 192.3643 | -0.39558 | 7.7156 | 2.0878 | 9.8034 | 0
[6 empty] | 11715 | -6.3554 | 227.8021 | 16.1693 | -3.8932 | 192.2297 | -0.39595 | 7.723 | 2.0909 | 9.8139 | 0
[6 empty] | 11730 | -6.3608 | 227.7031 | 16.1851 | -3.8774 | 192.0958 | -0.3963 | 7.7304 | 2.0939 | 9.8244 | 0
[6 empty] | 11745 | -6.3659 | 227.6039 | 16.2007 | -3.9243 | 191.9623 | -0.39664 | 7.7378 | 2.097 | 9.8349 | 0
[6 empty] | 11760 | -6.3716 | 227.5047 | 16.217 | -3.908 | 191.8273 | -0.39695 | 7.7452 | 2.1001 | 9.8453 | 0
[6 empty] | 11775 | -6.3771 | 227.4053 | 16.233 | -3.892 | 191.6928 | -0.39731 | 7.7526 | 2.1032 | 9.8558 | 0
[6 empty] | 11790 | -6.3823 | 227.3059 | 16.2487 | -3.8763 | 191.5588 | -0.39766 | 7.76 | 2.1063 | 9.8663 | 0
[6 empty] | 11805 | -6.3873 | 227.2064 | 16.2641 | -3.8609 | 191.4254 | -0.39798 | 7.7674 | 2.1094 | 9.8768 | 0
[6 empty] | 11820 | -6.3919 | 227.1068 | 16.2792 | -3.8458 | 191.2926 | -0.39829 | 7.7748 | 2.1125 | 9.8873 | 0
[6 empty] | 11835 | -6.3963 | 227.0072 | 16.2941 | -3.8309 | 191.1602 | -0.39858 | 7.7822 | 2.1156 | 9.8978 | 0
[6 empty] | 11850 | -6.4004 | 226.9075 | 16.3087 | -3.8163 | 191.0284 | -0.39885 | 7.7896 | 2.1187 | 9.9083 | 0
[6 empty] | 11865 | -6.4043 | 226.8077 | 16.323 | -3.8645 | 190.897 | -0.39911 | 7.7969 | 2.1218 | 9.9187 | 0
[6 empty] | 11880 | -6.4089 | 226.7079 | 16.3381 | -3.8494 | 190.764 | -0.39935 | 7.8043 | 2.1249 | 9.9292 | 0
[6 empty] | 11895 | -6.4132 | 226.608 | 16.3529 | -3.8346 | 190.6316 | -0.39964 | 7.8117 | 2.128 | 9.9397 | 0
[6 empty] | 11910 | -6.4173 | 226.508 | 16.3675 | -3.82 | 190.4996 | -0.39991 | 7.819 | 2.1311 | 9.9502 | 0
[6 empty] | 11925 | -6.4211 | 226.4079 | 16.3817 | -3.8058 | 190.3681 | -0.40016 | 7.8264 | 2.1342 | 9.9606 | 0
[6 empty] | 11940 | -6.4247 | 226.3078 | 16.3958 | -3.7917 | 190.2372 | -0.4004 | 7.8337 | 2.1373 | 9.9711 | 0
[6 empty] | 11955 | -6.428 | 226.2077 | 16.4096 | -3.7779 | 190.1067 | -0.40062 | 7.8411 | 2.1405 | 9.9815 | 0
[6 empty] | 11970 | -6.4311 | 226.1075 | 16.4231 | -3.8269 | 189.9766 | -0.40083 | 7.8484 | 2.1436 | 9.992 | 0
[6 empty] | 11985 | -6.4349 | 226.0072 | 16.4374 | -3.8126 | 189.8449 | -0.40102 | 7.8557 | 2.1467 | 10.0025 | 0
[6 empty] | 12000 | -6.4385 | 225.9069 | 16.4514 | -3.7986 | 189.7137 | -0.40126 | 7.8631 | 2.1498 | 10.0129 | 0
[6 empty] | 12015 | -6.4419 | 225.8065 | 16.4652 | -3.7848 | 189.583 | -0.40149 | 7.8704 | 2.1529 | 10.0233 | 0
[6 empty] | 12030 | -6.445 | 225.7061 | 16.4788 | -3.7712 | 189.4528 | -0.4017 | 7.8777 | 2.1561 | 10.0338 | 0
[6 empty] | 12045 | -6.4479 | 225.6056 | 16.4921 | -3.7579 | 189.323 | -0.40189 | 7.885 | 2.1592 | 10.0442 | 0
[6 empty] | 12060 | -6.4506 | 225.5051 | 16.5052 | -3.7448 | 189.1936 | -0.40207 | 7.8923 | 2.1623 | 10.0546 | 0
[6 empty] | 12075 | -6.453 | 225.4046 | 16.5181 | -3.7319 | 189.0648 | -0.40224 | 7.8996 | 2.1655 | 10.0651 | 0
[6 empty] | 12090 | -6.4553 | 225.304 | 16.5307 | -3.7818 | 188.9363 | -0.40239 | 7.9069 | 2.1686 | 10.0755 | 0
[6 empty] | 12105 | -6.4583 | 225.2033 | 16.5442 | -3.7683 | 188.8062 | -0.40253 | 7.9142 | 2.1717 | 10.0859 | 0
[6 empty] | 12120 | -6.461 | 225.1027 | 16.5573 | -3.7552 | 188.6765 | -0.40272 | 7.9215 | 2.1748 | 10.0963 | 0
[6 empty] | 12135 | -6.4636 | 225.0019 | 16.5703 | -3.7422 | 188.5473 | -0.40289 | 7.9287 | 2.178 | 10.1067 | 0
[6 empty] | 12150 | -6.4659 | 224.9012 | 16.583 | -3.7295 | 188.4185 | -0.40305 | 7.936 | 2.1811 | 10.1171 | 0
[6 empty] | 12165 | -6.4681 | 224.8004 | 16.5956 | -3.7169 | 188.2901 | -0.40319 | 7.9432 | 2.1843 | 10.1275 | 0
[6 empty] | 12180 | -6.47 | 224.6995 | 16.6079 | -3.7046 | 188.1622 | -0.40333 | 7.9505 | 2.1874 | 10.1379 | 0
[6 empty] | 12195 | -6.4717 | 224.5987 | 16.62 | -3.6925 | 188.0347 | -0.40345 | 7.9577 | 2.1905 | 10.1482 | 0
[6 empty] | 12210 | -6.4733 | 224.4978 | 16.6319 | -3.7431 | 187.9076 | -0.40356 | 7.9649 | 2.1937 | 10.1586 | 0
[6 empty] | 12225 | -6.4756 | 224.3969 | 16.6446 | -3.7304 | 187.7788 | -0.40365 | 7.9722 | 2.1968 | 10.169 | 0
[6 empty] | 12240 | -6.4777 | 224.2959 | 16.657 | -3.718 | 187.6504 | -0.4038 | 7.9794 | 2.2 | 10.1793 | 0
[6 empty] | 12255 | -6.4796 | 224.1949 | 16.6693 | -3.7057 | 187.5225 | -0.40393 | 7.9866 | 2.2031 | 10.1897 | 0
[6 empty] | 12270 | -6.4813 | 224.0939 | 16.6814 | -3.6936 | 187.395 | -0.40405 | 7.9938 | 2.2062 | 10.2 | 0
[6 empty] | 12285 | -6.4828 | 223.9929 | 16.6932 | -3.6818 | 187.2678 | -0.40415 | 8.001 | 2.2094 | 10.2104 | 0
[6 empty] | 12300 | -6.4842 | 223.8918 | 16.7049 | -3.6701 | 187.1411 | -0.40425 | 8.0082 | 2.2125 | 10.2207 | 0
[6 empty] | 12315 | -6.4853 | 223.7908 | 16.7163 | -3.6587 | 187.0148 | -0.40433 | 8.0153 | 2.2157 | 10.231 | 0
[6 empty] | 12330 | -6.4863 | 223.6897 | 16.7276 | -3.7099 | 186.8889 | -0.4044 | 8.0225 | 2.2188 | 10.2413 | 0
[6 empty] | 12345 | -6.488 | 223.5885 | 16.7397 | -3.6978 | 186.7612 | -0.40446 | 8.0297 | 2.222 | 10.2516 | 0
[6 empty] | 12360 | -6.4896 | 223.4874 | 16.7515 | -3.686 | 186.634 | -0.40457 | 8.0368 | 2.2251 | 10.2619 | 0
[6 empty] | 12375 | -6.491 | 223.3862 | 16.7632 | -3.6743 | 186.5071 | -0.40467 | 8.044 | 2.2283 | 10.2722 | 0
[6 empty] | 12390 | -6.4922 | 223.285 | 16.7747 | -3.6628 | 186.3807 | -0.40476 | 8.0511 | 2.2314 | 10.2825 | 0
[6 empty] | 12405 | -6.4932 | 223.1838 | 16.786 | -3.6515 | 186.2546 | -0.40483 | 8.0582 | 2.2346 | 10.2928 | 0
[6 empty] | 12420 | -6.4941 | 223.0826 | 16.7971 | -3.6404 | 186.129 | -0.4049 | 8.0653 | 2.2377 | 10.3031 | 0
[6 empty] | 12435 | -6.4947 | 222.9814 | 16.808 | -3.6295 | 186.0037 | -0.40495 | 8.0725 | 2.2409 | 10.3133 | 0
[6 empty] | 12450 | -6.4952 | 222.8801 | 16.8188 | -3.6812 | 185.8788 | -0.40499 | 8.0796 | 2.244 | 10.3236 | 0
[6 empty] | 12465 | -6.4965 | 222.7789 | 16.8304 | -3.6696 | 185.7521 | -0.40502 | 8.0867 | 2.2472 | 10.3338 | 0
[6 empty] | 12480 | -6.4977 | 222.6776 | 16.8417 | -3.6583 | 185.6258 | -0.4051 | 8.0937 | 2.2503 | 10.344 | 0
[6 empty] | 12495 | -6.4986 | 222.5763 | 16.8529 | -3.6471 | 185.4999 | -0.40517 | 8.1008 | 2.2535 | 10.3543 | 0
[6 empty] | 12510 | -6.4994 | 222.475 | 16.8639 | -3.6361 | 185.3744 | -0.40523 | 8.1079 | 2.2566 | 10.3645 | 0
[6 empty] | 12525 | -6.5 | 222.3737 | 16.8747 | -3.6253 | 185.2493 | -0.40528 | 8.1149 | 2.2598 | 10.3747 | 0
[6 empty] | 12540 | -6.5005 | 222.2723 | 16.8854 | -3.6146 | 185.1245 | -0.40532 | 8.122 | 2.2629 | 10.3849 | 0
[6 empty] | 12555 | -6.5008 | 222.171 | 16.8959 | -3.6041 | 185.0001 | -0.40535 | 8.129 | 2.2661 | 10.3951 | 0
[6 empty] | 12570 | -6.5009 | 222.0697 | 16.9062 | -3.7188 | 184.8761 | -0.40537 | 8.136 | 2.2692 | 10.4053 | 0
[6 empty] | 12585 | -6.5028 | 221.9683 | 16.9183 | -3.7067 | 184.7481 | -0.40538 | 8.1431 | 2.2724 | 10.4154 | 0
[6 empty] | 12600 | -6.5046 | 221.8669 | 16.9302 | -3.6948 | 184.6205 | -0.4055 | 8.1501 | 2.2755 | 10.4256 | 0
[6 empty] | 12615 | -6.5061 | 221.7655 | 16.9419 | -3.6831 | 184.4934 | -0.4056 | 8.1571 | 2.2787 | 10.4358 | 0
[6 empty] | 12630 | -6.5075 | 221.6641 | 16.9534 | -3.6716 | 184.3666 | -0.4057 | 8.1641 | 2.2818 | 10.4459 | 0
[6 empty] | 12645 | -6.5087 | 221.5627 | 16.9647 | -3.6603 | 184.2402 | -0.40579 | 8.1711 | 2.285 | 10.4561 | 0
[6 empty] | 12660 | -6.5097 | 221.4612 | 16.9759 | -3.6491 | 184.1143 | -0.40586 | 8.1781 | 2.2882 | 10.4662 | 0
[6 empty] | 12675 | -6.5105 | 221.3597 | 16.9868 | -3.7632 | 183.9887 | -0.40592 | 8.185 | 2.2913 | 10.4763 | 0
[6 empty] | 12690 | -6.5131 | 221.2582 | 16.9996 | -3.7504 | 183.8592 | -0.40597 | 8.192 | 2.2945 | 10.4865 | 0
[6 empty] | 12705 | -6.5155 | 221.1567 | 17.0121 | -3.7379 | 183.7301 | -0.40614 | 8.1989 | 2.2976 | 10.4966 | 0
[6 empty] | 12720 | -6.5177 | 221.0551 | 17.0244 | -3.7256 | 183.6015 | -0.40629 | 8.2059 | 2.3008 | 10.5067 | 0
[6 empty] | 12735 | -6.5197 | 220.9535 | 17.0365 | -3.7135 | 183.4733 | -0.40642 | 8.2128 | 2.3039 | 10.5168 | 0
[6 empty] | 12750 | -6.5215 | 220.8519 | 17.0484 | -3.7016 | 183.3455 | -0.40655 | 8.2198 | 2.3071 | 10.5269 | 0
[6 empty] | 12765 | -6.5231 | 220.7502 | 17.0601 | -3.6899 | 183.2181 | -0.40666 | 8.2267 | 2.3103 | 10.537 | 0
[6 empty] | 12780 | -6.5245 | 220.6485 | 17.0716 | -3.6784 | 183.0911 | -0.40676 | 8.2336 | 2.3134 | 10.5471 | 0
[6 empty] | 12795 | -6.5257 | 220.5468 | 17.0829 | -3.7921 | 182.9645 | -0.40685 | 8.2405 | 2.3166 | 10.5571 | 0
[6 empty] | 12810 | -6.5288 | 220.4451 | 17.096 | -3.779 | 182.834 | -0.40692 | 8.2474 | 2.3198 | 10.5672 | 0
[6 empty] | 12825 | -6.5315 | 220.3433 | 17.1088 | -3.7662 | 182.7039 | -0.40711 | 8.2543 | 2.3229 | 10.5773 | 0
[6 empty] | 12840 | -6.5341 | 220.2415 | 17.1214 | -3.7536 | 182.5743 | -0.40729 | 8.2612 | 2.3261 | 10.5873 | 0
[6 empty] | 12855 | -6.5365 | 220.1396 | 17.1339 | -3.7411 | 182.4451 | -0.40745 | 8.2681 | 2.3293 | 10.5974 | 0
[6 empty] | 12870 | -6.5386 | 220.0377 | 17.1461 | -3.7289 | 182.3164 | -0.40759 | 8.275 | 2.3324 | 10.6074 | 0
[6 empty] | 12885 | -6.5406 | 219.9358 | 17.1581 | -3.7169 | 182.1881 | -0.40773 | 8.2819 | 2.3356 | 10.6175 | 0
[6 empty] | 12900 | -6.5424 | 219.8338 | 17.1699 | -3.7051 | 182.0601 | -0.40785 | 8.2887 | 2.3388 | 10.6275 | 0
[6 empty] | 12915 | -6.5439 | 219.7318 | 17.1815 | -3.8185 | 181.9326 | -0.40796 | 8.2956 | 2.342 | 10.6375 | 0
[6 empty] | 12930 | -6.5473 | 219.6298 | 17.1948 | -3.8052 | 181.8012 | -0.40806 | 8.3024 | 2.3451 | 10.6475 | 0
[6 empty] | 12945 | -6.5504 | 219.5278 | 17.208 | -3.792 | 181.6703 | -0.40827 | 8.3093 | 2.3483 | 10.6576 | 0
[6 empty] | 12960 | -6.5533 | 219.4256 | 17.2209 | -3.7791 | 181.5398 | -0.40846 | 8.3161 | 2.3515 | 10.6676 | 0
[6 empty] | 12975 | -6.556 | 219.3235 | 17.2335 | -3.7665 | 181.4097 | -0.40864 | 8.3229 | 2.3547 | 10.6776 | 0
[6 empty] | 12990 | -6.5584 | 219.2213 | 17.246 | -3.754 | 181.2801 | -0.40881 | 8.3297 | 2.3578 | 10.6876 | 0
[6 empty] | 13005 | -6.5607 | 219.119 | 17.2583 | -3.7417 | 181.1509 | -0.40896 | 8.3365 | 2.361 | 10.6976 | 0
[6 empty] | 13020 | -6.5627 | 219.0168 | 17.2703 | -3.7297 | 181.0221 | -0.4091 | 8.3434 | 2.3642 | 10.7076 | 0
[6 empty] | 13035 | -6.5646 | 218.9145 | 17.2821 | -3.8429 | 180.8938 | -0.40923 | 8.3502 | 2.3674 | 10.7175 | 0
[6 empty] | 13050 | -6.5682 | 218.8121 | 17.2957 | -3.8293 | 180.7615 | -0.40935 | 8.3569 | 2.3706 | 10.7275 | 0
[6 empty] | 13065 | -6.5716 | 218.7097 | 17.3091 | -3.8159 | 180.6297 | -0.40957 | 8.3637 | 2.3738 | 10.7375 | 0
[6 empty] | 13080 | -6.5748 | 218.6073 | 17.3222 | -3.8028 | 180.4984 | -0.40978 | 8.3705 | 2.377 | 10.7475 | 0
[6 empty] | 13095 | -6.5777 | 218.5048 | 17.3351 | -3.7899 | 180.3675 | -0.40998 | 8.3773 | 2.3801 | 10.7574 | 0
[6 empty] | 13110 | -6.5804 | 218.4022 | 17.3478 | -3.7772 | 180.2371 | -0.41016 | 8.384 | 2.3833 | 10.7674 | 0
[6 empty] | 13125 | -6.5829 | 218.2997 | 17.3603 | -3.7647 | 180.1071 | -0.41033 | 8.3908 | 2.3865 | 10.7773 | 0
[6 empty] | 13140 | -6.5852 | 218.197 | 17.3725 | -3.8775 | 179.9775 | -0.41049 | 8.3976 | 2.3897 | 10.7873 | 0
[6 empty] | 13155 | -6.5893 | 218.0944 | 17.3865 | -3.8635 | 179.8441 | -0.41063 | 8.4043 | 2.3929 | 10.7972 | 0
[6 empty] | 13170 | -6.5931 | 217.9917 | 17.4002 | -3.8498 | 179.7111 | -0.41089 | 8.411 | 2.3961 | 10.8072 | 0
[6 empty] | 13185 | -6.5967 | 217.8889 | 17.4137 | -3.8363 | 179.5787 | -0.41112 | 8.4178 | 2.3993 | 10.8171 | 0
[6 empty] | 13200 | -6.6 | 217.786 | 17.427 | -3.823 | 179.4466 | -0.41135 | 8.4245 | 2.4025 | 10.827 | 0
[6 empty] | 13215 | -6.6031 | 217.6832 | 17.44 | -3.81 | 179.3151 | -0.41155 | 8.4312 | 2.4057 | 10.8369 | 0
[6 empty] | 13230 | -6.606 | 217.5802 | 17.4529 | -3.7971 | 179.1839 | -0.41175 | 8.4379 | 2.4089 | 10.8469 | 0
[6 empty] | 13245 | -6.6087 | 217.4772 | 17.4654 | -3.7846 | 179.0533 | -0.41193 | 8.4447 | 2.4121 | 10.8568 | 0
[6 empty] | 13260 | -6.6111 | 217.3742 | 17.4778 | -3.8972 | 178.923 | -0.41209 | 8.4514 | 2.4153 | 10.8667 | 0
[6 empty] | 13275 | -6.6154 | 217.2712 | 17.4919 | -3.8831 | 178.7889 | -0.41225 | 8.4581 | 2.4185 | 10.8766 | 0
[6 empty] | 13290 | -6.6193 | 217.168 | 17.5058 | -3.8692 | 178.6553 | -0.41251 | 8.4647 | 2.4217 | 10.8865 | 0
[6 empty] | 13305 | -6.623 | 217.0648 | 17.5194 | -3.8556 | 178.5221 | -0.41276 | 8.4714 | 2.4249 | 10.8964 | 0
[6 empty] | 13320 | -6.6265 | 216.9616 | 17.5328 | -3.8422 | 178.3894 | -0.41299 | 8.4781 | 2.4282 | 10.9063 | 0
[6 empty] | 13335 | -6.6298 | 216.8583 | 17.546 | -3.829 | 178.2572 | -0.41321 | 8.4848 | 2.4314 | 10.9162 | 0
[6 empty] | 13350 | -6.6328 | 216.7549 | 17.5589 | -3.8161 | 178.1254 | -0.41341 | 8.4915 | 2.4346 | 10.926 | 0
[6 empty] | 13365 | -6.6356 | 216.6515 | 17.5716 | -3.9284 | 177.9941 | -0.4136 | 8.4981 | 2.4378 | 10.9359 | 0
[6 empty] | 13380 | -6.6402 | 216.5481 | 17.586 | -3.914 | 177.8589 | -0.41378 | 8.5048 | 2.441 | 10.9458 | 0
[6 empty] | 13395 | -6.6445 | 216.4446 | 17.6002 | -3.8998 | 177.7242 | -0.41406 | 8.5114 | 2.4442 | 10.9557 | 0
[6 empty] | 13410 | -6.6486 | 216.341 | 17.6141 | -3.8859 | 177.59 | -0.41433 | 8.5181 | 2.4475 | 10.9655 | 0
[6 empty] | 13425 | -6.6524 | 216.2373 | 17.6278 | -3.8722 | 177.4562 | -0.41458 | 8.5247 | 2.4507 | 10.9754 | 0
[6 empty] | 13440 | -6.656 | 216.1336 | 17.6412 | -3.8588 | 177.323 | -0.41482 | 8.5313 | 2.4539 | 10.9852 | 0
[6 empty] | 13455 | -6.6593 | 216.0299 | 17.6544 | -3.8456 | 177.1902 | -0.41504 | 8.538 | 2.4571 | 10.9951 | 0
[6 empty] | 13470 | -6.6624 | 215.9261 | 17.6674 | -3.8326 | 177.0578 | -0.41525 | 8.5446 | 2.4604 | 11.005 | 0
[6 empty] | 13485 | -6.6653 | 215.8222 | 17.6801 | -3.9449 | 176.9259 | -0.41545 | 8.5512 | 2.4636 | 11.0148 | 0
[6 empty] | 13500 | -6.67 | 215.7183 | 17.6946 | -3.9304 | 176.7902 | -0.41563 | 8.5578 | 2.4668 | 11.0246 | 0
[6 empty] | 13515 | -6.6743 | 215.6143 | 17.7088 | -3.9162 | 176.6549 | -0.41592 | 8.5644 | 2.4701 | 11.0345 | 0
[6 empty] | 13530 | -6.6785 | 215.5103 | 17.7228 | -3.9022 | 176.5201 | -0.41619 | 8.571 | 2.4733 | 11.0443 | 0
[6 empty] | 13545 | -6.6824 | 215.4062 | 17.7365 | -3.8885 | 176.3858 | -0.41645 | 8.5776 | 2.4766 | 11.0541 | 0
[6 empty] | 13560 | -6.686 | 215.302 | 17.75 | -3.875 | 176.252 | -0.41669 | 8.5842 | 2.4798 | 11.064 | 0
[6 empty] | 13575 | -6.6894 | 215.1978 | 17.7632 | -3.8618 | 176.1187 | -0.41692 | 8.5908 | 2.483 | 11.0738 | 0
[6 empty] | 13590 | -6.6926 | 215.0935 | 17.7762 | -3.9738 | 175.9858 | -0.41713 | 8.5973 | 2.4863 | 11.0836 | 0
[6 empty] | 13605 | -6.6975 | 214.9891 | 17.791 | -3.959 | 175.849 | -0.41733 | 8.6039 | 2.4895 | 11.0934 | 0
[6 empty] | 13620 | -6.7022 | 214.8847 | 17.8054 | -3.9446 | 175.7128 | -0.41764 | 8.6105 | 2.4928 | 11.1032 | 0
[6 empty] | 13635 | -6.7066 | 214.7803 | 17.8197 | -3.9303 | 175.577 | -0.41793 | 8.617 | 2.496 | 11.113 | 0
[6 empty] | 13650 | -6.7108 | 214.6757 | 17.8336 | -3.9164 | 175.4418 | -0.4182 | 8.6236 | 2.4993 | 11.1228 | 0
[6 empty] | 13665 | -6.7147 | 214.5711 | 17.8473 | -3.9027 | 175.307 | -0.41846 | 8.6301 | 2.5025 | 11.1327 | 0
[6 empty] | 13680 | -6.7183 | 214.4664 | 17.8608 | -3.8892 | 175.1727 | -0.4187 | 8.6367 | 2.5058 | 11.1424 | 0
[6 empty] | 13695 | -6.7218 | 214.3617 | 17.874 | -3.876 | 175.0388 | -0.41893 | 8.6432 | 2.5091 | 11.1522 | 0
[6 empty] | 13710 | -6.725 | 214.2569 | 17.887 | -3.988 | 174.9054 | -0.41915 | 8.6497 | 2.5123 | 11.162 | 0
[6 empty] | 13725 | -6.7299 | 214.1521 | 17.9018 | -3.9732 | 174.7682 | -0.41935 | 8.6562 | 2.5156 | 11.1718 | 0
[6 empty] | 13740 | -6.7346 | 214.0471 | 17.9162 | -3.9588 | 174.6314 | -0.41966 | 8.6628 | 2.5188 | 11.1816 | 0
[6 empty] | 13755 | -6.7391 | 213.9422 | 17.9304 | -3.9446 | 174.4952 | -0.41995 | 8.6693 | 2.5221 | 11.1914 | 0
[6 empty] | 13770 | -6.7432 | 213.8371 | 17.9444 | -3.9306 | 174.3594 | -0.42023 | 8.6758 | 2.5254 | 11.2012 | 0
[6 empty] | 13785 | -6.7472 | 213.732 | 17.9581 | -3.9169 | 174.2242 | -0.42049 | 8.6823 | 2.5286 | 11.2109 | 0
[6 empty] | 13800 | -6.7509 | 213.6268 | 17.9716 | -3.9034 | 174.0894 | -0.42073 | 8.6888 | 2.5319 | 11.2207 | 0
[6 empty] | 13815 | -6.7543 | 213.5215 | 17.9848 | -4.0152 | 173.955 | -0.42096 | 8.6953 | 2.5352 | 11.2305 | 0
[6 empty] | 13830 | -6.7595 | 213.4163 | 17.9997 | -4.0003 | 173.8169 | -0.42118 | 8.7017 | 2.5385 | 11.2402 | 0
[6 empty] | 13845 | -6.7644 | 213.3109 | 18.0144 | -3.9856 | 173.6792 | -0.4215 | 8.7082 | 2.5417 | 11.25 | 0
[6 empty] | 13860 | -6.7691 | 213.2054 | 18.0288 | -3.9712 | 173.542 | -0.42181 | 8.7147 | 2.545 | 11.2597 | 0
[6 empty] | 13875 | -6.7735 | 213.0999 | 18.043 | -3.957 | 173.4054 | -0.4221 | 8.7212 | 2.5483 | 11.2695 | 0
[6 empty] | 13890 | -6.7776 | 212.9943 | 18.0569 | -3.9431 | 173.2692 | -0.42237 | 8.7276 | 2.5516 | 11.2792 | 0
[6 empty] | 13905 | -6.7816 | 212.8887 | 18.0705 | -3.9295 | 173.1335 | -0.42263 | 8.7341 | 2.5549 | 11.289 | 0
[6 empty] | 13920 | -6.7852 | 212.7829 | 18.0839 | -3.9661 | 172.9982 | -0.42288 | 8.7405 | 2.5582 | 11.2987 | 0
[6 empty] | 13935 | -6.7895 | 212.6772 | 18.0979 | -3.9521 | 172.8618 | -0.4231 | 8.747 | 2.5615 | 11.3085 | 0
[6 empty] | 13950 | -6.7934 | 212.5713 | 18.1116 | -3.9384 | 172.7257 | -0.42337 | 8.7534 | 2.5648 | 11.3182 | 0
[6 empty] | 13965 | -6.7972 | 212.4654 | 18.1251 | -3.9249 | 172.5902 | -0.42362 | 8.7599 | 2.5681 | 11.3279 | 0
[6 empty] | 13980 | -6.8007 | 212.3594 | 18.1383 | -3.9117 | 172.4551 | -0.42385 | 8.7663 | 2.5714 | 11.3376 | 0
[6 empty] | 13995 | -6.804 | 212.2534 | 18.1513 | -3.8987 | 172.3205 | -0.42407 | 8.7727 | 2.5747 | 11.3474 | 0
[6 empty] | 14010 | -6.807 | 212.1474 | 18.1641 | -3.8859 | 172.1863 | -0.42427 | 8.7791 | 2.578 | 11.3571 | 0
[6 empty] | 14025 | -6.8099 | 212.0412 | 18.1767 | -3.8733 | 172.0526 | -0.42446 | 8.7855 | 2.5813 | 11.3668 | 0
[6 empty] | 14040 | -6.8125 | 211.9351 | 18.189 | -3.911 | 171.9193 | -0.42464 | 8.7919 | 2.5846 | 11.3765 | 0
[6 empty] | 14055 | -6.8157 | 211.8289 | 18.2019 | -3.8981 | 171.7847 | -0.4248 | 8.7983 | 2.5879 | 11.3862 | 0
[6 empty] | 14070 | -6.8187 | 211.7226 | 18.2146 | -3.8854 | 171.6506 | -0.425 | 8.8047 | 2.5912 | 11.3959 | 0
[6 empty] | 14085 | -6.8215 | 211.6163 | 18.227 | -3.873 | 171.5168 | -0.42519 | 8.8111 | 2.5945 | 11.4056 | 0
[6 empty] | 14100 | -6.824 | 211.51 | 18.2393 | -3.8607 | 171.3836 | -0.42536 | 8.8175 | 2.5978 | 11.4153 | 0
[6 empty] | 14115 | -6.8264 | 211.4036 | 18.2513 | -3.8487 | 171.2507 | -0.42552 | 8.8238 | 2.6011 | 11.4249 | 0
[6 empty] | 14130 | -6.8286 | 211.2972 | 18.2632 | -3.8368 | 171.1182 | -0.42567 | 8.8302 | 2.6044 | 11.4346 | 0
[6 empty] | 14145 | -6.8305 | 211.1907 | 18.2748 | -3.8752 | 170.9862 | -0.42581 | 8.8366 | 2.6077 | 11.4443 | 0
[6 empty] | 14160 | -6.8331 | 211.0843 | 18.287 | -3.863 | 170.8528 | -0.42593 | 8.8429 | 2.611 | 11.4539 | 0
[6 empty] | 14175 | -6.8354 | 210.9777 | 18.299 | -3.851 | 170.7199 | -0.42609 | 8.8492 | 2.6143 | 11.4636 | 0
[6 empty] | 14190 | -6.8376 | 210.8712 | 18.3108 | -3.8392 | 170.5874 | -0.42624 | 8.8556 | 2.6177 | 11.4732 | 0
[6 empty] | 14205 | -6.8396 | 210.7646 | 18.3224 | -3.8276 | 170.4552 | -0.42637 | 8.8619 | 2.621 | 11.4829 | 0
[6 empty] | 14220 | -6.8414 | 210.658 | 18.3338 | -3.8162 | 170.3235 | -0.42649 | 8.8682 | 2.6243 | 11.4925 | 0
[6 empty] | 14235 | -6.8429 | 210.5513 | 18.3451 | -3.8049 | 170.1922 | -0.4266 | 8.8745 | 2.6276 | 11.5021 | 0
[6 empty] | 14250 | -6.8444 | 210.4446 | 18.3561 | -3.7939 | 170.0612 | -0.4267 | 8.8808 | 2.6309 | 11.5117 | 0
[6 empty] | 14265 | -6.8456 | 210.3379 | 18.3669 | -3.8331 | 169.9307 | -0.42679 | 8.8871 | 2.6343 | 11.5213 | 0
[6 empty] | 14280 | -6.8475 | 210.2312 | 18.3784 | -3.8216 | 169.7988 | -0.42687 | 8.8934 | 2.6376 | 11.5309 | 0
[6 empty] | 14295 | -6.8491 | 210.1245 | 18.3897 | -3.8103 | 169.6672 | -0.42698 | 8.8996 | 2.6409 | 11.5405 | 0
[6 empty] | 14310 | -6.8506 | 210.0177 | 18.4007 | -3.7993 | 169.5361 | -0.42709 | 8.9059 | 2.6442 | 11.5501 | 0
[6 empty] | 14325 | -6.8519 | 209.9109 | 18.4116 | -3.7884 | 169.4054 | -0.42718 | 8.9122 | 2.6475 | 11.5597 | 0
[6 empty] | 14340 | -6.853 | 209.8041 | 18.4223 | -3.7777 | 169.275 | -0.42726 | 8.9184 | 2.6509 | 11.5693 | 0
[6 empty] | 14355 | -6.854 | 209.6973 | 18.4329 | -3.7671 | 169.145 | -0.42733 | 8.9247 | 2.6542 | 11.5788 | 0
[6 empty] | 14370 | -6.8548 | 209.5904 | 18.4432 | -3.7568 | 169.0153 | -0.42739 | 8.9309 | 2.6575 | 11.5884 | 0
[6 empty] | 14385 | -6.8555 | 209.4836 | 18.4534 | -3.7966 | 168.886 | -0.42744 | 8.9371 | 2.6608 | 11.5979 | 0
[6 empty] | 14400 | -6.8567 | 209.3767 | 18.4642 | -3.7858 | 168.7554 | -0.42748 | 8.9433 | 2.6642 | 11.6075 | 0
[6 empty] | 14415 | -6.8578 | 209.2698 | 18.4749 | -3.7751 | 168.6251 | -0.42756 | 8.9495 | 2.6675 | 11.617 | 0
[6 empty] | 14430 | -6.8588 | 209.1629 | 18.4853 | -3.7647 | 168.4952 | -0.42763 | 8.9557 | 2.6708 | 11.6265 | 0
[6 empty] | 14445 | -6.8596 | 209.056 | 18.4956 | -3.7544 | 168.3656 | -0.42769 | 8.9619 | 2.6741 | 11.6361 | 0
[6 empty] | 14460 | -6.8602 | 208.949 | 18.5057 | -3.7443 | 168.2364 | -0.42774 | 8.9681 | 2.6775 | 11.6456 | 0
[6 empty] | 14475 | -6.8606 | 208.8421 | 18.5157 | -3.7343 | 168.1075 | -0.42778 | 8.9743 | 2.6808 | 11.6551 | 0
[6 empty] | 14490 | -6.8609 | 208.7351 | 18.5255 | -3.7745 | 167.979 | -0.42781 | 8.9804 | 2.6841 | 11.6646 | 0
[6 empty] | 14505 | -6.8619 | 208.6282 | 18.5359 | -3.7641 | 167.8491 | -0.42783 | 8.9866 | 2.6875 | 11.674 | 0
[6 empty] | 14520 | -6.8627 | 208.5212 | 18.5462 | -3.7538 | 167.7196 | -0.42788 | 8.9927 | 2.6908 | 11.6835 | 0
[6 empty] | 14535 | -6.8633 | 208.4142 | 18.5563 | -3.7437 | 167.5904 | -0.42793 | 8.9989 | 2.6941 | 11.693 | 0
[6 empty] | 14550 | -6.8638 | 208.3072 | 18.5662 | -3.7338 | 167.4616 | -0.42797 | 9.005 | 2.6974 | 11.7024 | 0
[6 empty] | 14565 | -6.8641 | 208.2002 | 18.576 | -3.724 | 167.3331 | -0.428 | 9.0111 | 2.7008 | 11.7119 | 0
[6 empty] | 14580 | -6.8643 | 208.0932 | 18.5856 | -3.7144 | 167.2049 | -0.42802 | 9.0172 | 2.7041 | 11.7213 | 0
[6 empty] | 14595 | -6.8643 | 207.9862 | 18.5951 | -3.7049 | 167.0771 | -0.42803 | 9.0233 | 2.7074 | 11.7307 | 0
[6 empty] | 14610 | -6.8642 | 207.8792 | 18.6044 | -3.7456 | 166.9496 | -0.42804 | 9.0294 | 2.7108 | 11.7402 | 0
[6 empty] | 14625 | -6.8648 | 207.7722 | 18.6143 | -3.7357 | 166.8207 | -0.42803 | 9.0355 | 2.7141 | 11.7496 | 0
[6 empty] | 14640 | -6.8652 | 207.6652 | 18.6241 | -3.7259 | 166.6921 | -0.42806 | 9.0416 | 2.7174 | 11.759 | 0
[6 empty] | 14655 | -6.8654 | 207.5582 | 18.6338 | -3.7162 | 166.5639 | -0.42809 | 9.0476 | 2.7207 | 11.7684 | 0
[6 empty] | 14670 | -6.8655 | 207.4511 | 18.6432 | -3.7068 | 166.436 | -0.4281 | 9.0537 | 2.7241 | 11.7777 | 0
[6 empty] | 14685 | -6.8655 | 207.3441 | 18.6526 | -3.6974 | 166.3084 | -0.42811 | 9.0597 | 2.7274 | 11.7871 | 0
[6 empty] | 14700 | -6.8653 | 207.2371 | 18.6618 | -3.6882 | 166.1812 | -0.42811 | 9.0657 | 2.7307 | 11.7965 | 0
[6 empty] | 14715 | -6.865 | 207.1301 | 18.6708 | -3.6792 | 166.0542 | -0.4281 | 9.0718 | 2.7341 | 11.8058 | 0
[6 empty] | 14730 | -6.8645 | 207.023 | 18.6797 | -3.7203 | 165.9276 | -0.42808 | 9.0778 | 2.7374 | 11.8152 | 0
[6 empty] | 14745 | -6.8648 | 206.916 | 18.6893 | -3.7107 | 165.7996 | -0.42805 | 9.0838 | 2.7407 | 11.8245 | 0
[6 empty] | 14760 | -6.8648 | 206.809 | 18.6987 | -3.7013 | 165.6719 | -0.42806 | 9.0898 | 2.7441 | 11.8338 | 0
[6 empty] | 14775 | -6.8648 | 206.702 | 18.7079 | -3.6921 | 165.5445 | -0.42807 | 9.0958 | 2.7474 | 11.8432 | 0
[6 empty] | 14790 | -6.8646 | 206.595 | 18.7171 | -3.6829 | 165.4175 | -0.42807 | 9.1018 | 2.7507 | 11.8525 | 0
[6 empty] | 14805 | -6.8643 | 206.488 | 18.726 | -3.674 | 165.2907 | -0.42805 | 9.1077 | 2.754 | 11.8618 | 0
[6 empty] | 14820 | -6.8638 | 206.3809 | 18.7349 | -3.6651 | 165.1643 | -0.42803 | 9.1137 | 2.7574 | 11.8711 | 0
[6 empty] | 14835 | -6.8632 | 206.2739 | 18.7436 | -3.6564 | 165.0381 | -0.428 | 9.1196 | 2.7607 | 11.8803 | 0
[6 empty] | 14850 | -6.8625 | 206.167 | 18.7521 | -3.6979 | 164.9123 | -0.42797 | 9.1256 | 2.764 | 11.8896 | 0
[6 empty] | 14865 | -6.8625 | 206.06 | 18.7613 | -3.6887 | 164.785 | -0.42792 | 9.1315 | 2.7674 | 11.8989 | 0
[6 empty] | 14880 | -6.8623 | 205.953 | 18.7704 | -3.6796 | 164.6581 | -0.42792 | 9.1374 | 2.7707 | 11.9081 | 0
[6 empty] | 14895 | -6.862 | 205.846 | 18.7793 | -3.6707 | 164.5315 | -0.42791 | 9.1433 | 2.774 | 11.9173 | 0
[6 empty] | 14910 | -6.8616 | 205.739 | 18.7881 | -3.6619 | 164.4051 | -0.42789 | 9.1492 | 2.7773 | 11.9266 | 0
[6 empty] | 14925 | -6.861 | 205.6321 | 18.7968 | -3.6532 | 164.2791 | -0.42786 | 9.1551 | 2.7807 | 11.9358 | 0
[6 empty] | 14940 | -6.8603 | 205.5251 | 18.8053 | -3.6447 | 164.1534 | -0.42783 | 9.161 | 2.784 | 11.945 | 0
[6 empty] | 14955 | -6.8595 | 205.4182 | 18.8137 | -3.6363 | 164.028 | -0.42779 | 9.1669 | 2.7873 | 11.9542 | 0
[6 empty] | 14970 | -6.8586 | 205.3112 | 18.822 | -3.678 | 163.9028 | -0.42774 | 9.1727 | 2.7907 | 11.9634 | 0
[6 empty] | 14985 | -6.8584 | 205.2043 | 18.8309 | -3.6691 | 163.7762 | -0.42768 | 9.1786 | 2.794 | 11.9726 | 0
[6 empty] | 15000 | -6.858 | 205.0974 | 18.8397 | -3.6603 | 163.65 | -0.42767 | 9.1844 | 2.7973 | 11.9817 | 0
[6 empty] | 15015 | -6.8575 | 204.9905 | 18.8484 | -3.6516 | 163.524 | -0.42764 | 9.1903 | 2.8006 | 11.9909 | 0
[6 empty] | 15030 | -6.8569 | 204.8836 | 18.8569 | -3.6431 | 163.3983 | -0.42761 | 9.1961 | 2.804 | 12.0001 | 0
[6 empty] | 15045 | -6.8561 | 204.7767 | 18.8653 | -3.6347 | 163.273 | -0.42757 | 9.2019 | 2.8073 | 12.0092 | 0
[6 empty] | 15060 | -6.8553 | 204.6698 | 18.8736 | -3.6264 | 163.1479 | -0.42753 | 9.2077 | 2.8106 | 12.0183 | 0
[6 empty] | 15075 | -6.8543 | 204.5629 | 18.8818 | -3.6182 | 163.0231 | -0.42747 | 9.2135 | 2.8139 | 12.0274 | 0
[6 empty] | 15090 | -6.8532 | 204.4561 | 18.8898 | -3.7769 | 162.8986 | -0.42741 | 9.2193 | 2.8173 | 12.0366 | 0
[6 empty] | 15105 | -6.8547 | 204.3493 | 18.9003 | -3.7664 | 162.7686 | -0.42734 | 9.2251 | 2.8206 | 12.0457 | 0
[6 empty] | 15120 | -6.8559 | 204.2424 | 18.9107 | -3.756 | 162.639 | -0.42743 | 9.2308 | 2.8239 | 12.0547 | 0
[6 empty] | 15135 | -6.857 | 204.1355 | 18.9208 | -3.7458 | 162.5097 | -0.42751 | 9.2366 | 2.8272 | 12.0638 | 0
[6 empty] | 15150 | -6.8579 | 204.0286 | 18.9308 | -3.7358 | 162.3808 | -0.42758 | 9.2423 | 2.8306 | 12.0729 | 0
[6 empty] | 15165 | -6.8587 | 203.9217 | 18.9407 | -3.726 | 162.2522 | -0.42764 | 9.2481 | 2.8339 | 12.082 | 0
[6 empty] | 15180 | -6.8593 | 203.8148 | 18.9504 | -3.7163 | 162.124 | -0.42769 | 9.2538 | 2.8372 | 12.091 | 0
[6 empty] | 15195 | -6.8598 | 203.7079 | 18.9599 | -3.8734 | 161.9961 | -0.42772 | 9.2596 | 2.8405 | 12.1001 | 0
[6 empty] | 15210 | -6.8627 | 203.6009 | 18.9719 | -3.8615 | 161.8628 | -0.42775 | 9.2653 | 2.8439 | 12.1091 | 0
[6 empty] | 15225 | -6.8655 | 203.4939 | 18.9837 | -3.8497 | 161.7299 | -0.42794 | 9.271 | 2.8472 | 12.1182 | 0
[6 empty] | 15240 | -6.868 | 203.3869 | 18.9952 | -3.8381 | 161.5974 | -0.42811 | 9.2767 | 2.8505 | 12.1272 | 0
[6 empty] | 15255 | -6.8704 | 203.2798 | 19.0066 | -3.8267 | 161.4653 | -0.42827 | 9.2824 | 2.8539 | 12.1362 | 0
[6 empty] | 15270 | -6.8725 | 203.1727 | 19.0178 | -3.8156 | 161.3336 | -0.42841 | 9.2881 | 2.8572 | 12.1453 | 0
[6 empty] | 15285 | -6.8745 | 203.0656 | 19.0288 | -3.8046 | 161.2023 | -0.42855 | 9.2938 | 2.8605 | 12.1543 | 0
[6 empty] | 15300 | -6.8763 | 202.9584 | 19.0396 | -3.7938 | 161.0714 | -0.42867 | 9.2994 | 2.8639 | 12.1633 | 0
[6 empty] | 15315 | -6.8779 | 202.8512 | 19.0502 | -3.9498 | 160.9408 | -0.42878 | 9.3051 | 2.8672 | 12.1723 | 0
[6 empty] | 15330 | -6.8819 | 202.744 | 19.0632 | -3.9368 | 160.8049 | -0.42888 | 9.3108 | 2.8705 | 12.1813 | 0
[6 empty] | 15345 | -6.8858 | 202.6367 | 19.0761 | -3.9239 | 160.6694 | -0.42914 | 9.3164 | 2.8739 | 12.1903 | 0
[6 empty] | 15360 | -6.8894 | 202.5294 | 19.0887 | -3.9113 | 160.5344 | -0.42937 | 9.3221 | 2.8772 | 12.1993 | 0
[6 empty] | 15375 | -6.8928 | 202.422 | 19.101 | -3.899 | 160.3998 | -0.4296 | 9.3277 | 2.8806 | 12.2083 | 0
[6 empty] | 15390 | -6.8959 | 202.3145 | 19.1132 | -3.8868 | 160.2656 | -0.42981 | 9.3333 | 2.8839 | 12.2172 | 0
[6 empty] | 15405 | -6.8989 | 202.207 | 19.1251 | -3.8749 | 160.1318 | -0.43001 | 9.339 | 2.8872 | 12.2262 | 0
[6 empty] | 15420 | -6.9016 | 202.0995 | 19.1368 | -4.0298 | 159.9985 | -0.43019 | 9.3446 | 2.8906 | 12.2352 | 0
[6 empty] | 15435 | -6.9068 | 201.9919 | 19.151 | -4.0157 | 159.8598 | -0.43036 | 9.3502 | 2.8939 | 12.2442 | 0
[6 empty] | 15450 | -6.9117 | 201.8842 | 19.1648 | -4.0018 | 159.7216 | -0.43069 | 9.3558 | 2.8973 | 12.2531 | 0
[6 empty] | 15465 | -6.9164 | 201.7765 | 19.1785 | -3.9882 | 159.5839 | -0.43099 | 9.3614 | 2.9006 | 12.2621 | 0
[6 empty] | 15480 | -6.9208 | 201.6687 | 19.1918 | -3.9748 | 159.4466 | -0.43128 | 9.367 | 2.904 | 12.271 | 0
[6 empty] | 15495 | -6.925 | 201.5608 | 19.205 | -3.9617 | 159.3098 | -0.43156 | 9.3726 | 2.9074 | 12.28 | 0
[6 empty] | 15510 | -6.929 | 201.4528 | 19.2179 | -3.9488 | 159.1735 | -0.43182 | 9.3782 | 2.9107 | 12.2889 | 0
[6 empty] | 15525 | -6.9327 | 201.3448 | 19.2306 | -3.9361 | 159.0376 | -0.43207 | 9.3838 | 2.9141 | 12.2979 | 0
[6 empty] | 15540 | -6.9362 | 201.2367 | 19.243 | -4.0903 | 158.9021 | -0.4323 | 9.3894 | 2.9174 | 12.3068 | 0
[6 empty] | 15555 | -6.9421 | 201.1286 | 19.2578 | -4.0755 | 158.7613 | -0.43252 | 9.395 | 2.9208 | 12.3158 | 0
[6 empty] | 15570 | -6.9477 | 201.0204 | 19.2724 | -4.0609 | 158.6211 | -0.43289 | 9.4005 | 2.9242 | 12.3247 | 0
[6 empty] | 15585 | -6.9531 | 200.9121 | 19.2867 | -4.0466 | 158.4813 | -0.43324 | 9.4061 | 2.9275 | 12.3336 | 0
[6 empty] | 15600 | -6.9582 | 200.8037 | 19.3007 | -4.0326 | 158.3421 | -0.43357 | 9.4117 | 2.9309 | 12.3426 | 0
[6 empty] | 15615 | -6.963 | 200.6952 | 19.3145 | -4.0188 | 158.2033 | -0.43389 | 9.4172 | 2.9343 | 12.3515 | 0
[6 empty] | 15630 | -6.9676 | 200.5867 | 19.328 | -4.0053 | 158.065 | -0.43419 | 9.4228 | 2.9377 | 12.3604 | 0
[6 empty] | 15645 | -6.972 | 200.478 | 19.3413 | -4.1587 | 157.9271 | -0.43448 | 9.4283 | 2.941 | 12.3693 | 0
[6 empty] | 15660 | -6.9787 | 200.3694 | 19.357 | -4.143 | 157.784 | -0.43475 | 9.4338 | 2.9444 | 12.3783 | 0
[6 empty] | 15675 | -6.9851 | 200.2606 | 19.3723 | -4.1277 | 157.6414 | -0.43517 | 9.4394 | 2.9478 | 12.3872 | 0
[6 empty] | 15690 | -6.9913 | 200.1517 | 19.3874 | -4.1126 | 157.4994 | -0.43557 | 9.4449 | 2.9512 | 12.3961 | 0
[6 empty] | 15705 | -6.9972 | 200.0427 | 19.4022 | -4.0978 | 157.3579 | -0.43595 | 9.4504 | 2.9546 | 12.405 | 0
[6 empty] | 15720 | -7.0028 | 199.9336 | 19.4167 | -4.0833 | 157.2168 | -0.43632 | 9.456 | 2.958 | 12.4139 | 0
[6 empty] | 15735 | -7.0081 | 199.8244 | 19.431 | -4.069 | 157.0763 | -0.43667 | 9.4615 | 2.9614 | 12.4229 | 0
[6 empty] | 15750 | -7.0132 | 199.7152 | 19.445 | -4.2217 | 156.9363 | -0.437 | 9.467 | 2.9648 | 12.4318 | 0
[6 empty] | 15765 | -7.0206 | 199.6058 | 19.4613 | -4.2054 | 156.791 | -0.43732 | 9.4725 | 2.9682 | 12.4407 | 0
[6 empty] | 15780 | -7.0278 | 199.4964 | 19.4773 | -4.1893 | 156.6463 | -0.43778 | 9.478 | 2.9716 | 12.4496 | 0
[6 empty] | 15795 | -7.0346 | 199.3868 | 19.4931 | -4.1736 | 156.5021 | -0.43823 | 9.4835 | 2.975 | 12.4585 | 0
[6 empty] | 15810 | -7.0411 | 199.2772 | 19.5085 | -4.1581 | 156.3584 | -0.43865 | 9.489 | 2.9784 | 12.4674 | 0
[6 empty] | 15825 | -7.0474 | 199.1674 | 19.5237 | -4.143 | 156.2153 | -0.43906 | 9.4945 | 2.9818 | 12.4763 | 0
[6 empty] | 15840 | -7.0533 | 199.0576 | 19.5385 | -4.1281 | 156.0728 | -0.43945 | 9.5 | 2.9852 | 12.4852 | 0
[6 empty] | 15855 | -7.059 | 198.9476 | 19.5531 | -4.2802 | 155.9307 | -0.43982 | 9.5055 | 2.9887 | 12.4941 | 0
[6 empty] | 15870 | -7.067 | 198.8376 | 19.5701 | -4.2633 | 155.7834 | -0.44018 | 9.511 | 2.9921 | 12.503 | 0
[6 empty] | 15885 | -7.0747 | 198.7274 | 19.5867 | -4.2466 | 155.6367 | -0.44068 | 9.5164 | 2.9955 | 12.5119 | 0
[6 empty] | 15900 | -7.0821 | 198.6171 | 19.603 | -4.2303 | 155.4905 | -0.44116 | 9.5219 | 2.9989 | 12.5208 | 0
[6 empty] | 15915 | -7.0892 | 198.5067 | 19.619 | -4.2144 | 155.3449 | -0.44162 | 9.5274 | 3.0024 | 12.5298 | 0
[6 empty] | 15930 | -7.096 | 198.3962 | 19.6347 | -4.1987 | 155.1999 | -0.44206 | 9.5328 | 3.0058 | 12.5387 | 0
[6 empty] | 15945 | -7.1025 | 198.2856 | 19.6501 | -4.1833 | 155.0554 | -0.44248 | 9.5383 | 3.0093 | 12.5476 | 0
[6 empty] | 15960 | -7.1087 | 198.1748 | 19.6652 | -4.3348 | 154.9114 | -0.44289 | 9.5438 | 3.0127 | 12.5565 | 0
[6 empty] | 15975 | -7.1172 | 198.064 | 19.6826 | -4.3174 | 154.7623 | -0.44328 | 9.5492 | 3.0162 | 12.5654 | 0
[6 empty] | 15990 | -7.1254 | 197.9531 | 19.6997 | -4.3003 | 154.6137 | -0.44381 | 9.5547 | 3.0196 | 12.5743 | 0
[6 empty] | 16005 | -7.1333 | 197.842 | 19.7165 | -4.2835 | 154.4657 | -0.44432 | 9.5601 | 3.0231 | 12.5832 | 0
[6 empty] | 16020 | -7.1409 | 197.7308 | 19.733 | -4.267 | 154.3183 | -0.44481 | 9.5656 | 3.0265 | 12.5921 | 0
[6 empty] | 16035 | -7.1481 | 197.6195 | 19.7491 | -4.2509 | 154.1714 | -0.44528 | 9.571 | 3.03 | 12.601 | 0
[6 empty] | 16050 | -7.1551 | 197.508 | 19.765 | -4.235 | 154.0251 | -0.44573 | 9.5764 | 3.0335 | 12.6099 | 0
[6 empty] | 16065 | -7.1617 | 197.3965 | 19.7805 | -4.3862 | 153.8794 | -0.44617 | 9.5819 | 3.0369 | 12.6188 | 0
[6 empty] | 16080 | -7.1707 | 197.2848 | 19.7984 | -4.3683 | 153.7284 | -0.44658 | 9.5873 | 3.0404 | 12.6277 | 0
[6 empty] | 16095 | -7.1793 | 197.1731 | 19.8159 | -4.3508 | 153.5781 | -0.44714 | 9.5927 | 3.0439 | 12.6366 | 0
[6 empty] | 16110 | -7.1875 | 197.0611 | 19.8331 | -4.3336 | 153.4284 | -0.44768 | 9.5982 | 3.0474 | 12.6455 | 0
[6 empty] | 16125 | -7.1955 | 196.9491 | 19.8499 | -4.3167 | 153.2792 | -0.44819 | 9.6036 | 3.0508 | 12.6544 | 0
[6 empty] | 16140 | -7.2031 | 196.8369 | 19.8665 | -4.3002 | 153.1307 | -0.44869 | 9.609 | 3.0543 | 12.6633 | 0
[6 empty] | 16155 | -7.2104 | 196.7246 | 19.8827 | -4.4506 | 152.9827 | -0.44916 | 9.6144 | 3.0578 | 12.6723 | 0
[6 empty] | 16170 | -7.2201 | 196.6122 | 19.9012 | -4.4321 | 152.8295 | -0.44962 | 9.6198 | 3.0613 | 12.6812 | 0
[6 empty] | 16185 | -7.2293 | 196.4997 | 19.9194 | -4.4139 | 152.677 | -0.45022 | 9.6253 | 3.0648 | 12.6901 | 0
[6 empty] | 16200 | -7.2382 | 196.387 | 19.9372 | -4.3961 | 152.5251 | -0.4508 | 9.6307 | 3.0683 | 12.699 | 0
[6 empty] | 16215 | -7.2468 | 196.2741 | 19.9547 | -4.3786 | 152.3738 | -0.45135 | 9.6361 | 3.0718 | 12.7079 | 0
[6 empty] | 16230 | -7.255 | 196.1612 | 19.9718 | -4.3615 | 152.2231 | -0.45189 | 9.6415 | 3.0754 | 12.7168 | 0
[6 empty] | 16245 | -7.2629 | 196.0481 | 19.9887 | -4.3447 | 152.073 | -0.4524 | 9.6469 | 3.0789 | 12.7257 | 0
[6 empty] | 16260 | -7.2705 | 195.9348 | 20.0052 | -4.4948 | 151.9235 | -0.45289 | 9.6523 | 3.0824 | 12.7347 | 0
[6 empty] | 16275 | -7.2804 | 195.8215 | 20.024 | -4.476 | 151.7688 | -0.45336 | 9.6577 | 3.0859 | 12.7436 | 0
[6 empty] | 16290 | -7.2899 | 195.708 | 20.0424 | -4.4576 | 151.6147 | -0.45398 | 9.663 | 3.0895 | 12.7525 | 0
[6 empty] | 16305 | -7.299 | 195.5944 | 20.0605 | -4.4395 | 151.4613 | -0.45457 | 9.6684 | 3.093 | 12.7614 | 0
[6 empty] | 16320 | -7.3078 | 195.4806 | 20.0782 | -4.4218 | 151.3086 | -0.45514 | 9.6738 | 3.0965 | 12.7704 | 0
[6 empty] | 16335 | -7.3163 | 195.3666 | 20.0956 | -4.4044 | 151.1564 | -0.45569 | 9.6792 | 3.1001 | 12.7793 | 0
[6 empty] | 16350 | -7.3244 | 195.2526 | 20.1126 | -4.3874 | 151.0048 | -0.45622 | 9.6846 | 3.1036 | 12.7882 | 0
[6 empty] | 16365 | -7.3322 | 195.1384 | 20.1294 | -4.5373 | 150.8538 | -0.45673 | 9.69 | 3.1072 | 12.7971 | 0
[6 empty] | 16380 | -7.3423 | 195.0241 | 20.1484 | -4.5183 | 150.6977 | -0.45721 | 9.6953 | 3.1107 | 12.8061 | 0
[6 empty] | 16395 | -7.352 | 194.9096 | 20.167 | -4.4996 | 150.5422 | -0.45784 | 9.7007 | 3.1143 | 12.815 | 0
[6 empty] | 16410 | -7.3614 | 194.795 | 20.1853 | -4.4813 | 150.3873 | -0.45845 | 9.7061 | 3.1179 | 12.8239 | 0
[6 empty] | 16425 | -7.3704 | 194.6803 | 20.2033 | -4.4634 | 150.2331 | -0.45903 | 9.7114 | 3.1214 | 12.8329 | 0
[6 empty] | 16440 | -7.3791 | 194.5654 | 20.2209 | -4.4458 | 150.0795 | -0.45959 | 9.7168 | 3.125 | 12.8418 | 0
[6 empty] | 16455 | -7.3874 | 194.4503 | 20.2381 | -4.5952 | 149.9265 | -0.46013 | 9.7221 | 3.1286 | 12.8507 | 0
[6 empty] | 16470 | -7.398 | 194.3352 | 20.2577 | -4.5757 | 149.7683 | -0.46065 | 9.7275 | 3.1322 | 12.8597 | 0
[6 empty] | 16485 | -7.4082 | 194.2199 | 20.2768 | -4.5565 | 149.6109 | -0.46131 | 9.7329 | 3.1358 | 12.8686 | 0
[6 empty] | 16500 | -7.418 | 194.1044 | 20.2956 | -4.5377 | 149.4541 | -0.46195 | 9.7382 | 3.1394 | 12.8776 | 0
[6 empty] | 16515 | -7.4275 | 193.9887 | 20.314 | -4.5193 | 149.2979 | -0.46256 | 9.7436 | 3.143 | 12.8865 | 0
[6 empty] | 16530 | -7.4366 | 193.8729 | 20.3321 | -4.5013 | 149.1424 | -0.46315 | 9.7489 | 3.1466 | 12.8955 | 0
[6 empty] | 16545 | -7.4454 | 193.757 | 20.3498 | -4.6502 | 148.9875 | -0.46372 | 9.7542 | 3.1502 | 12.9044 | 0
[6 empty] | 16560 | -7.4564 | 193.6409 | 20.3698 | -4.6302 | 148.8274 | -0.46427 | 9.7596 | 3.1538 | 12.9134 | 0
[6 empty] | 16575 | -7.4671 | 193.5247 | 20.3894 | -4.6106 | 148.6681 | -0.46496 | 9.7649 | 3.1574 | 12.9223 | 0
[6 empty] | 16590 | -7.4773 | 193.4083 | 20.4086 | -4.5914 | 148.5094 | -0.46562 | 9.7703 | 3.161 | 12.9313 | 0
[6 empty] | 16605 | -7.4872 | 193.2917 | 20.4274 | -4.5726 | 148.3514 | -0.46626 | 9.7756 | 3.1646 | 12.9402 | 0
[6 empty] | 16620 | -7.4967 | 193.175 | 20.4459 | -4.5541 | 148.194 | -0.46688 | 9.7809 | 3.1683 | 12.9492 | 0
[6 empty] | 16635 | -7.5059 | 193.0581 | 20.464 | -4.536 | 148.0373 | -0.46747 | 9.7862 | 3.1719 | 12.9582 | 0
[6 empty] | 16650 | -7.5147 | 192.9411 | 20.4818 | -4.8182 | 147.8812 | -0.46804 | 9.7916 | 3.1756 | 12.9671 | 0
[6 empty] | 16665 | -7.5278 | 192.824 | 20.5039 | -4.7961 | 147.7154 | -0.46859 | 9.7969 | 3.1792 | 12.9761 | 0
[6 empty] | 16680 | -7.5405 | 192.7066 | 20.5256 | -4.7744 | 147.5503 | -0.46941 | 9.8022 | 3.1828 | 12.9851 | 0
[6 empty] | 16695 | -7.5528 | 192.5891 | 20.5469 | -4.7531 | 147.386 | -0.4702 | 9.8075 | 3.1865 | 12.994 | 0
[6 empty] | 16710 | -7.5647 | 192.4713 | 20.5677 | -4.7323 | 147.2224 | -0.47097 | 9.8128 | 3.1902 | 13.003 | 0
[6 empty] | 16725 | -7.5761 | 192.3534 | 20.5881 | -4.7119 | 147.0596 | -0.47171 | 9.8182 | 3.1938 | 13.012 | 0
[6 empty] | 16740 | -7.5871 | 192.2353 | 20.6081 | -4.9919 | 146.8974 | -0.47242 | 9.8235 | 3.1975 | 13.021 | 0
[6 empty] | 16755 | -7.6025 | 192.117 | 20.6325 | -4.9675 | 146.7256 | -0.47311 | 9.8288 | 3.2012 | 13.03 | 0
[6 empty] | 16770 | -7.6173 | 191.9985 | 20.6563 | -4.9437 | 146.5547 | -0.47406 | 9.8341 | 3.2049 | 13.039 | 0
[6 empty] | 16785 | -7.6317 | 191.8798 | 20.6797 | -4.9203 | 146.3845 | -0.47499 | 9.8394 | 3.2086 | 13.048 | 0
[6 empty] | 16800 | -7.6455 | 191.7608 | 20.7025 | -4.8975 | 146.2152 | -0.47588 | 9.8447 | 3.2123 | 13.057 | 0
[6 empty] | 16815 | -7.6589 | 191.6416 | 20.725 | -4.875 | 146.0466 | -0.47675 | 9.8501 | 3.216 | 13.066 | 0
[6 empty] | 16830 | -7.6719 | 191.5222 | 20.747 | -5.153 | 145.8789 | -0.47759 | 9.8554 | 3.2197 | 13.0751 | 0
[6 empty] | 16845 | -7.6891 | 191.4026 | 20.7732 | -5.1268 | 145.7015 | -0.47839 | 9.8607 | 3.2234 | 13.0841 | 0
[6 empty] | 16860 | -7.7058 | 191.2827 | 20.7989 | -5.1011 | 145.5251 | -0.47947 | 9.866 | 3.2272 | 13.0932 | 0
[6 empty] | 16875 | -7.7219 | 191.1626 | 20.8241 | -5.0759 | 145.3495 | -0.48051 | 9.8713 | 3.2309 | 13.1022 | 0
[6 empty] | 16890 | -7.7375 | 191.0422 | 20.8488 | -5.0512 | 145.1749 | -0.48151 | 9.8766 | 3.2346 | 13.1113 | 0
[6 empty] | 16905 | -7.7527 | 190.9216 | 20.873 | -5.027 | 145.001 | -0.48249 | 9.882 | 3.2384 | 13.1203 | 0
[6 empty] | 16920 | -7.7673 | 190.8008 | 20.8967 | -5.3033 | 144.828 | -0.48343 | 9.8873 | 3.2421 | 13.1294 | 0
[6 empty] | 16935 | -7.7861 | 190.6797 | 20.9246 | -5.2754 | 144.6455 | -0.48434 | 9.8926 | 3.2459 | 13.1385 | 0
[6 empty] | 16950 | -7.8044 | 190.5583 | 20.952 | -5.248 | 144.464 | -0.48552 | 9.8979 | 3.2497 | 13.1476 | 0
[6 empty] | 16965 | -7.8221 | 190.4366 | 20.9788 | -5.2212 | 144.2833 | -0.48665 | 9.9032 | 3.2535 | 13.1567 | 0
[6 empty] | 16980 | -7.8392 | 190.3147 | 21.005 | -5.195 | 144.1037 | -0.48776 | 9.9086 | 3.2573 | 13.1658 | 0
[6 empty] | 16995 | -7.8558 | 190.1925 | 21.0307 | -5.4693 | 143.9249 | -0.48883 | 9.9139 | 3.2611 | 13.175 | 0
[6 empty] | 17010 | -7.8765 | 190.07 | 21.0606 | -5.4394 | 143.7367 | -0.48986 | 9.9192 | 3.2649 | 13.1841 | 0
[6 empty] | 17025 | -7.8966 | 189.9472 | 21.0899 | -5.4101 | 143.5495 | -0.49115 | 9.9245 | 3.2687 | 13.1933 | 0
[6 empty] | 17040 | -7.9162 | 189.8241 | 21.1186 | -5.3814 | 143.3633 | -0.49241 | 9.9299 | 3.2725 | 13.2024 | 0
[6 empty] | 17055 | -7.9351 | 189.7007 | 21.1467 | -5.3533 | 143.1781 | -0.49363 | 9.9352 | 3.2764 | 13.2116 | 0
[6 empty] | 17070 | -7.9534 | 189.577 | 21.1742 | -5.6258 | 142.9938 | -0.4948 | 9.9406 | 3.2802 | 13.2208 | 0
[6 empty] | 17085 | -7.9758 | 189.453 | 21.2058 | -5.5942 | 142.8002 | -0.49595 | 9.9459 | 3.2841 | 13.23 | 0
[6 empty] | 17100 | -7.9976 | 189.3287 | 21.2368 | -5.5632 | 142.6077 | -0.49735 | 9.9512 | 3.288 | 13.2392 | 0
[6 empty] | 17115 | -8.0187 | 189.204 | 21.2672 | -5.5328 | 142.4162 | -0.4987 | 9.9566 | 3.2918 | 13.2484 | 0
[6 empty] | 17130 | -8.0392 | 189.079 | 21.2969 | -5.5031 | 142.2258 | -0.50002 | 9.962 | 3.2957 | 13.2577 | 0
[6 empty] | 17145 | -8.0591 | 188.9537 | 21.326 | -5.474 | 142.0365 | -0.5013 | 9.9673 | 3.2996 | 13.2669 | 0
[6 empty] | 17160 | -8.0783 | 188.8281 | 21.3545 | -5.7455 | 141.8481 | -0.50254 | 9.9727 | 3.3035 | 13.2762 | 0
[6 empty] | 17175 | -8.1017 | 188.7021 | 21.3872 | -5.7128 | 141.6503 | -0.50374 | 9.978 | 3.3074 | 13.2855 | 0
[6 empty] | 17190 | -8.1244 | 188.5758 | 21.4191 | -5.6809 | 141.4537 | -0.50519 | 9.9834 | 3.3114 | 13.2948 | 0
[6 empty] | 17205 | -8.1463 | 188.4492 | 21.4504 | -5.6496 | 141.2582 | -0.50661 | 9.9888 | 3.3153 | 13.3041 | 0
[6 empty] | 17220 | -8.1677 | 188.3222 | 21.4811 | -5.6189 | 141.0638 | -0.50798 | 9.9941 | 3.3193 | 13.3134 | 0
[6 empty] | 17235 | -8.1884 | 188.1949 | 21.5111 | -5.8889 | 140.8704 | -0.50931 | 9.9995 | 3.3232 | 13.3228 | 0
[6 empty] | 17250 | -8.2131 | 188.0672 | 21.5452 | -5.8548 | 140.6678 | -0.5106 | 10.0049 | 3.3272 | 13.3321 | 0
[6 empty] | 17265 | -8.2371 | 187.9392 | 21.5786 | -5.8214 | 140.4663 | -0.51214 | 10.0103 | 3.3312 | 13.3415 | 0
[6 empty] | 17280 | -8.2604 | 187.8108 | 21.6113 | -5.7887 | 140.2659 | -0.51364 | 10.0157 | 3.3352 | 13.3509 | 0
[6 empty] | 17295 | -8.283 | 187.682 | 21.6433 | -5.7567 | 140.0667 | -0.51509 | 10.0211 | 3.3392 | 13.3603 | 0
[6 empty] | 17310 | -8.305 | 187.5529 | 21.6747 | -6.0253 | 139.8686 | -0.5165 | 10.0265 | 3.3432 | 13.3697 | 0
[6 empty] | 17325 | -8.331 | 187.4234 | 21.7101 | -5.9899 | 139.6612 | -0.51787 | 10.0319 | 3.3472 | 13.3791 | 0
[6 empty] | 17340 | -8.3562 | 187.2935 | 21.7447 | -5.9553 | 139.4551 | -0.51949 | 10.0373 | 3.3513 | 13.3886 | 0
[6 empty] | 17355 | -8.3807 | 187.1633 | 21.7787 | -5.9213 | 139.2501 | -0.52106 | 10.0427 | 3.3553 | 13.398 | 0
[6 empty] | 17370 | -8.4044 | 187.0326 | 21.8119 | -5.8881 | 139.0464 | -0.52259 | 10.0481 | 3.3594 | 13.4075 | 0
[6 empty] | 17385 | -8.4275 | 186.9016 | 21.8445 | -6.1555 | 138.8437 | -0.52407 | 10.0535 | 3.3635 | 13.417 | 0
[6 empty] | 17400 | -8.4545 | 186.7702 | 21.8811 | -6.1189 | 138.6319 | -0.52551 | 10.0589 | 3.3676 | 13.4265 | 0
[6 empty] | 17415 | -8.4808 | 186.6384 | 21.9169 | -6.0831 | 138.4213 | -0.5272 | 10.0644 | 3.3717 | 13.436 | 0
[6 empty] | 17430 | -8.5063 | 186.5062 | 21.9519 | -6.0481 | 138.212 | -0.52884 | 10.0698 | 3.3758 | 13.4456 | 0
[6 empty] | 17445 | -8.5311 | 186.3736 | 21.9863 | -6.0137 | 138.0038 | -0.53043 | 10.0753 | 3.3799 | 13.4552 | 0
[6 empty] | 17460 | -8.5552 | 186.2406 | 22.0199 | -6.2801 | 137.7969 | -0.53197 | 10.0807 | 3.384 | 13.4647 | 0
[6 empty] | 17475 | -8.5832 | 186.1072 | 22.0575 | -6.2425 | 137.5807 | -0.53347 | 10.0861 | 3.3882 | 13.4743 | 0
[6 empty] | 17490 | -8.6104 | 185.9734 | 22.0943 | -6.2057 | 137.3659 | -0.53522 | 10.0916 | 3.3924 | 13.4839 | 0
[6 empty] | 17505 | -8.6368 | 185.8392 | 22.1304 | -6.1696 | 137.1523 | -0.53692 | 10.0971 | 3.3965 | 13.4936 | 0
[6 empty] | 17520 | -8.6625 | 185.7046 | 22.1657 | -6.4343 | 136.94 | -0.53856 | 10.1025 | 3.4007 | 13.5032 | 0
[6 empty] | 17535 | -8.6921 | 185.5695 | 22.205 | -6.395 | 136.7186 | -0.54016 | 10.108 | 3.4049 | 13.5129 | 0
[6 empty] | 17550 | -8.7208 | 185.434 | 22.2434 | -6.3566 | 136.4985 | -0.54201 | 10.1135 | 3.4091 | 13.5226 | 0
[6 empty] | 17565 | -8.7488 | 185.2981 | 22.2811 | -6.3189 | 136.2797 | -0.5438 | 10.1189 | 3.4134 | 13.5323 | 0
[6 empty] | 17580 | -8.7759 | 185.1617 | 22.3179 | -6.2821 | 136.0623 | -0.54554 | 10.1244 | 3.4176 | 13.542 | 0
[6 empty] | 17595 | -8.8022 | 185.0249 | 22.354 | -6.546 | 135.8461 | -0.54723 | 10.1299 | 3.4219 | 13.5518 | 0
[6 empty] | 17610 | -8.8325 | 184.8877 | 22.394 | -6.506 | 135.6208 | -0.54888 | 10.1354 | 3.4261 | 13.5616 | 0
[6 empty] | 17625 | -8.8619 | 184.75 | 22.4332 | -6.4668 | 135.3969 | -0.55076 | 10.1409 | 3.4304 | 13.5713 | 0
[6 empty] | 17640 | -8.8904 | 184.6118 | 22.4716 | -6.4284 | 135.1743 | -0.5526 | 10.1464 | 3.4347 | 13.5811 | 0
[6 empty] | 17655 | -8.9182 | 184.4732 | 22.5091 | -6.6909 | 134.9531 | -0.55438 | 10.1519 | 3.439 | 13.591 | 0
[6 empty] | 17670 | -8.9498 | 184.3342 | 22.5506 | -6.6494 | 134.7228 | -0.55611 | 10.1575 | 3.4434 | 13.6008 | 0
[6 empty] | 17685 | -8.9805 | 184.1947 | 22.5912 | -6.6088 | 134.494 | -0.55808 | 10.163 | 3.4477 | 13.6107 | 0
[6 empty] | 17700 | -9.0104 | 184.0547 | 22.631 | -6.569 | 134.2666 | -0.56 | 10.1685 | 3.4521 | 13.6206 | 0
[6 empty] | 17715 | -9.0394 | 183.9142 | 22.6699 | -6.5301 | 134.0405 | -0.56186 | 10.1741 | 3.4564 | 13.6305 | 0
[6 empty] | 17730 | -9.0676 | 183.7733 | 22.708 | -6.792 | 133.8158 | -0.56367 | 10.1796 | 3.4608 | 13.6404 | 0
[6 empty] | 17745 | -9.0996 | 183.6319 | 22.75 | -6.75 | 133.582 | -0.56542 | 10.1851 | 3.4652 | 13.6504 | 0
[6 empty] | 17760 | -9.1308 | 183.4901 | 22.7911 | -6.7089 | 133.3497 | -0.56742 | 10.1907 | 3.4696 | 13.6603 | 0
[6 empty] | 17775 | -9.161 | 183.3477 | 22.8313 | -6.6687 | 133.1188 | -0.56936 | 10.1963 | 3.474 | 13.6703 | 0
[6 empty] | 17790 | -9.1904 | 183.2049 | 22.8707 | -6.9293 | 132.8893 | -0.57125 | 10.2018 | 3.4785 | 13.6803 | 0
[6 empty] | 17805 | -9.2237 | 183.0617 | 22.914 | -6.886 | 132.6509 | -0.57308 | 10.2074 | 3.483 | 13.6903 | 0
[6 empty] | 17820 | -9.256 | 182.9179 | 22.9564 | -6.8436 | 132.4139 | -0.57516 | 10.213 | 3.4874 | 13.7004 | 0
[6 empty] | 17835 | -9.2874 | 182.7736 | 22.9978 | -6.8022 | 132.1783 | -0.57717 | 10.2186 | 3.4919 | 13.7105 | 0
[6 empty] | 17850 | -9.3179 | 182.6288 | 23.0384 | -6.9616 | 131.9443 | -0.57913 | 10.2241 | 3.4964 | 13.7206 | 0
[6 empty] | 17865 | -9.3506 | 182.4835 | 23.0813 | -6.9187 | 131.7047 | -0.58103 | 10.2297 | 3.5009 | 13.7307 | 0
[6 empty] | 17880 | -9.3825 | 182.3378 | 23.1233 | -6.8767 | 131.4666 | -0.58307 | 10.2353 | 3.5055 | 13.7408 | 0
[6 empty] | 17895 | -9.4134 | 182.1915 | 23.1644 | -6.8356 | 131.2299 | -0.58506 | 10.2409 | 3.51 | 13.751 | 0
[6 empty] | 17910 | -9.4434 | 182.0448 | 23.2046 | -6.7954 | 130.9947 | -0.58699 | 10.2466 | 3.5146 | 13.7611 | 0
[6 empty] | 17925 | -9.4726 | 181.8975 | 23.244 | -6.756 | 130.7608 | -0.58886 | 10.2522 | 3.5192 | 13.7713 | 0
[6 empty] | 17940 | -9.501 | 181.7499 | 23.2825 | -6.7175 | 130.5283 | -0.59068 | 10.2578 | 3.5238 | 13.7816 | 0
[6 empty] | 17955 | -9.5285 | 181.6018 | 23.3203 | -6.6797 | 130.2971 | -0.59245 | 10.2634 | 3.5284 | 13.7918 | 0
[6 empty] | 17970 | -9.5552 | 181.4532 | 23.3573 | -6.6427 | 130.0672 | -0.59417 | 10.269 | 3.533 | 13.802 | 0
[6 empty] | 17985 | -9.5812 | 181.3043 | 23.3935 | -6.1065 | 129.8386 | -0.59583 | 10.2747 | 3.5376 | 13.8123 | 0
[6 empty] | 18000 | -9.5985 | 181.1549 | 23.4211 | -6.0789 | 129.6285 | -0.59745 | 10.2803 | 3.5423 | 13.8226 | 0
[6 empty] | 18015 | -9.6153 | 181.0053 | 23.4482 | -6.0518 | 129.4193 | -0.59853 | 10.2859 | 3.5469 | 13.8328 | 0
[6 empty] | 18030 | -9.6316 | 180.8554 | 23.4747 | -6.0253 | 129.211 | -0.59958 | 10.2915 | 3.5516 | 13.8431 | 0
[6 empty] | 18045 | -9.6473 | 180.7052 | 23.5007 | -5.9993 | 129.0036 | -0.60059 | 10.2971 | 3.5563 | 13.8534 | 0
[6 empty] | 18060 | -9.6625 | 180.5548 | 23.5262 | -3.9024 | 128.7972 | -0.60157 | 10.3028 | 3.561 | 13.8637 | 0
[6 empty] | 18075 | -9.6448 | 180.4042 | 23.5188 | -3.9098 | 128.6629 | -0.60252 | 10.3084 | 3.5656 | 13.874 | 0
[6 empty] | 18090 | -9.6273 | 180.2538 | 23.5116 | -3.3455 | 128.5283 | -0.60142 | 10.3139 | 3.5703 | 13.8843 | 0
[6 empty] | 18105 | -9.6012 | 180.1038 | 23.4957 | -3.3614 | 128.4132 | -0.60033 | 10.3195 | 3.575 | 13.8945 | 0
[6 empty] | 18120 | -9.5756 | 179.9541 | 23.4803 | -2.8055 | 128.2975 | -0.5987 | 10.325 | 3.5796 | 13.9047 | 0
[6 empty] | 18135 | -9.5415 | 179.8048 | 23.4563 | -2.8294 | 128.201 | -0.5971 | 10.3305 | 3.5843 | 13.9148 | 0
[6 empty] | 18150 | -9.508 | 179.6561 | 23.4329 | -2.2813 | 128.1036 | -0.59498 | 10.336 | 3.5889 | 13.9249 | 0
[6 empty] | 18165 | -9.4663 | 179.5078 | 23.4013 | -2.313 | 128.0251 | -0.59289 | 10.3414 | 3.5935 | 13.935 | 0
[6 empty] | 18180 | -9.4254 | 179.3603 | 23.3704 | -1.7725 | 127.9455 | -0.59029 | 10.3468 | 3.5981 | 13.9449 | 0
[6 empty] | 18195 | -9.3764 | 179.2133 | 23.3313 | -1.8116 | 127.8845 | -0.58774 | 10.3522 | 3.6027 | 13.9549 | 0
[6 empty] | 18210 | -9.3285 | 179.0672 | 23.2932 | -1.8497 | 127.8221 | -0.58468 | 10.3575 | 3.6072 | 13.9647 | 0
[6 empty] | 18225 | -9.2815 | 178.9217 | 23.256 | -1.3154 | 127.7585 | -0.58169 | 10.3628 | 3.6118 | 13.9745 | 0
[6 empty] | 18240 | -9.2266 | 178.777 | 23.2108 | -1.3606 | 127.7132 | -0.57876 | 10.368 | 3.6163 | 13.9843 | 0
[6 empty] | 18255 | -9.1728 | 178.6332 | 23.1667 | -1.4047 | 127.6664 | -0.57534 | 10.3732 | 3.6207 | 13.994 | 0
[6 empty] | 18270 | -9.1202 | 178.4902 | 23.1237 | -1.4477 | 127.618 | -0.57198 | 10.3784 | 3.6252 | 14.0036 | 0
[6 empty] | 18285 | -9.0686 | 178.348 | 23.0817 | -0.91826 | 127.5682 | -0.5687 | 10.3835 | 3.6296 | 14.0131 | 0
[6 empty] | 18300 | -9.0093 | 178.2067 | 23.0318 | -0.96815 | 127.5366 | -0.56549 | 10.3886 | 3.634 | 14.0226 | 0
[6 empty] | 18315 | -8.9512 | 178.0662 | 22.9832 | -1.0168 | 127.5033 | -0.56179 | 10.3936 | 3.6384 | 14.032 | 0
[6 empty] | 18330 | -8.8943 | 177.9267 | 22.9356 | -1.0644 | 127.4683 | -0.55817 | 10.3986 | 3.6427 | 14.0413 | 0
[6 empty] | 18345 | -8.8387 | 177.788 | 22.8893 | -1.1107 | 127.4317 | -0.55462 | 10.4035 | 3.647 | 14.0505 | 0
[6 empty] | 18360 | -8.7843 | 177.6502 | 22.844 | -0.58457 | 127.3934 | -0.55115 | 10.4084 | 3.6513 | 14.0597 | 0
[6 empty] | 18375 | -8.7221 | 177.5133 | 22.7909 | -0.63767 | 127.3733 | -0.54776 | 10.4132 | 3.6556 | 14.0688 | 0
[6 empty] | 18390 | -8.6612 | 177.3773 | 22.7391 | -0.6895 | 127.3514 | -0.54388 | 10.4181 | 3.6598 | 14.0779 | 0
[6 empty] | 18405 | -8.6016 | 177.2423 | 22.6885 | -0.74009 | 127.3276 | -0.54008 | 10.4228 | 3.664 | 14.0868 | 0
[6 empty] | 18420 | -8.5434 | 177.1082 | 22.6391 | -0.78946 | 127.3022 | -0.53637 | 10.4275 | 3.6682 | 14.0957 | 0
[6 empty] | 18435 | -8.4864 | 176.975 | 22.5909 | -0.83765 | 127.275 | -0.53274 | 10.4322 | 3.6723 | 14.1046 | 0
[6 empty] | 18450 | -8.4306 | 176.8427 | 22.5439 | -0.88468 | 127.2462 | -0.52918 | 10.4369 | 3.6765 | 14.1133 | 0
[6 empty] | 18465 | -8.376 | 176.7113 | 22.498 | -0.93058 | 127.2157 | -0.5257 | 10.4414 | 3.6805 | 14.122 | 0
[6 empty] | 18480 | -8.3226 | 176.5807 | 22.4532 | -0.40395 | 127.1837 | -0.5223 | 10.446 | 3.6846 | 14.1306 | 0
[6 empty] | 18495 | -8.2614 | 176.451 | 22.4005 | -0.4566 | 127.1698 | -0.51897 | 10.4505 | 3.6886 | 14.1392 | 0
[6 empty] | 18510 | -8.2015 | 176.3222 | 22.3491 | -0.508 | 127.1541 | -0.51515 | 10.455 | 3.6926 | 14.1476 | 0
[6 empty] | 18525 | -8.1429 | 176.1943 | 22.299 | -0.55817 | 127.1366 | -0.51142 | 10.4594 | 3.6966 | 14.1561 | 0
[6 empty] | 18540 | -8.0856 | 176.0674 | 22.25 | -0.60714 | 127.1174 | -0.50777 | 10.4638 | 3.7006 | 14.1644 | 0
[6 empty] | 18555 | -8.0295 | 175.9414 | 22.2022 | -0.65494 | 127.0965 | -0.50419 | 10.4682 | 3.7045 | 14.1727 | 0
[6 empty] | 18570 | -7.9747 | 175.8162 | 22.1556 | -0.70159 | 127.074 | -0.5007 | 10.4725 | 3.7084 | 14.1809 | 0
[6 empty] | 18585 | -7.921 | 175.6919 | 22.11 | -0.74712 | 127.0498 | -0.49727 | 10.4768 | 3.7123 | 14.189 | 0
[6 empty] | 18600 | -7.8685 | 175.5684 | 22.0656 | -0.79156 | 127.0241 | -0.49393 | 10.481 | 3.7161 | 14.1971 | 0
[6 empty] | 18615 | -7.8171 | 175.4457 | 22.0222 | -0.26351 | 126.9969 | -0.49065 | 10.4852 | 3.7199 | 14.2051 | 0
[6 empty] | 18630 | -7.7579 | 175.3239 | 21.9709 | -0.31478 | 126.9878 | -0.48745 | 10.4893 | 3.7237 | 14.2131 | 0
[6 empty] | 18645 | -7.6999 | 175.2029 | 21.9209 | -0.36483 | 126.977 | -0.48376 | 10.4935 | 3.7275 | 14.2209 | 0
[6 empty] | 18660 | -7.6433 | 175.0829 | 21.872 | -0.41368 | 126.9644 | -0.48014 | 10.4975 | 3.7312 | 14.2288 | 0
[6 empty] | 18675 | -7.6067 | 175 | 21.8408 | -0.44488 | 126.9502 | -0.33153 | 10.5004 | 3.7338 | 14.2341 | 7.6067
[6 empty] | 18690 | -7.6136 | 175 | 21.8478 | -0.43792 | 126.9349 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6136
[6 empty] | 18705 | -7.6205 | 175 | 21.8546 | -0.43107 | 126.9198 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6205
[6 empty] | 18720 | -7.6272 | 175 | 21.8614 | -0.42433 | 126.905 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6272
[6 empty] | 18735 | -7.6339 | 175 | 21.868 | -0.41769 | 126.8904 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6339
[6 empty] | 18750 | -7.6404 | 175 | 21.8746 | -0.41116 | 126.876 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6404
[6 empty] | 18765 | -7.6468 | 175 | 21.881 | -0.40473 | 126.8618 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6468
[6 empty] | 18780 | -7.6532 | 175 | 21.8873 | -0.3984 | 126.8479 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6532
[6 empty] | 18795 | -7.6594 | 175 | 21.8936 | -0.39216 | 126.8342 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6594
[6 empty] | 18810 | -7.6655 | 175 | 21.8997 | -0.38603 | 126.8207 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6655
[6 empty] | 18825 | -7.6716 | 175 | 21.9057 | -0.37999 | 126.8074 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6716
[6 empty] | 18840 | -7.6775 | 175 | 21.9117 | 0.19738 | 126.7943 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6775
[6 empty] | 18855 | -7.6744 | 175 | 21.9086 | -0.37713 | 126.8011 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6744
[6 empty] | 18870 | -7.6803 | 175 | 21.9145 | 0.20019 | 126.7881 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6803
[6 empty] | 18885 | -7.6772 | 175 | 21.9113 | 0.19706 | 126.795 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6772
[6 empty] | 18900 | -7.6741 | 175 | 21.9083 | -0.37745 | 126.8018 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6741
[6 empty] | 18915 | -7.68 | 175 | 21.9142 | 0.19988 | 126.7888 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.68
[6 empty] | 18930 | -7.6769 | 175 | 21.911 | 0.19676 | 126.7957 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6769
[6 empty] | 18945 | -7.6738 | 175 | 21.908 | -0.37775 | 126.8025 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6738
[6 empty] | 18960 | -7.6797 | 175 | 21.9139 | 0.19959 | 126.7895 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6797
[6 empty] | 18975 | -7.6766 | 175 | 21.9108 | 0.19647 | 126.7963 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6766
[6 empty] | 18990 | -7.6735 | 175 | 21.9077 | -0.37804 | 126.8031 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6735
[6 empty] | 19005 | -7.6794 | 175 | 21.9136 | 0.19931 | 126.7901 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6794
[6 empty] | 19020 | -7.6763 | 175 | 21.9105 | 0.19619 | 126.797 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6763
[6 empty] | 19035 | -7.6733 | 175 | 21.9074 | -0.37831 | 126.8037 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6733
[6 empty] | 19050 | -7.6792 | 175 | 21.9133 | 0.19904 | 126.7907 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6792
[6 empty] | 19065 | -7.6761 | 175 | 21.9102 | 0.19592 | 126.7975 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6761
[6 empty] | 19080 | -7.673 | 175 | 21.9071 | -0.37857 | 126.8043 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.673
[6 empty] | 19095 | -7.6789 | 175 | 21.9131 | 0.19878 | 126.7913 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6789
[6 empty] | 19110 | -7.6758 | 175 | 21.91 | 0.19567 | 126.7981 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6758
[6 empty] | 19125 | -7.6727 | 175 | 21.9069 | -0.37882 | 126.8048 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6727
[6 empty] | 19140 | -7.6787 | 175 | 21.9128 | 0.19854 | 126.7918 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6787
[6 empty] | 19155 | -7.6756 | 175 | 21.9097 | 0.19543 | 126.7986 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6756
[6 empty] | 19170 | -7.6725 | 175 | 21.9067 | -0.37905 | 126.8053 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6725
[6 empty] | 19185 | -7.6784 | 175 | 21.9126 | 0.1983 | 126.7923 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6784
[6 empty] | 19200 | -7.6753 | 175 | 21.9095 | 0.1952 | 126.7991 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6753
[6 empty] | 19215 | -7.6723 | 175 | 21.9064 | -0.37928 | 126.8058 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6723
[6 empty] | 19230 | -7.6782 | 175 | 21.9124 | 0.19808 | 126.7928 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6782
[6 empty] | 19245 | -7.6751 | 175 | 21.9093 | 0.19498 | 126.7996 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6751
[6 empty] | 19260 | -7.6721 | 175 | 21.9062 | -0.3795 | 126.8063 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6721
[6 empty] | 19275 | -7.678 | 175 | 21.9122 | 0.19787 | 126.7933 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.678
[6 empty] | 19290 | -7.6749 | 175 | 21.9091 | -0.37665 | 126.8001 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6749
[6 empty] | 19305 | -7.6808 | 175 | 21.915 | 0.20067 | 126.7871 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6808
[6 empty] | 19320 | -7.6777 | 175 | 21.9118 | 0.19753 | 126.794 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6777
[6 empty] | 19335 | -7.6746 | 175 | 21.9087 | -0.37699 | 126.8008 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6746
[6 empty] | 19350 | -7.6805 | 175 | 21.9146 | 0.20033 | 126.7878 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6805
[6 empty] | 19365 | -7.6773 | 175 | 21.9115 | 0.1972 | 126.7947 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6773
[6 empty] | 19380 | -7.6743 | 175 | 21.9084 | -0.37731 | 126.8015 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6743
[6 empty] | 19395 | -7.6802 | 175 | 21.9143 | 0.20002 | 126.7885 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6802
[6 empty] | 19410 | -7.677 | 175 | 21.9112 | 0.19689 | 126.7954 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.677
[6 empty] | 19425 | -7.6739 | 175 | 21.9081 | -0.37762 | 126.8022 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6739
[6 empty] | 19440 | -7.6799 | 175 | 21.914 | 0.19972 | 126.7892 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6799
[6 empty] | 19455 | -7.6767 | 175 | 21.9109 | 0.19659 | 126.7961 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6767
[6 empty] | 19470 | -7.6737 | 175 | 21.9078 | -0.37791 | 126.8028 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6737
[6 empty] | 19485 | -7.6796 | 175 | 21.9137 | 0.19943 | 126.7898 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6796
[6 empty] | 19500 | -7.6764 | 175 | 21.9106 | 0.19631 | 126.7967 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6764
[6 empty] | 19515 | -7.6734 | 175 | 21.9075 | -0.37819 | 126.8034 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6734
[6 empty] | 19530 | -7.6793 | 175 | 21.9134 | 0.19915 | 126.7904 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6793
[6 empty] | 19545 | -7.6762 | 175 | 21.9103 | 0.19604 | 126.7973 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6762
[6 empty] | 19560 | -7.6731 | 175 | 21.9073 | -0.37846 | 126.804 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6731
[6 empty] | 19575 | -7.679 | 175 | 21.9132 | 0.19889 | 126.791 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.679
[6 empty] | 19590 | -7.6759 | 175 | 21.9101 | 0.19578 | 126.7979 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6759
[6 empty] | 19605 | -7.6729 | 175 | 21.907 | -0.37871 | 126.8046 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6729
[6 empty] | 19620 | -7.6788 | 175 | 21.9129 | 0.19864 | 126.7916 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6788
[6 empty] | 19635 | -7.6757 | 175 | 21.9098 | 0.19554 | 126.7984 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6757
[6 empty] | 19650 | -7.6726 | 175 | 21.9068 | -0.37895 | 126.8051 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6726
[6 empty] | 19665 | -7.6785 | 175 | 21.9127 | 0.1984 | 126.7921 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6785
[6 empty] | 19680 | -7.6754 | 175 | 21.9096 | 0.1953 | 126.7989 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6754
[6 empty] | 19695 | -7.6724 | 175 | 21.9065 | -0.37918 | 126.8056 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6724
[6 empty] | 19710 | -7.6783 | 175 | 21.9125 | 0.19818 | 126.7926 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6783
[6 empty] | 19725 | -7.6752 | 175 | 21.9094 | 0.19508 | 126.7994 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6752
[6 empty] | 19740 | -7.6722 | 175 | 21.9063 | -0.3794 | 126.8061 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6722
[6 empty] | 19755 | -7.6781 | 175 | 21.9122 | 0.19796 | 126.7931 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6781
[6 empty] | 19770 | -7.675 | 175 | 21.9091 | 0.19486 | 126.7999 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.675
[6 empty] | 19785 | -7.672 | 175 | 21.9061 | -0.37961 | 126.8066 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.672
[6 empty] | 19800 | -7.6779 | 175 | 21.912 | 0.19775 | 126.7935 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6779
[6 empty] | 19815 | -7.6748 | 175 | 21.9089 | -0.37677 | 126.8003 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6748
[6 empty] | 19830 | -7.6807 | 175 | 21.9148 | 0.20055 | 126.7874 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6807
[6 empty] | 19845 | -7.6776 | 175 | 21.9117 | 0.19742 | 126.7943 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6776
[6 empty] | 19860 | -7.6745 | 175 | 21.9086 | -0.3771 | 126.801 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6745
[6 empty] | 19875 | -7.6804 | 175 | 21.9145 | 0.20023 | 126.7881 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6804
[6 empty] | 19890 | -7.6772 | 175 | 21.9114 | 0.1971 | 126.795 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6772
[6 empty] | 19905 | -7.6741 | 175 | 21.9083 | -0.37742 | 126.8017 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6741
[6 empty] | 19920 | -7.6801 | 175 | 21.9142 | 0.19992 | 126.7888 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6801
[6 empty] | 19935 | -7.6769 | 175 | 21.9111 | 0.19679 | 126.7956 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6769
[6 empty] | 19950 | -7.6738 | 175 | 21.908 | -0.37772 | 126.8024 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6738
[6 empty] | 19965 | -7.6798 | 175 | 21.9139 | 0.19962 | 126.7894 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6798
[6 empty] | 19980 | -7.6766 | 175 | 21.9108 | 0.1965 | 126.7963 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6766
[6 empty] | 19995 | -7.6736 | 175 | 21.9077 | -0.37801 | 126.803 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6736
[6 empty] | 20010 | -7.6795 | 175 | 21.9136 | 0.19934 | 126.79 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6795
[6 empty] | 20025 | -7.6764 | 175 | 21.9105 | 0.19622 | 126.7969 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6764
[6 empty] | 20040 | -7.6733 | 175 | 21.9074 | -0.37828 | 126.8036 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6733
[6 empty] | 20055 | -7.6792 | 175 | 21.9134 | 0.19907 | 126.7906 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6792
[6 empty] | 20070 | -7.6761 | 175 | 21.9102 | 0.19595 | 126.7975 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6761
[6 empty] | 20085 | -7.673 | 175 | 21.9072 | -0.37854 | 126.8042 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.673
[6 empty] | 20100 | -7.6789 | 175 | 21.9131 | 0.19881 | 126.7912 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6789
[6 empty] | 20115 | -7.6758 | 175 | 21.91 | 0.1957 | 126.798 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6758
[6 empty] | 20130 | -7.6728 | 175 | 21.9069 | -0.37879 | 126.8048 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6728
[6 empty] | 20145 | -7.6787 | 175 | 21.9128 | 0.19856 | 126.7917 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6787
[6 empty] | 20160 | -7.6756 | 175 | 21.9097 | 0.19546 | 126.7986 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6756
[6 empty] | 20175 | -7.6725 | 175 | 21.9067 | -0.37903 | 126.8053 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6725
[6 empty] | 20190 | -7.6785 | 175 | 21.9126 | 0.19833 | 126.7923 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6785
[6 empty] | 20205 | -7.6754 | 175 | 21.9095 | 0.19523 | 126.7991 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6754
[6 empty] | 20220 | -7.6723 | 175 | 21.9065 | -0.37926 | 126.8058 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6723
[6 empty] | 20235 | -7.6782 | 175 | 21.9124 | 0.1981 | 126.7927 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6782
[6 empty] | 20250 | -7.6751 | 175 | 21.9093 | 0.19501 | 126.7996 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6751
[6 empty] | 20265 | -7.6721 | 175 | 21.9062 | -0.37947 | 126.8063 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6721
[6 empty] | 20280 | -7.678 | 175 | 21.9122 | 0.19789 | 126.7932 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.678
[6 empty] | 20295 | -7.6749 | 175 | 21.9091 | -0.37663 | 126.8 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6749
[6 empty] | 20310 | -7.6808 | 175 | 21.915 | 0.20069 | 126.7871 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6808
[6 empty] | 20325 | -7.6777 | 175 | 21.9118 | 0.19755 | 126.794 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6777
[6 empty] | 20340 | -7.6746 | 175 | 21.9087 | -0.37697 | 126.8008 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6746
[6 empty] | 20355 | -7.6805 | 175 | 21.9146 | 0.20035 | 126.7878 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6805
[6 empty] | 20370 | -7.6774 | 175 | 21.9115 | 0.19722 | 126.7947 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6774
[6 empty] | 20385 | -7.6743 | 175 | 21.9084 | -0.37729 | 126.8015 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6743
[6 empty] | 20400 | -7.6802 | 175 | 21.9143 | 0.20004 | 126.7885 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6802
[6 empty] | 20415 | -7.677 | 175 | 21.9112 | 0.19691 | 126.7954 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.677
[6 empty] | 20430 | -7.674 | 175 | 21.9081 | -0.3776 | 126.8022 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.674
[6 empty] | 20445 | -7.6799 | 175 | 21.914 | 0.19973 | 126.7892 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6799
[6 empty] | 20460 | -7.6767 | 175 | 21.9109 | 0.19661 | 126.796 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6767
[6 empty] | 20475 | -7.6737 | 175 | 21.9078 | -0.37789 | 126.8028 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6737
[6 empty] | 20490 | -7.6796 | 175 | 21.9137 | 0.19945 | 126.7898 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6796
[6 empty] | 20505 | -7.6765 | 175 | 21.9106 | 0.19633 | 126.7967 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6765
[6 empty] | 20520 | -7.6734 | 175 | 21.9075 | -0.37817 | 126.8034 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6734
[6 empty] | 20535 | -7.6793 | 175 | 21.9135 | 0.19917 | 126.7904 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6793
[6 empty] | 20550 | -7.6762 | 175 | 21.9103 | 0.19605 | 126.7973 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6762
[6 empty] | 20565 | -7.6731 | 175 | 21.9073 | -0.37844 | 126.804 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6731
[6 empty] | 20580 | -7.679 | 175 | 21.9132 | 0.19891 | 126.791 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.679
[6 empty] | 20595 | -7.6759 | 175 | 21.9101 | 0.1958 | 126.7978 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6759
[6 empty] | 20610 | -7.6729 | 175 | 21.907 | -0.3787 | 126.8046 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6729
[6 empty] | 20625 | -7.6788 | 175 | 21.9129 | 0.19866 | 126.7915 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6788
[6 empty] | 20640 | -7.6757 | 175 | 21.9098 | 0.19555 | 126.7984 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6757
[6 empty] | 20655 | -7.6726 | 175 | 21.9068 | -0.37894 | 126.8051 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6726
[6 empty] | 20670 | -7.6786 | 175 | 21.9127 | 0.19842 | 126.7921 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6786
[6 empty] | 20685 | -7.6755 | 175 | 21.9096 | 0.19531 | 126.7989 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6755
[6 empty] | 20700 | -7.6724 | 175 | 21.9065 | -0.37917 | 126.8056 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6724
[6 empty] | 20715 | -7.6783 | 175 | 21.9125 | 0.19819 | 126.7926 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6783
[6 empty] | 20730 | -7.6752 | 175 | 21.9094 | 0.19509 | 126.7994 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6752
[6 empty] | 20745 | -7.6722 | 175 | 21.9063 | -0.37939 | 126.8061 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6722
[6 empty] | 20760 | -7.6781 | 175 | 21.9123 | 0.19797 | 126.793 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6781
[6 empty] | 20775 | -7.675 | 175 | 21.9092 | 0.19488 | 126.7998 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.675
[6 empty] | 20790 | -7.672 | 175 | 21.9061 | -0.3796 | 126.8065 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.672
[6 empty] | 20805 | -7.6779 | 175 | 21.9121 | 0.19777 | 126.7935 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6779
[6 empty] | 20820 | -7.6748 | 175 | 21.909 | -0.37676 | 126.8003 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6748
[6 empty] | 20835 | -7.6807 | 175 | 21.9149 | 0.20057 | 126.7873 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6807
[6 empty] | 20850 | -7.6776 | 175 | 21.9117 | 0.19743 | 126.7942 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6776
[6 empty] | 20865 | -7.6745 | 175 | 21.9086 | -0.37709 | 126.801 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6745
[6 empty] | 20880 | -7.6804 | 175 | 21.9145 | 0.20024 | 126.788 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6804
[6 empty] | 20895 | -7.6772 | 175 | 21.9114 | 0.19711 | 126.7949 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6772
[6 empty] | 20910 | -7.6742 | 175 | 21.9083 | -0.37741 | 126.8017 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6742
[6 empty] | 20925 | -7.6801 | 175 | 21.9142 | 0.19993 | 126.7887 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6801
[6 empty] | 20940 | -7.6769 | 175 | 21.9111 | 0.1968 | 126.7956 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6769
[6 empty] | 20955 | -7.6739 | 175 | 21.908 | -0.37771 | 126.8024 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6739
[6 empty] | 20970 | -7.6798 | 175 | 21.9139 | 0.19963 | 126.7894 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6798
[6 empty] | 20985 | -7.6766 | 175 | 21.9108 | 0.19651 | 126.7963 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6766
[6 empty] | 21000 | -7.6736 | 175 | 21.9077 | -0.378 | 126.803 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6736
[6 empty] | 21015 | -7.6795 | 175 | 21.9136 | 0.19935 | 126.79 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6795
[6 empty] | 21030 | -7.6764 | 175 | 21.9105 | 0.19623 | 126.7969 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6764
[6 empty] | 21045 | -7.6733 | 175 | 21.9074 | -0.37827 | 126.8036 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6733
[6 empty] | 21060 | -7.6792 | 175 | 21.9134 | 0.19908 | 126.7906 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6792
[6 empty] | 21075 | -7.6761 | 175 | 21.9102 | 0.19596 | 126.7975 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6761
[6 empty] | 21090 | -7.673 | 175 | 21.9072 | -0.37853 | 126.8042 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.673
[6 empty] | 21105 | -7.679 | 175 | 21.9131 | 0.19882 | 126.7912 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.679
[6 empty] | 21120 | -7.6758 | 175 | 21.91 | 0.19571 | 126.798 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6758
[6 empty] | 21135 | -7.6728 | 175 | 21.9069 | -0.37878 | 126.8048 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6728
[6 empty] | 21150 | -7.6787 | 175 | 21.9129 | 0.19857 | 126.7917 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6787
[6 empty] | 21165 | -7.6756 | 175 | 21.9098 | 0.19546 | 126.7985 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6756
[6 empty] | 21180 | -7.6725 | 175 | 21.9067 | -0.37902 | 126.8053 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6725
[6 empty] | 21195 | -7.6785 | 175 | 21.9126 | 0.19834 | 126.7922 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6785
[6 empty] | 21210 | -7.6754 | 175 | 21.9095 | 0.19523 | 126.7991 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6754
[6 empty] | 21225 | -7.6723 | 175 | 21.9065 | -0.37925 | 126.8058 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6723
[6 empty] | 21240 | -7.6782 | 175 | 21.9124 | 0.19811 | 126.7927 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6782
[6 empty] | 21255 | -7.6751 | 175 | 21.9093 | 0.19501 | 126.7995 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6751
[6 empty] | 21270 | -7.6721 | 175 | 21.9062 | -0.37947 | 126.8063 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6721
[6 empty] | 21285 | -7.678 | 175 | 21.9122 | 0.1979 | 126.7932 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.678
[6 empty] | 21300 | -7.6749 | 175 | 21.9091 | -0.37663 | 126.8 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6749
[6 empty] | 21315 | -7.6808 | 175 | 21.915 | 0.20069 | 126.787 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6808
[6 empty] | 21330 | -7.6777 | 175 | 21.9118 | 0.19755 | 126.794 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6777
[6 empty] | 21345 | -7.6746 | 175 | 21.9088 | -0.37696 | 126.8007 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6746
[6 empty] | 21360 | -7.6805 | 175 | 21.9146 | 0.20036 | 126.7878 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6805
[6 empty] | 21375 | -7.6774 | 175 | 21.9115 | 0.19723 | 126.7947 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6774
[6 empty] | 21390 | -7.6743 | 175 | 21.9084 | -0.37729 | 126.8015 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6743
[6 empty] | 21405 | -7.6802 | 175 | 21.9143 | 0.20004 | 126.7885 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6802
[6 empty] | 21420 | -7.6771 | 175 | 21.9112 | 0.19691 | 126.7954 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6771
[6 empty] | 21435 | -7.674 | 175 | 21.9081 | -0.37759 | 126.8021 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.674
[6 empty] | 21450 | -7.6799 | 175 | 21.914 | 0.19974 | 126.7891 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6799
[6 empty] | 21465 | -7.6768 | 175 | 21.9109 | 0.19662 | 126.796 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6768
[6 empty] | 21480 | -7.6737 | 175 | 21.9078 | -0.37789 | 126.8028 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6737
[6 empty] | 21495 | -7.6796 | 175 | 21.9137 | 0.19945 | 126.7898 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6796
[6 empty] | 21510 | -7.6765 | 175 | 21.9106 | 0.19633 | 126.7966 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6765
[6 empty] | 21525 | -7.6734 | 175 | 21.9075 | -0.37817 | 126.8034 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6734
[6 empty] | 21540 | -7.6793 | 175 | 21.9135 | 0.19918 | 126.7904 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6793
[6 empty] | 21555 | -7.6762 | 175 | 21.9103 | 0.19606 | 126.7972 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6762
[6 empty] | 21570 | -7.6731 | 175 | 21.9073 | -0.37844 | 126.804 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6731
[6 empty] | 21585 | -7.6791 | 175 | 21.9132 | 0.19891 | 126.791 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6791
[6 empty] | 21600 | -7.6759 | 175 | 21.9101 | 0.1958 | 126.7978 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6759
[6 empty] | 21615 | -7.6729 | 175 | 21.907 | -0.37869 | 126.8045 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6729
[6 empty] | 21630 | -7.6788 | 175 | 21.9129 | 0.19866 | 126.7915 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6788
[6 empty] | 21645 | -7.6757 | 175 | 21.9098 | 0.19556 | 126.7984 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6757
[6 empty] | 21660 | -7.6726 | 175 | 21.9068 | -0.37893 | 126.8051 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6726
[6 empty] | 21675 | -7.6786 | 175 | 21.9127 | 0.19842 | 126.792 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6786
[6 empty] | 21690 | -7.6755 | 175 | 21.9096 | 0.19532 | 126.7989 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6755
[6 empty] | 21705 | -7.6724 | 175 | 21.9065 | -0.37916 | 126.8056 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6724
[6 empty] | 21720 | -7.6783 | 175 | 21.9125 | 0.1982 | 126.7925 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6783
[6 empty] | 21735 | -7.6752 | 175 | 21.9094 | 0.1951 | 126.7994 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6752
[6 empty] | 21750 | -7.6722 | 175 | 21.9063 | -0.37939 | 126.8061 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6722
[6 empty] | 21765 | -7.6781 | 175 | 21.9123 | 0.19798 | 126.793 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6781
[6 empty] | 21780 | -7.675 | 175 | 21.9092 | 0.19488 | 126.7998 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.675
[6 empty] | 21795 | -7.672 | 175 | 21.9061 | -0.3796 | 126.8065 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.672
[6 empty] | 21810 | -7.6779 | 175 | 21.9121 | 0.19777 | 126.7935 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6779
[6 empty] | 21825 | -7.6748 | 175 | 21.909 | -0.37675 | 126.8003 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6748
[6 empty] | 21840 | -7.6807 | 175 | 21.9149 | 0.20057 | 126.7873 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6807
[6 empty] | 21855 | -7.6776 | 175 | 21.9117 | 0.19743 | 126.7942 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6776
[6 empty] | 21870 | -7.6745 | 175 | 21.9086 | -0.37708 | 126.801 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6745
[6 empty] | 21885 | -7.6804 | 175 | 21.9145 | 0.20024 | 126.788 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6804
[6 empty] | 21900 | -7.6772 | 175 | 21.9114 | 0.19711 | 126.7949 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6772
[6 empty] | 21915 | -7.6742 | 175 | 21.9083 | -0.3774 | 126.8017 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6742
[6 empty] | 21930 | -7.6801 | 175 | 21.9142 | 0.19993 | 126.7887 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6801
[6 empty] | 21945 | -7.6769 | 175 | 21.9111 | 0.1968 | 126.7956 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6769
[6 empty] | 21960 | -7.6739 | 175 | 21.908 | -0.3777 | 126.8024 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6739
[6 empty] | 21975 | -7.6798 | 175 | 21.9139 | 0.19963 | 126.7894 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6798
[6 empty] | 21990 | -7.6766 | 175 | 21.9108 | 0.19651 | 126.7962 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6766
[6 empty] | 22005 | -7.6736 | 175 | 21.9077 | -0.37799 | 126.803 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6736
[6 empty] | 22020 | -7.6795 | 175 | 21.9136 | 0.19935 | 126.79 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6795
[6 empty] | 22035 | -7.6764 | 175 | 21.9105 | 0.19623 | 126.7969 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6764
[6 empty] | 22050 | -7.6733 | 175 | 21.9074 | -0.37827 | 126.8036 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6733
[6 empty] | 22065 | -7.6792 | 175 | 21.9134 | 0.19908 | 126.7906 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6792
[6 empty] | 22080 | -7.6761 | 175 | 21.9103 | 0.19596 | 126.7974 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6761
[6 empty] | 22095 | -7.673 | 175 | 21.9072 | -0.37853 | 126.8042 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.673
[6 empty] | 22110 | -7.679 | 175 | 21.9131 | 0.19882 | 126.7912 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.679
[6 empty] | 22125 | -7.6758 | 175 | 21.91 | 0.19571 | 126.798 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6758
[6 empty] | 22140 | -7.6728 | 175 | 21.9069 | -0.37878 | 126.8047 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6728
[6 empty] | 22155 | -7.6787 | 175 | 21.9129 | 0.19857 | 126.7917 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6787
[6 empty] | 22170 | -7.6756 | 175 | 21.9098 | 0.19547 | 126.7985 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6756
[6 empty] | 22185 | -7.6725 | 175 | 21.9067 | -0.37902 | 126.8053 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6725
[6 empty] | 22200 | -7.6785 | 175 | 21.9126 | 0.19834 | 126.7922 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6785
[6 empty] | 22215 | -7.6754 | 175 | 21.9095 | 0.19524 | 126.7991 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6754
[6 empty] | 22230 | -7.6723 | 175 | 21.9065 | -0.37925 | 126.8058 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6723
[6 empty] | 22245 | -7.6783 | 175 | 21.9124 | 0.19811 | 126.7927 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6783
[6 empty] | 22260 | -7.6752 | 175 | 21.9093 | 0.19502 | 126.7995 | 0 | 10.5004 | 3.7338 | 14.2341 | 7.6752
174
[8 empty] | 173 | [6 empty] | 14.3631
[8 empty] | 172 | [6 empty] | 14.4267
[8 empty] | 171 | [6 empty] | 14.4897
[8 empty] | 170 | [6 empty] | 14.5522
[8 empty] | 169 | [6 empty] | 14.6142
[8 empty] | 168 | [6 empty] | 14.6755
[8 empty] | 167 | [6 empty] | 14.7364
[8 empty] | 166 | [6 empty] | 14.7967
[8 empty] | 165 | [6 empty] | 14.8566
[8 empty] | 164 | [6 empty] | 14.9158
[8 empty] | 163 | [6 empty] | 14.9746
[8 empty] | 162 | [6 empty] | 15.0329
[8 empty] | 161 | [6 empty] | 15.0907
[8 empty] | 160 | [6 empty] | 15.148
[8 empty] | 159 | [6 empty] | 15.2048
[8 empty] | 158 | [6 empty] | 15.2611
[8 empty] | 157 | [6 empty] | 15.3169
[8 empty] | 156 | [6 empty] | 15.3723
[8 empty] | 155 | [6 empty] | 15.4272
[8 empty] | 154 | [6 empty] | 15.4817
[8 empty] | 153 | [6 empty] | 15.5357
[8 empty] | 152 | [6 empty] | 15.5893
[8 empty] | 151 | [6 empty] | 15.6424
[8 empty] | 150 | [6 empty] | 15.6951
[8 empty] | 149 | [6 empty] | 15.7473
[8 empty] | 148 | [6 empty] | 15.7992
[8 empty] | 147 | [6 empty] | 15.8506
[8 empty] | 146 | [6 empty] | 15.9016
[8 empty] | 145 | [6 empty] | 15.9522
[8 empty] | 144 | [6 empty] | 16.0025
[8 empty] | 143 | [6 empty] | 16.0523
[8 empty] | 142 | [6 empty] | 16.1017
[8 empty] | 141 | [6 empty] | 16.1508
[8 empty] | 140 | [6 empty] | 16.1994
[8 empty] | 139 | [6 empty] | 16.2477
[8 empty] | 138 | [6 empty] | 16.2957
[8 empty] | 137 | [6 empty] | 16.3432
[8 empty] | 136 | [6 empty] | 16.3904
[8 empty] | 135 | [6 empty] | 16.4373
[8 empty] | 134 | [6 empty] | 16.4838
[8 empty] | 133 | [6 empty] | 16.53
[8 empty] | 132 | [6 empty] | 16.5758
[8 empty] | 131 | [6 empty] | 16.6213
[8 empty] | 130 | [6 empty] | 16.6665
[8 empty] | 129 | [6 empty] | 16.7114
[8 empty] | 128 | [6 empty] | 16.7559
[8 empty] | 127 | [6 empty] | 16.8002
[8 empty] | 126 | [6 empty] | 16.8441
[8 empty] | 125 | [6 empty] | 16.8877
[8 empty] | 124 | [6 empty] | 16.9311
[8 empty] | 123 | [6 empty] | 16.9741
[8 empty] | 122 | [6 empty] | 17.0168
[8 empty] | 121 | [6 empty] | 17.0593
[8 empty] | 120 | [6 empty] | 17.1015
[8 empty] | 119 | [6 empty] | 17.1434
[8 empty] | 118 | [6 empty] | 17.1851
[8 empty] | 117 | [6 empty] | 17.2265
[8 empty] | 116 | [6 empty] | 17.2676
[8 empty] | 115 | [6 empty] | 17.3084
[8 empty] | 114 | [6 empty] | 17.3491
[8 empty] | 113 | [6 empty] | 17.3894
[8 empty] | 112 | [6 empty] | 17.4296
[8 empty] | 111 | [6 empty] | 17.4694
[8 empty] | 110 | [6 empty] | 17.5091
[8 empty] | 109 | [6 empty] | 17.5485
[8 empty] | 108 | [6 empty] | 17.5877
[8 empty] | 107 | [6 empty] | 17.6267
[8 empty] | 106 | [6 empty] | 17.6655
[8 empty] | 105 | [6 empty] | 17.704
[8 empty] | 104 | [6 empty] | 17.7424
[8 empty] | 103 | [6 empty] | 17.7805
[8 empty] | 102 | [6 empty] | 17.8184
[8 empty] | 101 | [6 empty] | 17.8562
[8 empty] | 100 | [6 empty] | 17.8937
[8 empty] | 99 | [6 empty] | 17.931
[8 empty] | 98 | [6 empty] | 17.9682
[8 empty] | 97 | [6 empty] | 18.0052
[8 empty] | 96 | [6 empty] | 18.042
[8 empty] | 95 | [6 empty] | 18.0786
[8 empty] | 94 | [6 empty] | 18.1151
93
Note on D13: Roger Smith: Guess based on my own designs
Note on D14: Roger Smith: placeholder. May eway off. Assume fiberglass conductivity =0.04 W/mK
Note on J45: Roger Smith: Conduction from ambient + radiation from ambient - conduction to cooler
Note on K45: Roger Smith: delta T = time increment * net heat into detector / heat capacity; clamped at tmeperature setpoint to represent an ideal temprature servo
Note on N45: deltaT = neat heat into cooling head * time increment / heat capacity at cooling head
Scroll to bottom of table to see thermal circuit diagram
Ta
Rp
Rw
Rd
Tsp
Radiation(Ta,Td)
Cooling(Tc)
Td
Tc
Cd
Cc
idealized servo
Note that radiative load on window will be reduced some by radiative cooling of the window. How significant is this? I.e. what is the expected temperature profile on the inner surface?

## vignetting
f number | 2.5
Beam entering camera
| X | Y | Z | min X | min Y
| (mm) | (mm) | (mm) | (mm) | (mm)
CCDs | 194.32 | 124.88 | 0 | 194.32 | 124.88
window |  |  | 25 | 204.32 | 134.88 |  |  | Outer surface
shutter | 219.075 | 168.275 | 35 | 208.32 | 138.88 | ok | ok
filter | 238.125 | 152.4 | 60 | 218.32 | 148.88 | ok | ok
Beam obscuration
| X | Y | area | obscuration
| (m) | (m) | (sq.m)
Pupil (Schmidt Corrector) |  |  | 1.1675
Dewar base | 0.3556 | 0.3556 | 0.12645 | 10.83%
CFH shutter | 0.4064 | 0.4572 | 0.18581 | 15.92%
Bonn 2 blade shutter | 0.3556 | 0.45664 | 0.16238 | 13.91%
Sci In Tech split 2 blade shutter | 0.3556 | 0.3937 | 0.14 | 11.99%
Dual filter mechanism | 0.3556 | 0.50135 | 0.17828 | 15.27%
One amp or two:
amps | 1 | 1 | 2 | 2
P shift | 0.0001 | 0.0001 | 0.0001 | 0.0001 | s
read | 6.5e-06 | 3e-06 | 6.5e-06 | 3e-06 | s
active area | 7.7672 | 7.7672 | 6.7963 | 6.7963 | sq.deg
read time | 56.0286 | 26.0801 | 28.7389 | 13.4849 | s
exposure time (s) | sq.deg/hr | sq.deg/hr | sq.deg/hr | sq.deg/hr
25 | 345.0886 | 547.4153 | 455.2894 | 635.75
50 | 263.7217 | 367.534 | 310.7328 | 385.3952
75 | 213.4041 | 276.6324 | 235.8494 | 276.5079
100 | 179.211 | 221.7799 | 190.0495 | 215.595
150 | 135.7192 | 158.8029 | 136.8855 | 149.6577
200 | 109.2145 | 123.6819 | 106.9637 | 114.6066
250 | 91.3706 | 101.2823 | 87.7767 | 92.8584
Scattering off side of shutter and filter:
Max Beam angle for scattering | 0.4318 | 3
tan(angle) | 0.14393
angle | 8.1905
Height of flat side |  | 0.15945
min Total baffle depth | 0.02295
or…
pitch of cuts (weatherboard) |  | 0.01252 |  |  | half inch
min depth of cuts | 0.0018021
height of CCD dewar |  | 0.16422 | [5 empty] | However, the back plate will need to extend past the dewar wall and thus should form the primary baffle.
shutter overhang required | 0.023637 | [6 empty] | Cut the edge of back plate at an angle to give slope > 0.144
Actual shutter overhang in travel dir | 0.033 |  |  |  | ok
Actual shutter overhang orthogonal | 0.0254 |  |  |  | just ok

## Sensor locations
Heater servo, Td | center of detector support plate
Heat clamp at corner
CCD heat starp clamp, at center
Spreader tip
Spreader center | at cold head
Floatng shield?
Dewar case
Also need:
Re-route heater wiring as 3 banks of four 150ohm resistors
Microtemp thermal fuse, mounted in series with heater, using CCD heat strap clamping screw and thick wires to sensor on either side; crimp sockets
Shield, copper sheet, folded to make rectangular box, (electropolished?) and gold plated.
For getter container, McMaster Carr 85385T117, 304 SS wire cloth, 36% open area, 0.0015" opening size, $20/sq.ft |  | 0.0381 | mm
For washing out particles, part #85385T92 has 0.018" opening and 59% open area | 304 SS Wire Cloth 42 X 42 Mesh, .0055" Wire Diameter

## Vacuum
Approx volume | 9.4226 | liter
| Pfeiffer TMU 071, MVP 035-2 |  | Adixen/Alcatel
Pump roughing | 2.1 | m^3/hr |  | 166.6667
roughing time at atm | 16.1531 | s
Low pressure speed | 60 | l/s
0.05 | mbar
38 | mtorr
Sieve dimensions
Z | 22.22 | mm
X | 25.4
Y | 50.8
Volume | 28.6709 | ml
| 1.7496 | cu.in
2.5
0.67
3.2889
Axial motions at cryotiger head due to contraction:
Axial length | 145.5 | mm
Integrated CTE of SS304 | 0.0012 | assuming linear temperature profile and delata T to mid point
Contraction | 0.1746 | mm
Cu spreader distance | 95 | mm from support point nearest inlet of cryotiger |  |  | approx
Integrated CTE of Cu | 0.00247 | this number is for BeCu actually
Contraction | 0.23465 | mm
Support to center of Moly | 99 | mm
Integrated CTE | 0.000733
Contraction | 0.072567
Differential motion at head | -0.012517 | mm
Radial motion at cryotiger head due to contraction
Moly support length | 52.07 | mm
Integrated CTE of Moly | 0.00236 |  | 300K to 183K | 0.00236 | for Al | 0.000507 | for Moly | 0.00163 | for 304SS
Contraction | 0.12289 | mm
G10 post | 54.21 | mm
Integrated CTE of G10 | 0.000633 | assuming linear temperature profile; delta T to mid point
Contraction | 0.034315 | mm | 300K to average(300K,183K)=242K
G10 flexure | 40 | mm
Integrated CTE of G10 | 0.00153 | assuming linear temperature profile; delta T to mid point
Contraction | 0.0612 | mm | 300K to average(128K, 183K)=155K
Heat Spreader | 23 | mm
Integrated CTE of Cu | 0.00247 |  | 300K to 128K
Contraction | 0.05681 | mm
Motion towards window | 0.02944 | mm
Elongatin of SS tube at center of cryotiger
OD | 0.5 | inch
Wall thickness | 0.01 | inch
crossectional area | 0.0077754 | sq.in
| 5.0164e-06 | sq.m
length | 0.14 | m
2.4
Youngs modulus at 155K | 210.495 | Gpa
peak axial load | 25 | lb
| 111.23 | N
strain at peak axial load | 14.7473 | um
Activated carbon getter
diameter | 6.35 | cm
active depth | 2.032 | cm | guess
volume | 64.3518 | ml

## demister
Dry gas bleed in front of window.
Keep consdensation off window by filling space between filter and window with dry air
Storage Cylinder Volume | 40 | l
Initial storage pressure | 2000 | psi
vol at atm | 5714.2857 | l
1 cu.ft | 28317.0165 | ml
atm flow rate | 8 | ml/s
| 1.0171 | std cu.ft/hr
time till empty | 8.2672 | days
One week per cylinder change is a bit short, so investigate air dryers.
Example:
http://www.airdryers.com/prrd.htm