4/27/2007

Directions on how to run WCS and/or normal AO with the WCS installed

Note: These instructions apply to the subaperture relay also known as the WCS (Well Corrected Subaperture) relay, or originally, as the FQPM experiment.

- To run "normal" AO

  - start AO system

    - In /opt/ao, “ll bin” and verify that the link points to /home/aocm/ao3/bin

      - If it doesn’t, in /opt/ao

        - "rm –f bin"

        - "ln –s /home/aocm/ao3/bin"

    - From any directory, "start_ao -tcs"

  - Once the system is booted then

    - move fq_stim to 500,000 (this makes sure the relay stimulus is out)

    - move fq_relay to 0 (this makes sure the relay is out and we are looking at the full telescope)

  - Use the AO system as normal

  - DM Mask

There should not be one installed. Note only AO team members are authorized to remove or install a DM mask.

  - PHARO: run xpharo

- To run "WCS" AO

  - start AO system

    - In /opt/ao

      - "rm –f bin"

      - "ln –s /home/aousr/fqpm/bin"

    - From any directory, “start_ao –tcs”

  - Once the system is booted then

    - move fq_stim to 500,000 (this makes sure the relay stimulus is out)

    - move fq_relay to 34,500 (this makes sure the relay is in)

  - Load correct pixel gains file

I think you want to use pixel_gains_no_obs, although you might also want to try pixel_gains_all.

load wfp=ho, pixel_gains=aocp:/pixel_gains/pixel_gains_no_obs

  - Faint light performance/optimization

You can use play all the usual games as for the “normal” AO system, but if/when you try the optimal reconstructors use alphaxrxxxu.txt, the u is for no central obscuration for WCS.

  - DM Mask

The circular non-obscured mask should be installed on the DM if using WCS relay. Note only AO team members are authorized to remove or install a DM mask.

  - PHARO: run xpharo_fqpm

It might make sense to make a file called DEFAULT_WCS so that once the AO system is booted this can easily be restored.

- Switching between “normal” and “WCS” builds without restarting the whole AO system NOT TESTED

  - Exit TAO (the AO GUI)

  - cd to “/home/aocm/ao3/bin” or “/home/aousr/ao3/fqpm” as appropriate for the mode you are switching to

  - “./tao –ch aocp –th tcs” to restart tao

  - When TAO is exited PHARO will go into network simulation mode, once TAO is restarted the AO connection can be re-established by de-selecting simulate network from one of the XPHARO pulldown menus

  - restore the appropriate AO file DEFAULT or DEFAULT_WCS

  - DM Mask: There should be no DM mask for normal AO and the ciricular non-obscured mask for “WCS” AO. Note only AO team members are authorized to remove or install a DM mask.

  - PHARO:  exit PHARO GUI and re-start xpharo or xpharo_fqpm as appropriate.