BTO Control Math

C. Shelton

20 September, 2005 (rev 8)

- Overview

This document describes the 200-inch Beam Transfer Optics (BTO) from a control perspective, and derives the necessary control relations for its tracking features.  Figure 1 (Antonin’s) shows the BTO schematically and labels the relevant pieces.

Figure 1.  BTO Layout and Definitions

- Conventions

The signals from the quad cells Q1, Q2, Q3 are designated as q1, q2, q3, which are assumed to be already normalized and expressed as a unitless real number between -1.0 and +1.0.  Quad cell calibrations are called c1, c2, c3, are defined as the normalized signal change per mm, and are a function of spot size.  These signals and calibrations have x and y components.

The coordinate system used here follows the optical path (raytrace convention) and moves with the telescope.  The x axis is parallel to the dec axis at all locations, and the y axis is perpendicular to it and the local beam direction.  For example, at Q3, with the telescope at zenith, the x axis is east-west (horizontal) and the y axis is vertical.  At Q2, the x axis is east-west, the y axis is north-south.  As all of the BTO starting with M1b lies in a plane, the x direction will also be referred to as the out-of-plane direction.

The sign convention for x is that positive is west everywhere.  The sign convention for y is that positive is

up from M1 to M2;

north from M2 to M3;

up from M3 to M4;

south from M4 to M5/FSM;

down at Q3;

down at Q1.

At Q2, which is mounted out of plane, y is north and x is up.

- Input and Output Mapping

The relationship between the logical coordinate system above and the way the mirror actuators and the quad cell sensors are physically mounted is best described with pictures.  In these pictures, any actuator that needs its polarity reversed in software is labeled “reversed”.  The quad cell polarities are reversed, if needed, at the acquisition terminal block (see Table 1).

Figure 2 shows the two-inch version of the M1b mirror with its motors labeled x and y.   The four-inch version has a different lever arm (scaling), but requires no other change.  For both versions, the horizontal axis is x, connected to channel 1, the vertical is y, connected to channel 2, and there are no sign reversals.

Figure 2.  M1b Actuators

Figure 3 shows the trolley assembly, with quad cell Q1, and the M2 mirror with its two actuators M2x and M2y, and the M2rot rotary stage.  Positive rotation of M2r is indicated by the arrow; this is a +y deflection.  M2y creates –y deflection, additive to M2rot.  M2x is non-reversed.  M2x, M2y, and M2rot appear on channels 5, 6, 7, respectively.

Figure 3.  Trolley Assembly

Figure 4 shows the mounting arrangement of Q2 and M3.  Both axes of M3 are sign-reversed.  The x axis  appears on channel 1, y on channel 2.

Figure 4.  Q2 and M3

Figure 5 shows M4, Q3, the camera, iris, and fast steering mirror (FSM) on the side of the laser launch telescope (LLT).

Figure 5.  Laser Launch Telescope, with M4, Q3, FSM.

Figure 6 shows the orientation of the quad cells.  The A/D channel sequence for all Q’s is: first, pin 2 or “short” axis, then pin 6 or “long” axis, then pin 7, the sum signal.

Figure 6.  Quad Cell Orientation and Polarity

Table 1 summarizes the physical to logical mapping of the quad cells.  A negative sign means the signal should be inverted.

Table 1.  Quad Cell Axes Mapping

|  | Channel number, axis | Channel number, axis |
| --- | --- | --- |
| Q1 | 00, X | 01, -Y |
| Q2 | 03, -Y | 04, -X |
| Q3 | 08, -X | 09, -Y |

- Partitioning the Control

The BTO optics at M1 and before are fixed to the ground in the coude room.  So, for M1 only, define an x′, y′ which are rotated with respect to x, y by the hour angle of the telescope.  On the optical path from M1 to M2, a desired optical tilt in x, y coordinates is then rotated into x′, y′ coordinates before calculating mechanical tilts to be commanded to M1.

By applying this coordinate transformation to M1 commands, the control problem then neatly separates into non-interacting x and y parts.  These can be analyzed as two independent three-sensor, three-actuator systems.

- Mechanical and Optical Angles

Commands to the actuators are expressed two ways.  The first is an actuator position in microns, with symbols M1, M2, M3.  The second is as an optical deviation in radians, with symbols m1, m2, m3.  For M2 and M3, the conversion between the two is:

Mx=mxDx2cos(θ)My=myDy2

Here, D is the lever arm of the actuator, and θ is the angle of incidence.  It is assumed that the mirror tilt out of plane (x) is small, but that the mirror in-plane angle is arbitrary and has the value θ.  The value of θ for M2 is (90 + dec) /2; for M3 it is (90+15.1)/2, or 52.55 degrees.

For M1, the hour angle rotation must be included:

M1x=m1xcos(h)+m1ysin(h)D1x2cos(θ)M1y=−m1xsin(h)+m1ycos(h)D1y2

Here h is the hour angle of the telescope, in radians.  For M1b using the 2-inch mount, the measured θ is (90-3.55)/2, or 43.22 degrees.

Table 2 list the lever arms of the various mounts.  For the selected mounts, the x and y lever arms are identical.

Table 2.  Mirror Mount Lever Arms

|  | Mount type (Newport) | Lever arm (microns) |
| --- | --- | --- |
| M1b, 2 inch | U200-AC | 56100 |
| M1b, 4 inch | 605-4 | 51100 |
| M2 | U400-AC2X | 107700 |
| M3 | 605-4 | 51100 |

- Interaction and Control Matrices

The interaction matrix is simple enough: M1 affects Q1, Q2 and Q3; M2 affects Q2 and Q3; and M3 affects Q3.  Calling rij the distance from Mi to Qj, and cj the quad cell calibrations, the interaction matrix for either x or y can be written out as follows:

q1q2q3=r11c1r12c2r13c30r22c2r23c300r33c3m1m2m3

The inverse of this interaction matrix is

m1m2m3=1r11c1−r12r11r22c1r12r23−r13r22r11r22r33c101r22c2−r23r22r33c2001r33c3q1q2q3

As r11, r22 and r33 are always positive, the control matrix is non-singular iff the quad cell gains are all non-zero.  The control law can be expressed algebraically as follows:

m1=1r11c1q1−r12r22q2+r12r23−r13r22r22r33q3m2=1r22c2q2−r23r33q3m3=1r33c3q3

The above relations may be used unchanged for a system consisting only of M1 and Q1 by setting q2 and q3 to zero.  They may be used for a system consisting only of M1, M2, Q1, Q2 by setting q3 to zero.  This may be useful for initialization, or testing and integration.

If the beam misses a quad cell at any point in the beam train, there should be no attempt to control downstream actuators even if a valid quad cell signal exists farther on.  This can be enforced by setting to zero the first invalid q and all downstream q’s .  Doing so should create a system which automatically initializes itself as well as possible, short of invoking some search procedure.

- Sensor Gains

The cj are calculated using the assumptions that the laser beam is single-mode gaussian and has a waist at the LLT iris (basically, Q3) of 6.0 mm 1/e2 diameter.  The spot sizes at Q1 and Q2 are not very different from that at Q3; a small correction for standard Gaussian beam propagation is done for zenith only, and the variation in spot size with declination is neglected.  For Q1 and Q3 only, the y axis gain constant is lower because of the double reflection from their beamsplitters.   Q2 does not have this effect because it looks at leakage through the M2 mirror.  The details of this calculation are in the accompanying spread sheet “bto beam control v7.xls”   The resulting x and y gain constants are listed in Table 3.

Table 3.  Quad Cell Gain Constants

|  | X gain (mm-1) | Y gain (mm-1) |
| --- | --- | --- |
| c1 | 0.490 | 0.189 |
| c2 | 0.520 | 0.520 |
| c3 | 0.533 | 0.180 |

- Spacings

Call the radial spacing from the 200-inch optic axis to the M2-M3 optic axis D.  Viswa has a value of 3531.6 mm for D; independently and more roughly, I have measured 3455.0 mm. Viswa’s value will be used here.  Call the latitude of the 200-inch (+33.356 degrees) “lat”, and telescope declination “dec”. The calculations also assume that Q1 is at about the same location as M2, Q2 is at M3, Q3 is at M5/FSM.

With a combination of measurements and data from drawings, Viswa and I came up with the following spacings, at zenith.

M1-M2	12725 mm

M2-M3	13716 mm

M3-M4	3048 mm

M4-M5	1276 mm

This plus some geometry leads to the following values for the distances rij from mirror Mi to sensor Qj.  The units are mm.

r11=12725+D1cos(lat)−1cos(dec)=16963−3531.6cos(dec)r22=13716−Dtan(lat)−tan(dec)=16041+3531.6tan(dec)r33=4324r12=r11+r22r13=r12+r33r23=r22+r33

All of these distances except r33 change with declination.

The accompanying spreadsheet “bto control v6.xls” tabulates values of M1-M2 and M2-M3 distance, M2 angle of incidence, and M2 projected clear aperture for a range of zenith angles and declinations.

- Finishing the Controller

The use of the above equations and numbers results in a set of actuator position errors in microns.  To make a traditional integrate-only (PID minus the P and D) controller, these errors are multiplied by –k, where k is a dimensionless gain constant less than one, and used to increment the current commanded actuator positions.  Software limits are then imposed on the resulting positions, then position commands are sent out.  Adding amenities such as saving /restoring positions from file, loop on/off commands, background subtraction, piping data to a monitoring function, and the like, finish the controller.