BTO Control Math

- Conventions

The signals from the quad cells are designated as q1, q2, q3, and are assumed to be already normalized and expressed as a real number between -1.0 and +1.0.  These have x and y components.  Quad cell calibration is expressed as normalized signal change per mm, and is a function of spot size.  For this calculation, the beam is assumed to be single-mode gaussian and have a waist at the LLT iris of 6.0 mm 1/e2 diameter.  This results in quad cell calibration constants ranging from 0.18 mm-1 to 0.52 mm-1 at the different positions and axes.

The coordinate system follows the optical path (raytrace convention) and thus moves with the telescope.  The x axis is parallel to the dec axis at all locations, and the y axis is perpendicular to it and the local beam direction.  For example, at Q3, with the telescope at zenith, the x axis is east-west (horizontal) and the y axis is vertical.  At Q2, the x axis is east-west, the y axis is north-south.  As all of the BTO starting with M1 (M1b on Antonin’s chart) lies in a plane, the x direction will also be referred to as the out-of-plane direction.

The BTO optics at M1 and before are fixed to the ground in the coude room.  So, for M1 only, define an x′, y′ which are rotated with respect to x, y by the hour angle of the telescope.  On the optical path from M1 to M2, a desired optical tilt in x, y coordinates is then rotated into x′, y′ coordinates before calculating mechanical tilts to be commanded to M1.

By applying this coordinate transformation to M1 commands, the control problem then neatly separates into non-interacting x and y parts.  These can be analyzed as two independent three-sensor, three-actuator systems.

Commands to the actuators are expressed two ways.  The first is an actuator position in microns, with symbols M1, M2, M3.  The second is as an optical deviation in radians, with symbols m1, m2, m3.  For M2 and M3, the conversion between the two is:

Mx=mxDx2cos(θ)My=myDy2

Here, D is the lever arm of the actuator, and θ is the angle of incidence.  It is assumed that the mirror tilt out of plane (x) is small, but that the mirror in-plane angle is arbitrary and has the value θ.

For M1, the hour angle rotation must be included:

M1x=m1xcos(h)+m1ysin(h)D1x2cos(θ)M1y=−m1xsin(h)+m1ycos(h)D1y2

Here h is the hour angle of the telescope, in radians.  For M1, θ=(lat+90)/2, where lat=33.356 degrees, assuming that the path from M1a to M1b is vertical.

The lever arms D1 and D3 are 51100 microns, for both x and y; these mounts are Newport type 605-4.  The lever arm D2 is 107700 microns, for both x and y; this mount is a Newport U400-AC2X.

Interaction and Control Matrices

The interaction matrix is simple enough: M1 affects Q1, Q2 and Q3; M2 affects Q2 and Q3; and M3 affects Q3.  Calling rij the distance from Mi to Qj, and cj the quad cell calibrations, the interaction matrix for either x or y can be written out as follows:

q1q2q3=r11c1r12c2r13c30r22c2r23c300r33c3m1m2m3

The inverse of this interaction matrix is

m1m2m3=1r11c1−r12r11r22c1r12r23−r13r22r11r22r33c101r22c2−r23r22r33c2001r33c3q1q2q3

As r11, r22 and r33 are always positive, the control matrix is non-singular iff the quad cell gains are all non-zero.  The control law can be expressed algebraically as follows:

m1=1r11c1q1−r12r22q2+r12r23−r13r22r22r33q3m2=1r22c2q2−r23r33q3m3=1r33c3q3

- Components of the Matrices