Etalon Considerations

C. Shelton

6/2/2005

Ed’s proposal to use a 100 um stroke actuator for remote tuning of the laser is particularly attractive.  It replaces coarse motor control plus fine PZT control with all electrical, full span PZT, control.  This note is a consideration of stroke requirements and other implications.

Executive Summary

A classical standard for the control range of etalon tuning of a laser is 1.2 to 1.5 free spectral range (FSR).  This range allows unconditional acquisition of lock regardless of etalon temperature or other initial conditions, and allows for automatic “step-back” of the etalon by exactly 1.0 FSR upon reaching end of travel, resulting in seamless and unconditional tracking of line center.  We have previously discussed putting the actuator at the pivot position (not x or y) of the existing etalon mount, to shorten the effective lever arm to 26.9mm and improve the angular stroke. For a 100 um actuator stroke in this position to correspond to 1.2 FSR, the etalon must then be given a static angular offset of 3.67 degrees along the actuator tilt direction.  At this angle, there is a transverse beam walk-off of 40.2 um per etalon round trip.  This may compromise beam quality, the degree of which must be determined experimentally.

If the large angular offset cannot be accepted, there some backup possibilities.  One is to stuff all three positions of the etalon mount with 100 um stroke actuators, doubling the angular stroke and halving the angular offset and beam walk-off.  Another is to dispense with full range acquisition and tracking and only require a detune capability and lesser tracking range.  This would still require a manual adjustment per night, but would provide remote detune and a possibly adequate level of tracking during the night.

Discussion

Figure 1 shows etalon transmission as a function of actuator displacement.  The three cases are for different absolute phase offsets.  The absolute phase offset is 2π times the fractional part of 2nd/. For our purposes it is unknown but slowly varying.  The left end of the chart is normal incidence, with magnitude of the tilt increasing to the right.  Etalon tilt can be at any orientation to the x and y axes; in the case considered it is 45 degrees out of the plane of the laser.  These equations supporting these figures are in the appendix.  Other parameters are: 0.6 reflectivity; 1.06 um wavelength, 1.0 mm thickness of fused silica.

An equation for FSR as a function of tilt angle can be derived by computing dφ/d using the equation for phase φ in the appendix, then finding the actuator displacement that gives a phase change of 2π.

This gives

FSR = a/(2 n d sin()),  where

FSR=actuator displacement for 2π phase change

= the wavelength of the light

n = the index of refraction of the etalon material

d = the thickness of the etalon

 = the angle of the incoming light beam

a = the lever arm of the mount

Transverse beam walk-off per bounce is approximately 2d(n-1)/n.

Phase offset = 0.35 rad

Phase offset = 1.35 rad

Phase offset = 2.40 rad

Fig 1 Etalon transmission vs angle for three etalon phase offsets

Appendix

Etalon Equations

From: http://www.tecoptics.com/etalons/theory.htm

The transmission of an ideal etalon is

where

T = transmission

R = reflectivity of the mirrors

φ = the roundtrip phase change of the light ray

If any phase change at the mirror surfaces is ignored then

where

= the wavelength of the light

n = the index of refraction of the material between the mirrors

d = the distance between the mirrors

 = the angle of the incoming light beam

The above figure plots the etalon spectral transmission. The distance between adjacent peaks is the Free Spectral Range (D) and the width (FWHM) of each peak is the resolution (d). The Free Spectral Range can be written three ways:

Another useful concept for etalons is the finesse (F). This dimensionless parameter is the ratio of the free spectral range to the peak width.

For an ideal etalon, only the mirror reflectivity determines the finesse.