Palomar Adaptive Optics Test Plan

| Title | LGS Acquisition and characterization |
| --- | --- |
| Date | 12/06/06 |
| Lead | A. Bouchez,M. Troy |
| Time requested | 1 hr |
| Required conditions | Clear. |

Purpose Acquire LGS, focus LLT, determine LGS magnitude and spot size,

Test procedure - Setup - Configure AO to look at sky - Configure telescope for sky (open mirror cover, etc) – zenith only - Telescope at zenith, at best NGS focus. - Acquisition camera V filter not installed. - Move LLT mirror to its default position (determined during LLT alignment) [-11,53] - Load best available flatmap - Check white light position on HOWFS, - Move SSMs to center if necessary, then move back to sky - Acq focused at 90 km altitude (NGS focus-2450) move acq_z 12350 (for NGS use 14600) - LLT focused to 90 km altitude (NGS focus - 300). bto_control “move llt_focus 11730” - check/set laser focus to correct value bto_control “move laser_focus 7000” - Setup acquisition camera: - Start up IDL program acqview - Set integration time to 2s - In an appropriate experiment directory: IDL> ao_plot_vid_image - Perform final safety checks and propagate laser - Acquire LGS on Acq - If LGS is not in the Acq. FOV, use ellipticity of dichroic spots (apex points to LGS) and Raleigh gradient (brighter towards LGS) to steer it in: offset llt_a +X=up; offset llt_b +X=left - Calibrate UTT mirror throw (one-time calibration) - Center LGS in acquisition camera field. - Record image with ACQVIEW and note time tag in log in Results section. - Record images dithering laser in a cross pattern (see log). - Manually center laser behind reflective spot. - Focus on Na layer - Move LGS to clear region of ACQ. field: offset llt_b +30 - Rough focus laser spot: bto_control “offset llt_focus …” in steps of 100 move acq_z … in steps of 300 - Detune laser, take and save a 2s background image. - Focus Acq - ao_focus_loop, current-600, 300, 5, ‘acq’, ‘save_name’, ‘sky_name’, time=6.0 - ao_read_focus, data,file=’save_name’ - move acq_z best_focus - Focus LLT - ao_focus_loop, current-100, 50, 5, ‘llt’, ‘save_name’, ‘sky_name’, time=6.0 - ao_read_focus, data,file=’save_name’ - bto_control “move llt_focus XXX” - Repeat steps 4.3 and 4.4 again if necessary - Tune center wavelength - This depends…. - Take photometry - Install V filter in Acq. camera - Record and save a final set of detuned and tuned 2s Acq images for measuring photometry and spot size. - Image photometric calibrator - Go to photometric standard NGS, Landolt ~mV=11.0 - Move acq_z to NGS focus (move acq_z 14600) - Move star to approximate position of LGS - Take two images, with a ~10” dither between them; record star name and image time tags.

Results and conclusions UTT mirror calibration (section 4) Mirror commands; llt_a position; llt_b position; Image time tag; x centroid (pix); y centroid (pix). centered; -19; 19.9; 1165393512; 468.3; 267.3. offset llt_a 20; 0.9; 19.9; 1165393596; 455.2; 380.8. offset llt_a -40; 0.9; 59.9; 1165393656; 271.7; 357.3. offset llt_a 20; -19; 59.9; 1165393694; 285.7; 238.1. Optimal focus settings (section 5) UT Date; Dec 6 2006. LLT focus (µm); 11730. Acq_z (µm); 12350. LGS FWHM (pix); 11.5 (!). Photometric calibration (section 7) UT Date; Dec 6 2006. Star ID; Landolt 98-185. Image time tags; landolt98-185_1. landolt98-185_2. Star FWHM (pix); 10. * Fainter star of the pair (mV=10.54) – 2s exposures