Palomar Adaptive Optics Test Plan

| Title | NGS AO Checkout (12/06 v2.2) |
| --- | --- |
| Date | Every AO engineering night |
| Lead | A. Bouchez |
| Time requested | 20 minutes. Start after 6º twilight. |
| Required conditions | N/A |

Purpose Verify acceptable performance of the NGS AO system for subsequent engineering tests, and measuring seeing, on-axis Strehl and isoplanatic angle.

Test procedure - Choose a binary star near zenith from the list below. - Acquire the primary with the AO system with an initial frame rate of 500Hz. Adjust the frame rate to keep counts above 200 and below saturation. - Take a wavefront sensor sky. - Perform both coarse and fine WFS-DM registration. - Verify that co_default2 is loaded and that focus offloading is on. - Close the TT and DM loops. - Create a flat map on the star. - IDL> ao_make_dm_flat_map, filename - Set the flat map to default when complete. - While waiting for the flat map, check signal levels on PHARO: - Initial setup: 25mas FOV, filter Ks, ND 1%, standard cross, 2s integration. - Adjust the exposure time and filter to get 10-15k peak counts on brighter star. - With the loops closed, adjust tel. pointing to center both stars in the PHARO FOV. - Open the TT and DM loops. - Repeat the lenslet to DM registration. - Offset to sky and take a new WFS background. - Record 3 PHARO sky frames. - Offset back to star and close TT and DM loops. - Record 3 images with PHARO. Note the directory, frame numbers, and Strehl. - Open the DM loop only. - Record 1 image with PHARO, 30s integration. Note frame number and FWHM. - Load the latest white light flat map load wfp=ho, dm_calib_pos=aocp:/tables/default/wl_flat_map - Record 1 image with PHARO, 30s integration. Note frame number and FWHM. - Record the telescope focus.

Results (use one column per night) UT date; Dec 6. Weather conditions; GORGEOUS!. PHARO directory; /scr1/2006dec06. Target WDS #; 75567. 13. Sky frames; ph0000-0002. 15. Closed-loop frames; ph0003-0005. 15. Closed-loop Strehl; 62%. 17. Open-loop frame; ph0006. 17. Open-loop FWHM; 0.4”. 18. WL flat frame. 19. WL flat FWHM. 20. Telescope focus; 60.32.

Star list

Result of a VizieR search of the Washinton Double Star Catalogue (I/237) with the following constraints:  (DE2000: "18:20:00..48:20:00")  AND   (Sep1: "8..15")  AND   (Sep2: "8..15")  AND   (MagA: "5..7.5")  AND   (MagB: ">8")

This list is available in the Palomar format at: /export/home/pharo/lgs/wds2.lst

The digit in the first column provides a link to the Vizier page on the star.

Full    recno    RA2000  DE2000  pa1   Sep1   MagA   MagB

"h:m:s" "d:m:s"  deg arcsec   mag    mag

1      1632  00 32.8  +28 17  160    8.7   6.30  11.26

2      4755  01 41.3  +25 45   32   11.3   6.30  10.90

3      6043  02 09.6  +42 51  183   10.9   7.20  11.00

4      7194  02 38.3  +37 44  260    9.2   6.16  11.30

5      8161  03 01.5  +32 25    8    8.2   6.90   8.40

6     10866  04 09.6  +31 39  174   12.8   6.90  14.10

7     11686  04 28.9  +30 22   36   14.2   6.38   8.28

8     13555  05 09.8  +28 02   27   11.8   6.00   8.65

9     13915  05 15.4  +32 41  342   12.6   5.02  11.10

10     15769  05 43.0  +33 19  204   11.5   6.80  11.80

12     23030  07 06.2  +24 52   50   14.1   7.10  11.10

15     37648  11 38.7  +45 07  266   10.5   6.46   8.39

16     39044  12 17.5  +28 56  275    8.2   5.70  10.20

17     39405  12 28.1  +44 48  161   10.4   7.50   8.12

18     40881  13 10.9  +21 14  316   10.9   6.82  10.80

19     40900  13 11.5  +21 55  288   10.6   6.82  10.90

20     42896  14 06.7  +34 47   71   14.2   7.00  10.20

22     52898  18 03.9  +26 39  193   12.4   7.00  12.00

23     57496  19 03.9  +34 09  295   11.1   7.30  12.80

24     57902  19 08.8  +34 46  271   12.9   7.00   8.60

25     60632  19 41.3  +30 43   29    9.1   7.30  13.40

26     61322  19 49.0  +19 09  311    8.6   5.04   9.03

27     64828  20 24.2  +29 00  134   13.0   7.20  11.20

31     68328  21 05.7  +47 48  315   10.4   7.45  12.00

32     69741  21 24.1  +25 19  330    8.5   6.20  12.20

33     74560  22 46.0  +19 15  270    8.7   7.47  10.80