Palomar Adaptive Optics Test Plan

| Title | BTO setup |
| --- | --- |
| Version | 2.1 |
| Date released | 4/23/2007 |
| Lead | A. Bouchez |
| Time requested | 1 hr during the first afternoon of each observing run. |
| Required conditions | Dome closed |

Purpose Power up BTO, align mirrors using 660nm stimulus beam, save default positions. Align LLT to 660nm stimulus beam.

Test procedure Personnel required: BTO operator + 1 personnel in dome if necessary. Request that Coude block be unlocked before beginning procedure. Power up Coude hardware: Uncover BTO Coude optics. Remove blue FSM mirror cover. Open ports in Coude bench enclosure. Power up ESP300 and FSM controllers. Turn on all 3 motors. Power up power strip behind Coude bench (laser & diagnostics electronics). Power up 660nm laser using the following procedure: Unplug black wire from power supply. Turn on power supply. Check that power supply is set to 3.75V, 0.75A. Turn off power supply and reattach wire. Power on laser. Power up Coude lab video camera, intercom, and both microphones. Test intercom volume. Coude controls adjusts volume as heard from data room... Power up Dome hardware: Verify that telescope focus is nominal for AO (~61 mm) Power up all BTO electronics in dome from a computer in the data room (String). telnet viswa_power > /on C2 > /on C3 > /on trolley > /on bto > /x Wait ~3 minutes for BTO computer to boot. Start up BTO GUI on a computer in the data room. ssh -Y aousr@198.202.125.152 (lgs.palomar.caltech.edu) bto_gui > reset all (Note: this moves the laser focus stage on laser bench!) > set trigger=internal If laser trigger is not on, check servo state rate Prepare trolley optics for operation Move trolley to +5200000. Remove optics covers. Dust optics with Aero-Duster Align BTO. Connect TAO to BTO (under TAO menu bar...) Alternatively, > set hadec=tcs Set Track ON to move BTO mirrors to last zenith position. Open coude block. Set servo on to erect the BTO If no light is visible on Q1... Visually check Coude beam path. Verify that stimulus beam is precisely centered on both Coude irises. Verify that telescope pointing data is OK (see step 8) Open servo loops and save default positions (for 660nm laser) > File/set zenith and save DEFAULT Align LLT to laser (in prime focus) Remove LLT primary cover. Verify that UTT mirror controller is powered on. Lock BTO servo loop with high speed on Q3. Adjust the Q3 beamsplitter to center the laser on the UTT mirror. Adjust the top fold mirror to center the laser on it. Install reticles in the LLT primary mirror hole and on the LLT secondary. Iterate between UTT mirror adjustments (performed by the AO operator) and adjustments of the LLT final fold mirror to center the laser on both retticles. Remove the reticles and verify that the beam pattern on the dome ceiling is evenly illuminated and centered on the LLT secondary shadow. Open the BTO servo loop. Verify that the laser beam path in prime focus is clear of obstructions and everything appears ready for the high-power 589nm laser.

Results and conclusions