Palomar Adaptive Optics Test Plan

| Title | Abbreviated BTO checkout |
| --- | --- |
| Version | 1.0 |
| Date released | 8/26/07 |
| Lead | A. Bouchez |
| Time requested | 15 min. |
| Required conditions | Dome closed |

Purpose Use only on later nights of an observing run, when no major changes have been made to the laser or BTO. - Verify BTO alignment and functionality. - Test Q3 safety interlock.

Test procedure Personnel required: BTO operator, safety officer. Setup: - Dome closed. - LLT primary mirror cover removed and LLT ready for high power laser (visually check diagnostics bench, cables). - 589nm laser at high power (>5W), checked for obstructions (power meter, etc.) - Stimulus laser powered on. - Safety interlock system tested, all zones active. - Remove Coude block. - Arm alarm. Verify BTO alignment and functionality - Lock BTO servo on the stimulus laser: - Set Laser to “Stimulus” in BTO GUI. - Set all quad cells enabled and High-Speed to Q3. - “Track On”. - Wait for “Track: completed.” message in message window. - “Servo On”. - Verify BTO servo locked with high-speed on Q3. - Under File menu, choose “Set Zenith and Save DEFAULT”. - Set laser to 589nm. - Set trigger to "Laser". - Open the laser shutter. - If Q3 interlock closes laser shutter... - Reset BTO alarm, try again (twice). - Give up and follow full BTO alignment procedure starting at step 11. - Record Flux in BTO GUI. - “Set Zenith and Save DEFAULT”. - Shutter laser [continued on next page] Test Q3 interlock - Disarm alarm and install Coude block. - Arm alarm system. - Test-fire laser to demonstrate BTO interlock functionality. - Reset alarm on BTO.

Results (use one column per night)