Palomar Adaptive Optics Test Plan

| Title | Laser Acquisition on science nights |
| --- | --- |
| Version | 5.1 |
| Date released | 4/4/07 |
| Lead | A. Bouchez,M. Troy |
| Time requested | 30 min. |
| Required conditions | N/A |

Purpose Acquire LGS, focus LLT, determine LGS magnitude and spot size,

Test procedure Prepare to project laser Verify Space that Command has been called Telescope at zenith. Dome open. Mirror cover open. Raise windscreen. Spotters on station. Radar on and adjusted. ASCAM and IRCAM on and alarms reset. Remove the Coude block BTO configured for 589nm laser at zenith. Laser alarm system ready. AO Setup Check stimulus position on HOWFS. Move SSMs to center stimulus on HOWFS if necessary. Configure AO to look at sky Verify LLT FSM mirror position (last good position, previous night) Confirm that best available flatmap is loaded. Set TAO to LGS mode Move LGS_X to NGS position. Move ACQ_Z to LGS position (9950 um) Focus LLT to last good position (or NGS focus - 300). bto move llt_focus 11780 check/set laser focus to correct value bto move laser_focus 7000 (4/4/07 UT) Set Acq ON, integration = 2s. In an appropriate experiment directory: IDL> ao_plot_vid_image Project laser. Move LGS to a clear location on Acq. offset llt_a +X=up; offset llt_b +X=left If LGS is not in the Acq. FOV, use ellipticity of dichroic spots (apex points to LGS). Focus LLT on Na layer Manually offset LLT focus in 30um steps (bto offset llt_focus...), checking FWHM with ACQVIEW Gaussian fit at each position: -30, 0, 30 Determine laser focus (after major laser changes) Block laser, save a 2s background image using ao_plot_vid_image. Rough focus laser spot, looking at image in acq. camera bto_control “offset laser_focus …” in steps of 2000 ao_focus_loop, 0, 2000, 7, ‘laser_focus’,’save_name, sky=’sky_name’,time=6.0 bto move laser_focus XXX Save a final image of the LGS. Note file name below.

Results and conclusions From 5/07 run: LGS acq_z = 10275 llt_focus = 11865 laser_focus = 10000 llt mirror = 75, 10