Palomar Adaptive Optics Test Plan

| Title | Laser power-up |
| --- | --- |
| Version | 2.1 |
| Date released | 8/8/2007 |
| Lead | V. Velur |
| Time requested | 1 hr prior to laser use |
| Required conditions | N/A |

Purpose - Safely power up laser. - Optimize laser power, frequency tuning. - If necessary, reduce power for alignment procedures.

Test procedure A. Power up laser - Top up water in large chiller. - Switch on both chillers (button/switch on front of each). - Check “load center Coude camera room” circuit breaker. Reset any necessary ( 10 &12 often trip when shutting laser down) - Press white button on laser controller, labeled “Laser Flow”. - Pressure should read ~55 PSI, flow should read ~0.7 Gal/Min. - On temp control laptop, right-click on each window and restart display (on 1.06 AOM, 1.06 Etalon, 1.32 AOM, 1.32 LBO). Wait for temps to stabilize. - Verify that both IR beams are blocked (two flipping blocks, one next to 1.32 diodes, one near SFG) - Verify that laser function generator is set to 1 (=10 Hz pulses) - Switch photodiodes on (1.06, 1.32, 1.06 diode). - On both diode power supplies (Analogue Modules): - Turn key switch to on. - When alarm stops, hit “A” (pulse mode operation) - Press “run” (High power IR – green or purple goggles required) - Turn up laser function generator dial to 30 (=300 Hz pulses). - Verify that both lasers are lasing. Pulse shape should now look reasonable. - Verify that LBO temp display is at ~40.1 C (lowest left in rack). - Press “enable” on Na cell temp controller (lowest right in rack). - Check 589nm unmodelocked power. - Turn on power meter and insert in beam just ahead of output optics. - Zero power meter. - Open IR beam blocks (High power 589nm – purple goggles required) - Record power. - Verify that IR beams shapes are good in spiricon display (possibly obsolete!) - Mode lock IR lasers. - Turn “1.06 power adjust” knob clockwise until you reach black mark (=2). - Turn “1.32 power adjust” knob clockwise until just above black mark (=4). - Adjust 1.32 cavity length using micrometer near diodes to clean up mode locking (make clean comb). Can also use x-y aperture in bow-tie for final cleanup (caution - touchy!) - Adjust “Phase Control” knob until 589nm power is maximized. - Turn on chirping: Set “output” knob of pulse genertator such that output of PZT driver reads ~100 mV peak-to-peak. - Repeat step 16.3 to clean up mode locking. - Frequency adjustment (when Na cell has reached ~69 C) - Adjust horizontal actuator on 1.06 etalon to maximize fluorescence as seen in oscilloscope output. - Record power and remove power meter from beam. - Laser is ready for observing. To change to low power for alignment, dephase the IR lasers by adjusting the “phase contol” knob to achieve desired power. Laser optimization - Laser optimization, level 1 - Adjust the “Phase Control” knob until 589nm power is maximized. - Adjust the 1.32 intracavity lens x-y position, to maximize total 589nm power and minimize 1.32 laser turn-on time (the two should be correlated). - Adjust the 1.06 cavity length using the manual actuator on the 1.06 output coupler z stage, to maximize total 589nm power and minimize 1.06 laser turn-on time. Beware of burning your fingers! - IR laser optimization, level 2 (expert users only) (possibly obsolete!) - Block 1.32 cavity using intracavity flip mount. Remove 1.32 etalon by removing post from mount. Open intracavity flip mount. - Maximize 1.32 power by adjusting x-y position of intracavity lens (use oscilloscope trace) - Replace etalon as in step 1. - Block 1.06 cavity using intracavity flip mount. Remove 1.06 pinhole by unscrewing pinhole from x-y lens mount. Open intracavity flip mount. - Maximize 1.06 power by output coupler z stage (use oscilloscope trace). - Replace etalon as in step 11.4

Results and conclusions 8/8/09 startup 1.32: 8.20 W after adjusting x & y of intracavity lens. Some squiggle in trace. 1.06: 13.6 W with no adjustments. Would adjust x-y of aperture in bow-tie if nec. 589nm: 3.15 W w/ no mode locking. 7.50 W with chirping on, after all adjustment. Note: Chirping destabilized 1.32 cavity length, required careful adjustment of micrometer.