Palomar Adaptive Optics Test Plan

| Title | LGS Target Acquisition (SWIFT) |
| --- | --- |
| Version | 1.0 |
| Date released | 8/8/2009 |
| Lead | A. Bouchez |
| Time requested | 12 min. per target |
| Required conditions | Sufficiently clear for laser projection |

Purpose Acquire science target, lock all loops and hand over system to observer.

Test procedure - Slew to a V=6 SAO star near the LGS target. - Setup AO system for NGS. - Rotate the Cass ring if needed. - Set TAO mode to NGS. - Click "SETUP FOR NGS" with HOWFS at 500 Hz. - Stop chopper if running. (R/X) - Acquire and register HOWFS. - Move star behind reflective spot using hand paddle. - Click "OFFSET + TAKE HOWFS BKG". - Close TT loop. - Register DM (coarse and fine). - Close DM loop. - Make a new telescope flat map and load. - Ask observer to center star at the desired location for LGS acquisition. - Save AO config. - TX telescope. - Slew to LGS target tip/tilt star. - Request laser propagation. - Set up AO system for LGS. - Set TAO mode to LGS. - Click "SETUP FOR LGS with HOWFS at appropriate value (150 Hz). - Start chopper (R/X/E/E). - Acquire LGS - Center laser on reflective spot using UTTPADDLE. - Record centered position of LLT FSM. - Close the DM loop on the laser. - Move LGS_X to LGS position. - Update reconstructor with AORECON if Cass ring angle has changed. - Click "OFFSET + TAKE HOWFS BKG". - Close UTT loop. - Verify that HOWFS framerate is correct. If necessary: Open loop, change framerate, take an offset background, close UTT loop. - Close DM loop. - Acquire the NGS and close TT loop. - Set LOWFS framerate appropriate for NGS. - Move ACQ_Z to NGS position. - Set Acq to integration time appropriate for NGS. - Identify NGS in ACQVIEW field and send LOWFS. - Click "OFFSET + TAKE LOWFS BKG". - Close TTM loop. - Verify that LOWFS framerate is correct. If necessary: Open TTM loop, change LOWFS framerate, take an offset background, close TTM loop. - Close focus loop (pulldown on LGSFOC). - Close focus loop on LGSFOC tool, gain=1.0, update 10s. - Wait for focus to converge to <0.5 mm, then set gain=0.25, update 30s. - Ask observer to position NGS at desired location in the field. - Record PSF image cube if desired - Zero the telescope offsets - Offset to science target - Set tip/tilt loop gain to 0. - Ask observers to offset telescope to science target. - Once telescope and LOWFS are in position, resume tip/tilt control If total offset <45” - Reset the tip/tilt loops gain to the correct value (default: 0.7) If total offset >45” - Open the tip/tilt loop - Click “OFFSET + TAKE LOWFS BKG” - Reset the tip/tilt loop gain to the correct value (default: 0.7) - Close the tip/tilt loop - Close the focus control loop - Hand over control to observer

Results and conclusions