Palomar Adaptive Optics Test Plan

| Title | BTO setup |
| --- | --- |
| Version | 2.0 |
| Date released | 4/3/2007 |
| Lead | A. Bouchez |
| Time requested | 1 hr during the first afternoon of each observing run. |
| Required conditions | Dome closed |

Purpose - Power up BTO, align mirrors using 660nm stimulus beam, save default positions. - Align LLT to 660nm stimulus beam.

Test procedure Personnel required: AO operator (lead), BTO operator, 2x prime focus personnel. Power up Coude hardware: - Uncover BTO Coude bench. - Power on ESP300 and FSM controllers. - Power on 660nm laser. Power supply should be set to 3.75V, 0.75A. Follow procedures to ensure diode safety! Power up Dome hardware: - Verify that telescope focus is nominal for AO (~61 mm) - Telnet to ''viswa_power" and turn on all devices. - /on C2 - /on C3 - /on trolley - /on bto - /x - Disable BTO firewall - ssh root@bto.palomar.caltech.edu - /etc/init.d/ittables stop - Start up BTO command line. - ssh -X aousr@bto.palomar.caltech.edu - bto_cmd - > reset all (Note: this moves laser focus stage on laser bench!) - > set trigger=software - Move trolley to +5200000 and remove optics covers. Align BTO. - Connect TAO to BTO (under TAO menu bar...) - Alternatively, > set hadec=tcs - Set Track ON to move BTO mirrors to last zenith position. - Open coude block. - Erect the BTO by closing the servo loop to Q1, then Q2, then Q3. - If no light is visible on Q1... - Visually check Coude beam path. - Verify that stimulus beam is precisely centered on both Coude irises. - Verify that telescope pointing data is OK (see step 8) - Open servo loops and save default positions (for 660nm laser) - > set zenith - > save DEFAULT Align LLT to laser (in prime focus) - Remove LLT primary cover. - Verify that UTT mirror controller is powered on. - Lock BTO servo loop with high speed on Q3. - Adjust the Q3 beamsplitter to center the laser on the UTT mirror. - Adjust the top fold mirror to center the laser on it. - Install reticles in the LLT primary mirror hole and on the LLT secondary. - Iterate between UTT mirror adjustments (performed by the AO operator) and adjustments of the LLT final fold mirror to center the laser on both retticles. - Remove the reticles and verify that the beam pattern on the dome ceiling is evenly illuminated and centered on the LLT secondary shadow. - Open the BTO servo loop. - Verify that the laser beam path in prime focus is clear of obstructions and everything appears ready for the high-power 589nm laser.

Results and conclusions