## Slide 1: SED Machine: Summary of Potential Mechanical Improvements
November 7, 2016
Reston Nash

## Slide 2: Primary Objectives of Upgrade
Stabilize the IFU Lens Barrel.
Move electronics/cabling outside enclosure.
“Seal” the enclosure.
Reinstate IFU camera focus stage.

## Slide 3: Secondary Objectives of Upgrade
Add kinematic features to the cameras.
  Repeatably mount them.
Clean optics.
Add baffling.

## Slide 4: General Strategy
Work with what we have.
  Retrofits not new design.
Assume no further upgrades will be done.
  Try not to leave anything on the wish list.
Don’t plan on breaking the alignment.
  If we do then we can switch one of the lenses.

## Slide 5: Background
CAD file seems pretty representative.

## Slide 6: Background
Lenslet Barrel
Prism
IFU Barrel

## Slide 7: 1) Optical Stability
Both lens barrels should be rigidly clamped.
Current springs are too flimsy.
Rigid clamps on IFU Barrel
Currently: Springs are clearly not very tense
CAD Model

## Slide 8: 2) Electronics
The electronics were designed to be inside the enclosure.
  This leads to a big mess, no matter how hard you try.

## Slide 9: 3) “Seal” the enclosure
Currently there are a lot of holes in the enclosure, so the instrument is fairly exposed.
Making an airtight seal is almost impossible.
  Usually fancy instruments flow dry air or nitrogen through the volume, but that could be a lot of effort.
The enclosure should be made as air tight as possible, then use a desiccant and see how long it takes to saturate.

## Slide 10: “Seal” the enclosure
Most enclosure mounting holes are unused.
etc.

## Slide 11: 4) Focusing Strategy
The tracking (RC) camera and the IFU camera need to both be in focus.
  Ideally, M2 would be used to focus the RC camera,
  Then the IFU focus stage would focus the IFU camera.
  Currently IFU focus stage is loaded by gravity (?), and not springs.

## Slide 12: Camera Focus Mechanisms
IFU Camera
Motorized Nudger
Flexures
IFU Lens Barrel
Stage springs

## Slide 13: Baffling
Mixtures of plates, paper, and tape.

## Slide 14: Open Questions
What is the point of the RC camera enclosure? (Safety, likely not required to make light tight)
Do we want to redo the focus mechanism. (no, keep it and just add springs. It doesn’t have to be pretty)

## Slide 15: RC Camera Enclosure
Seems light tight
Seems unnecessary

## Slide 16: Minimum Parts to keep:
The enclosure should be redone.
A lot of the structure seems unnecessary.

## Slide 17: Potential New Enclosure Geometry
Electronics

## Slide 18: Next Steps
The CAD design for any upgrades shouldn’t take very long.
  New enclosure, new focus stage, some cabling and electronics mounting plates.
Need to decide the functionality we want.
Schedule is driven by ZTF downtime.
  6 weeks (?), seems like a comfortable amount of time.