## Slide 1: Questions on IFPAC_Requiement_Document

## Slide 2: 2.1.1 Biases : Summary of Capabilities
This card can control up to 2 CCDs
12 Biases must be supplied for each CCD and total 24 Biases per card (not 48 Biases)
16 Biases Voltage range = 0 to 27V, Generated by Op-amp AD8513 is working on 0-30V Power.
Current driving capability is +/- 50 mA for both
+/-12V and +27V Bias Voltages
Provision of filtering Biases using Power Line Ferrite bid and 1.5uF Feedthrough Capacitor on Backplane Card.

## Slide 3: 2.1.2 Testing Procedures Bias Voltage Drive Current
The maximum deviation between Loaded and Unloaded Bias Voltage should be as per CCD Data sheet and not RMS Bias voltage noise which is 1mV RMS

## Slide 4: 2.2.1 Summary of Capability of Clock Lines
Number of Clocks 30, swing +/- 10 V
One High Voltage Clock i.e. 0 to 55 V swing and not -10V to + 60 V
Do we need high voltage clock for our CCD ?

## Slide 5: 2.3.1 Signal Chain summary of Capabilities
Mainly written about ADC and not complete Video chain
The differential voltage range for the ADC input is from 0 V to + 4.096V.