## Slide 1: Filter Exchanger Requirements
Roger Smith
2016-01-22

## Slide 2: Overview
Yes, Requirements are documented (11 pages):
  http://www.oir.caltech.edu/twiki_oir/pub/Palomar/ZTF/Requirements/CIN_670_ZTF_Filter_and_Exchanger_Requirements.docx
Let’s look at driving requirements, and note some others.
Michael Porter will show a requirements compliance matrix.

## Slide 3: Requirements
Safety (personnel, telescope, instrument, filters)
  by design and by test …. metrics/protocols TBD but “handoff” is allowed
   cost & schedule
Functionality
  Automated exchange between 3 filters or none (all stowed).
  Manual change of filter set by day (safe, easy)
  Mechanical protection of stowed filters from dust or humans.
Survey rate
  Exchange time <90s [60s goal]
  Beam obstruction <1% on axis, 2% off axis [0.5%, 1% goals]
  Reliability: MTBF > 33,000 exchanges; down time < 2 night/yr
  Impact on tube seeing: surface temperature with 3C of ambient.
  Positioning accuracy: +-0.25mm in X,Y, or Z.

## Slide 4: Safety Requirements
Designs are not permitted different risk posture so this translates to cost and schedule required to meet need.
Safety for personnel
  Interlocks
  Limit forces?
Safety for telescope
  Interlocks, redundancy fro safe handoff
  Verify by repetitive test in various orientations.
Safety for instrument
  Force limits
  Soft frame surface (cannot scratch window)
Safety for filters
  All of the above
  Enclosed when stowed to protect from dust and harm by humans

## Slide 5: Functional Requirements
Automated exchange between 3 filters
  Or observe with no filter (all stowed)
  Easy upgrade to more filters is desirable but not a requirement.
Manual change of filter set by day
  Need infrastructure to make this safe (easy)

## Slide 6: Requirements affecting survey rate
Exchange time <90s [60s goal]
  Cooperative telescope motion allowed but counted as positioning overhead.
  Currently a maximum is specified but it the average which is more interesting for survey rate. Maybe the goal is for the average.
Beam obstruction <1% on axis, 2% off axis [0.5%, 1% goals]
Reliability: MTBF > 33,000 exchanges; down time < 2 night/yr
  Verification method is not yet specified.
Impact on tube seeing: surface temperature with 3C of ambient.
Positioning accuracy: +-0.25mm in X,Y, or Z. Tilt TBD.

## Slide 7: Concerns
No formal requirement yet:
Obstruction of human access to telescope interior.
Mass of filter exchanger
Mass of electronics
Heat dissipation in electronic enclosure
No back drive (part of safety requirements)
Knowledge of positioning preserved across power cycle