## Slide 1: ZTF Gasket and window stress under 1 atm pressure
Callahan
10/23/2014

## Slide 2: Model names
CAD model 1100=ZTF0000
FEA model: Window with gasket

## Slide 3: Boundary Conditions 1 ATM

## Slide 4: Window stress Max Principal Design Strength of Fused Silca 8.9MPa (FOS=4)

## Slide 5: Gasket stress Von Mises ~13 MPa Strength of NytekCF = 52MPa [7500 psi]

## Slide 6: Gasket compression 5 microns

## Slide 7: 11/7/2014
Increase window size by 6mm to accommodate 3mm spacing between CCD’s.
Change gasket to G10 so that a heating circuit can be embedded.

## Slide 8: Apply 1 atmosphere 14.7 psi to top surface inside O-ring (pressure inside surface)
Frictional contacts between window and gasket and gasket and cryostat enclosure. COF=0.2 to solve quickly and allow surfaces to open and close.
Elastic supports added to window solve.

## Slide 9: Boundary conditions

## Slide 10: Maximum Principal stress in window. 9.7MPa

## Slide 11: Determine factor of safety vs. probability of success
Use MathCad model: Design Strength-Fused Silica.

## Slide 12: Strength of Fused Silica Window vs. Probability of success. (95% probability of success shown by markers)

## Slide 13: Factor of safety at 9.9MPa vs probability of success. Markers show 95% probability of success has a FOS=4.4

## Slide 14: Another way to look at this is to calcuate the probability of success at 9.9MPa.
The probability of survival can be calculated at 9.7MPa using the above MathCad model: 1 part in 5 x 10E7.

## Slide 15: Cryostat assembly Gravity loads

## Slide 16: 1 g vertical loads max deflection 1.2 microns
Requirement: 580 micron profile of window w.r.t. optical axis.
Loadcase: 1g vertical
Supported: at 3x hexapod locations
20 Kg

## Slide 17: Von Mises stress 1 atmosphere pressure
Wt: 27 Kg

## Slide 18: Von Mises stress Max <3000 psi Yield strength 304 stainless = 31200 psi (Matweb.com)

## Slide 19: Z-deformation 113 microns