IRMS: MOSFIRE on TMT

TMT.INS.CDD.07.001.REL01

Ian Powell, Jenny Atwood, Peter Byrnes

HIA

17 Apr 2007

Table of ContentS

1	Introduction	4

1.1	Purpose	4

1.2	Scope	4

1.3	Applicable Documents	4

1.4	Reference Documents	4

1.5	Change Record	5

1.6	Abbreviations and Acronyms	5

1.7	Glossary	5

1.8	Verb Convention	5

1.9	Acknowledgements	5

2	IRMS introduction and top-level requirements	6

3	Optical DESCRIPTIONs	7

3.1	MOSFIRE on KECK	7

3.2	MOSFIRE on TMT (with no modifications)	7

3.3	MOSFIRE on TMT (with modifield field lens)	8

3.4	PERFORMANCE COMPARISON OF MOSFIRE on KECK and TMT	8

4	MECHANICAL ISSUES	23

4.1	General Description	23

4.2	Interface Modifications	23

4.3	Impact of modifying pupil location	26

4.4	Constraints affecting field lens position	27

4.5	Recommended Solution:  Displaced CSU	29

5	Recommendations	30

Figures

Figure 3-1	Schematic of sideview of MOSFIRE (KECK) - imaging mode.  Light enters from the left and passes through the slit mask mechanism to the large field lens and then to a fold mirror.	9

Figure 3-2	Top view of MOSFIRE (KECK) after the fold mirror - imaging mode	9

Figure 3-3	Foot-prints associated with beams from a total of 9 field points at (a) one of the field lens surfaces, and (b) 36x36mm detector, assuming a 4.8x4.8’ FOV and a spectral range from λ: 0.97 to 2.40μm	10

Figure 3-4	Spot diagrams associated with KECK + MOSFIRE for 4.8x4.8’ FOV - imaging mode	10

Figure 3-5	Spot diagrams associated with KECK + MOSFIRE for 4.8x1.8’ FOV - spectral mode	11

Figure 3-6	Foot-prints associated with beams from a total of 9 field points at reference surface 328mm in front of fold mirror/grating location for (a) imaging mode with 4.8x4.8’ FOV (λ: 0.97 to 2.40μm) and (b) spectral mode with 4.8x1.8’ FOV (λ: 2.00 to 2.35μm)	11

Figure 3-7	Spot diagrams associated with pupil imagery of KECK + MOSFIRE for 4.8x4.8’ FOV - imaging mode	12

Figure 3-8	Distortion mapping associated with pupil imagery of KECK + MOSFIRE for 4.8x4.8’ FOV-imaging mode	12

Figure 3-9	Top view of MOSFIRE (after fold mirror) + beam trajectories following TMT and NFIRAOS combination - imaging mode	13

Figure 3-10	Foot-prints associated with beams from a total of 9 field points with reference surface half way between last element of collimator & filter (a) imaging mode with 2.0x2.0’ FOV (λ: 0.97 to 2.40μm) & (b) spectral mode with 2.0x0.6’ FOV (λ: 2.00 to 2.35μm)	13

Figure 3-11	Spot diagrams associated with TMT + NFIRAOS + MOSFIRE for 2’ circular FOV - imaging mode	14

Figure 3-12	Spot diagrams associated with TMT + NFIRAOS + MOSFIRE for 2.0x0.6’ elliptical FOV-spectral mode	14

Figure 3-13	Spot diagrams associated with pupil imagery of TMT + NFIRAOS + MOSFIRE for 2’ circular  FOV - imaging mode	15

Figure 3-14	Distortion mapping associated with pupil imagery of TMT + NFIRAOS + MOSFIRE for 2’ circular FOV - imaging mode	15

Figure 3-15	Top view of MOSFIRE with modified field lens + beam trajectories - imaging mode	16

Figure 3-16	Spot diagrams associated with TMT + NFIRAOS + MOSFIRE (modified field lens) for 2.0’ circular FOV - imaging mode	16

Figure 3-17	Spot diagrams associated with TMT + NFIRAOS + MOSFIRE (modified field lens) for 2.0x0.6’ elliptical FOV - spectral mode	17

Figure 3-18	Foot-prints associated with beams from a total of 9 field points with reference surface half way between last element of collimator & filter (a) imaging mode with 2.0x2.0’ FOV (λ: 0.97 to 2.40μm) & (b) spectral mode with 2.0x0.6’ FOV (λ: 2.00 to 2.35μm)	17

Figure 3-19	Spot diagrams associated with pupil imagery of TMT + NFIRAOS + MOSFIRE (modified field lens) for 2’ circular FOV - imaging mode	18

Figure 3-20	Distortion mapping associated with pupil imagery of TMT + NFIRAOS + MOSFIRE (modified field lens) for 2’ circular FOV - imaging mode	18

Figure 4-1	MOSFIRE/IRMS with Generic NFIRAOS Instrument Envelope Superimposed (in red)	24

Figure 4-2	MOSFIRE/IRMS Situated on Lateral Instrument Port of NFIRAOS	25

Figure 4-3	Section View showing Science Beam and MOSFIRE/IRMS Input Optics	26

Figure 4-4	View of MOSFIRE’s Optical System Components (from MOSFIRE PDR Report)	27

Figure 4-5	Section View showing Field Lens (A), Internal Structure Stiffening Tube (B), Collimator Lenses (C), and Internal Bulkhead A (D)	28

Figure 4-6	Section through Internal Structure showing Filter Wheel (grey) Clearance to Stiffening Tube (purple)	28

Tables

Table 3-1	Lens prescription of KECK telescope and MOSFIRE combination	19

Table 4-1	Field Lens Position Approximate Space Envelope for Alternative Re-engineering Cases.  Parenthetical Values Indicate Combination of Each Case with the Modified Cell	29

# Introduction

In December 2006, SAC recommended that a clone of the Keck MOSFIRE instrument (behind NFIRAOS) should become one of the first light instruments. This is an initial examination of how feasible it is to interface MOSFIRE and NFIRAOS from both optical and mechanical perspectives.

## Purpose

Understand the mechanical and optical details of the Keck MOSFIRE instrument and identify possible TMT interface issues. Recommend any possible new requirements, if any, that NFIRAOS must meet in order to make this possible.

## Scope

As necessary, concepts and requirements flow down from the TMT Science Requirements Document (SRD).

As necessary, new requirements implied by the current document flow down into:

- TMT Observatory Requirements Document (ORD)

- TMT Observatory Architecture Document (OAD)

## Applicable Documents

Applicable documents are containing information that shall be applied in the current document. Examples are higher level requirements documents, standards, rules and regulations.

| AD01 | Science-based Requirements Document v15 | TMT.PSC.DRD.05.001.REL15 |
| --- | --- | --- |
| AD02 |  |  |

## Reference Documents

Reference documents are containing information complementing, explaining, detailing, or otherwise supporting the information included in the current document.

| RD01 | Keck MOSFIRE PDR |  |
| --- | --- | --- |
| RD02 |  |  |
| RD03 |  |  |
| RD04 |  |  |

## Change Record

| Version | Date | Remarks |
| --- | --- | --- |
| Draft 0 | 16 Jan 2007 | Internal working draft |
| Draft 1 | 20 Jan 2007 | First internal release for comment |
| Draft 2 | 30 Jan 2007 | Second internal release for comment |
| Draft 2d | 16 April 2007 | Mechanical Interface sections updated based on PDR MOSFIRE design |

## Abbreviations and Acronyms

AIV	Assembly, Integration, and Verification
AO	Adaptive Optics
ELT	Extremely Large Telescope
ESO	European Southern Observatory
IRMOS	Near IR Multi-Object Spectrograph

IRMS	InfraRed Multi-Slit Spectrograph
ISM	International Staff Member
ISO	International Standards Organization
IT	Information Technology
LRU	Line Replaceable Unit
LSM	Local Staff Member
LTQ	Long-Term Queue
LTS	Long-Term Schedule
MDT	Mean Down-Time
MIS	Management Information System
MTBF	Mean Time Between Failures
NFIRAOS	Narrow Field Infrared Adaptive Optics System
NRC	National Research Council of Canada
SW	Software
TBC	To Be Confirmed
TMT	Thirty Meter Telescope
TBD	To Be Determined
UPS	Uninterruptible Power Supply
VLT	Very Large Telescope
VO	Virtual Observatory

## Glossary

TO BE COMPLETED

## Verb Convention

‘Shall’ is used whenever a statement expresses a convention that is binding.  The verbs ‘should’ and ‘may’ express non-mandatory provisions.  ‘Will’ is used to express a declaration of purpose on the part of the design activity.

## Acknowledgements

Chuck Steidel, Bob Weber

# IRMS introduction and top-level requirements

The TMT SAC has consistently ranked some form of multi-object NIR spectroscopy (“IRMOS”) as one of their top priorities for first light. However, it became apparent from the IRMOS feasibility studies that a MOAO based deployable IFU system would be too risky and too expensive and so SAC recommended that a clone of the much simpler Keck MOSFIRE instrument be implemented behind NFIRAOS instead. This instrument would be named “IRMS” – InfraRed Multi-Slit imaging spectrograph.

On Keck, the MOSFIRE instrument is presented with a curved focal plane, approximately 2100mm concave radius, together with a pupil located near the secondary mirror roughly 20m from the focal plane.  On the contrary, NFIRAOS delivers a flat focal plane with a pupil around 0.5Km away.  In order to minimize background, IR instruments must image the pupil on a cold Lyot stop and so, given the above, obvious questions are how can the exit pupil of NFIRAOS be imaged onto a Lyot stop and what will that, and the different focal plane geometry, do to the IRMS image quality? A strong requirement is to minimize modifications to the MOSFIRE design in order to minimize cost, risk and schedule of IRMS.

As part of the interface to NFIRAOS, IRMS will have to provide (ideally) 3 NGS wavefront sensors that can access significant portions of the 2’ field. Ideally these should be identical to similar systems incorporated into IRIS (but their mechanical precision need not be as accurate since the image quality is an order of magnitude lower on average over the wide 2’ field).

# Optical DESCRIPTIONs

## MOSFIRE on KECK

MOSFIRE was specifically designed to work with one of the KECK telescopes.  The lens prescription of the complete optical system is listed in table 1.  Depending on whether a fold mirror or diffraction grating is employed imaging can be either ordinary polychromatic 2-dimensional imagery or imagery of a line object as a function of wavelength - spectral mode operation.  The instrument comprises collimator, followed by fold mirror/grating, camera objective and a 2 dimensional Hawaii2-RG detector with 18mu pixels.  Both collimator and camera objective have solely refractive elements.  The collimator has a focal length of around 1800mm and an effective aperture of F/15 to match the output from KECK, while the camera objective has a focal length of 250mm and an aperture of F/2.  The combination is such that there is a 7.2x reduction from the focal plane of the KECK telescope to the detector.  A FOV of 6.14 x 6.14’ in object space is mapped onto an area 36 x 36mm at the detector.  This is the size of field for ordinary imagery mode, but is reduced in one dimension in order to accommodate the spectral spread.  Spectral coverage is from roughly 0.9 to 2.5μm, although not simultaneously.  A choice of filter can be located between collimator and fold mirror/grating together with a Lyot stop for reducing unwanted radiation passing through to the detector.

Figure 3-1 illustrates a schematic of the complete optical package of MOSFIRE located behind KECK’s focal plane.  Figure 3-2 illustrates an expanded top view of said package minus the large field lens.   Ray-tracing the arrangement indicated that the field lens was in fact undersize to handle the full 6.14’ square FOV, restricting the FOV to a diameter of 6.8’ or an inscribed square of a 4.8’ on a side.  This is evident from Figure 3-3 which depicts foot-prints at one of the field lens surfaces of beams associated with the 9 field points in the spot diagrams as given in Figure 3-4.  Ray distributions for each of these field points are 24 (radial) x 24 (tangential) for the central and extreme wavelengths.  Performance of the KECK + MOSFIRE combination for the imaging mode can be appreciated from the spot diagrams depicted in Figure 3-4 which are for a 4.8’ square FOV and the full spectral range from λ equals 0.97 to 2.40μm.  Performance plots associated with its spectral imaging counterpart are depicted in Figure 3-5 for a 4.8 x 1.8’ FOV and reduced spectral range from λ equals 2.00 to 2.35μm.  It can be clearly seen from Figure 3-2 that the image of the pupil (primary mirror of KECK) is located close to the reference surface located 328mm in front of the fold mirror.  Foot-prints at this surface of beams associated with the same 9 field points are depicted in Figure 3-6a and Figure 3-6b respectively for the imaging and spectral mode of operation.   Spot diagrams associated with the imagery of primary mirror onto Lyot stop are depicted in Figure 3-7.  Mapping for this imagery is graphically represented in Figure 3-8, indicating 0.029% rms and 0.074% maximum distortion over the full circular aperture.

## MOSFIRE on TMT (with no modifications)

Figure 3-9 illustrates the MOSFIRE package introduced behind the TMT + NFIRAOS combination together with associated beam trajectories for a 2.0’ circular FOV.  The image of the pupil (primary mirror of the TMT) is seen to lie half way between the last optical element within the collimator lens group and filter.  Foot-prints at the location of beams associated with the same 9 field points are depicted in Figure 3-10a and Figure 3-10b respectively for the imaging and spectral mode of operation.  For comparison purposes similar analyses were undertaken on TMT configuration as were carried out on the KECK system.  Assuming identical spectral content as above, spot diagrams for the imaging and spectral mode of operations are depicted in Figure 3-11 and Figure 3-12; a 2.0’ circular FOV for the former, while a 2.0 x 0.6’ elliptical FOV for the latter.  Assuming a Lyot stop can be placed at the location mentioned above, the similarity in performance between MOSFIRE on TMT with that on KECK would suggest that this course of action would be the most convenient for incorporating MOSFIRE into TMT.   Spot diagrams associated with the imagery of primary mirror onto Lyot stop are depicted in Figure 3-13.  Mapping  for this imagery is graphically represented in Figure 3-14, indicating 0.120% rms and 0.228% maximum distortion over the full circular aperture. Unfortunately, as discussed below, it isn’t feasible to relocate the Lyot stop to the desired location and consequentially we next explore the possibility of modifying the field lens.

## MOSFIRE on TMT (with modifield field lens)

Decreasing the power of the field lens to force the pupil image to coincide with the same location as found with the KECK telescope and optimizing its shape and location to yield maximum performance resulted conveniently in a plano-convex element of the same size as the original but with the spherical surface having a radius of curvature of 518.0mm.  This modification is also accompanied by a longitudinal displacement of the entire MOSFIRE assembly of 20mm from the NFIRAOS focal plane.   A schematic of this modified arrangement, minus field lens, is illustrated in Figure 3-15 with its associated spot diagrams depicted in Figure 3-16 (imaging mode) and Figure 3-17 (spectral mode).  Foot-prints at the surface in question of beams associated with the same 9 field points are depicted in Figure 3-18a and Figure 3-18b respectively for the imaging and spectral modes of operation.  If the Lyot stop has to be located at its original location, simply modifying the field lens as described above would represent the next favourable possible course of action for incorporating MOSFIRE into TMT.   Spot diagrams associated with the imagery of primary mirror onto Lyot stop are depicted in Figure 3-19.  Mapping for this imagery is graphically represented in Figure 3-20, indicating 0.123% rms and 0.256% maximum distortion over the full circular aperture.

## PERFORMANCE COMPARISON OF MOSFIRE on KECK and TMT

It would appear from the various spot diagrams for the three systems that performance is very much the same for both regular and pupil imagery.  Rms spot sizes for the central and extreme field points for regular imaging are respectively 0.006 and 0.014mm for KECK, 0.005 and 0.013mm for TMT (original field lens) and  0.005 and 0.015mm for TMT (modified field lens).  For the spectral imaging counterpart, these values are as follows:  0.011 and 0.030mm, 0.011 and 0.025mm and 0.011 and 0.029mm.  These calculations are based on a spectral weighting of unity for the central and extreme wavelengths and for the spectral mode, dispersion has been ignored.  Pupil imagery associated with the three arrangements is also similar as clearly seen from the spot diagrams which vary from 1.1 to 1.5mm rms over the entire area of the Lyot aperture.  Pupil distortion is seen to be more prevalent in the TMT configurations, mainly because of the pair of off-axis parabolic mirrors, the relay optics for the NFIRAOS system.  Finally because of pupil mismatch, a small amount of vignetting can be expected with the non-modified TMT system.  Exactly how much would obviously depend on the clear apertures of all the optical components within the system.  Certainly a small increase in the size of the mirror/grating would help reduce this problem. In any case, it appears that the field lens will have to be modified and so this is not an issue.

Figure 3-1	Schematic of side view of MOSFIRE (KECK) - imaging mode.  Light enters from the left and passes through the slit mask mechanism to the large field lens and then to a fold mirror.

Figure 3-2	Top view of MOSFIRE (KECK) after the fold mirror - imaging mode

Figure 3-3	Foot-prints associated with beams from a total of 9 field points at (a) one of the field lens surfaces, and (b) 36x36mm detector, assuming a 4.8x4.8’ FOV and a spectral range from λ: 0.97 to 2.40μm

Figure 3-4	Spot diagrams associated with KECK + MOSFIRE for 4.8x4.8’ FOV - imaging mode

Figure 3-5	Spot diagrams associated with KECK + MOSFIRE for 4.8x1.8’ FOV - spectral mode

(a)		  			                             (b)

Figure 3-6	Foot-prints associated with beams from a total of 9 field points at reference surface 328mm in front of fold mirror/grating location for (a) imaging mode with 4.8x4.8’ FOV (λ: 0.97 to 2.40μm) and (b) spectral mode with 4.8x1.8’ FOV (λ: 2.00 to 2.35μm)

Figure 3-7	Spot diagrams associated with pupil imagery of KECK + MOSFIRE for 4.8x4.8’ FOV - imaging mode

Figure 3-8	Distortion mapping associated with pupil imagery of KECK + MOSFIRE for 4.8x4.8’ FOV-imaging mode

Figure 3-9	Top view of MOSFIRE (after fold mirror) + beam trajectories following TMT and NFIRAOS combination - imaging mode

(a)	(b)

Figure 3-10	Foot-prints associated with beams from a total of 9 field points with reference surface half way between last element of collimator & filter (a) imaging mode with 2.0x2.0’ FOV (λ: 0.97 to 2.40μm) & (b) spectral mode with 2.0x0.6’ FOV (λ: 2.00 to 2.35μm)

Figure 3-11	Spot diagrams associated with TMT + NFIRAOS + MOSFIRE for 2’ circular FOV - imaging mode

Figure 3-12	Spot diagrams associated with TMT + NFIRAOS + MOSFIRE for 2.0x0.6’ elliptical FOV-spectral mode

Figure 3-13	Spot diagrams associated with pupil imagery of TMT + NFIRAOS + MOSFIRE for 2’ circular 
FOV - imaging mode

Figure 3-14	Distortion mapping associated with pupil imagery of TMT + NFIRAOS + MOSFIRE for 2’ circular FOV - imaging mode

Figure 3-15	Top view of MOSFIRE with modified field lens + beam trajectories - imaging mode

Figure 3-16	Spot diagrams associated with TMT + NFIRAOS + MOSFIRE (modified field lens) for 2.0’ circular FOV - imaging mode

Figure 3-17	Spot diagrams associated with TMT + NFIRAOS + MOSFIRE (modified field lens) for 2.0x0.6’ elliptical FOV - spectral mode

(a)	(b)

Figure 3-18	Foot-prints associated with beams from a total of 9 field points with reference surface half way between last element of collimator & filter (a) imaging mode with 2.0x2.0’ FOV (λ: 0.97 to 2.40μm) & (b) spectral mode with 2.0x0.6’ FOV (λ: 2.00 to 2.35μm)

Figure 3-19	Spot diagrams associated with pupil imagery of TMT + NFIRAOS + MOSFIRE (modified field lens) for 2’ circular FOV - imaging mode

Figure 3-20	Distortion mapping associated with pupil imagery of TMT + NFIRAOS + MOSFIRE (modified field lens) for 2’ circular FOV - imaging mode

Table 3-1	Lens prescription of KECK telescope and MOSFIRE combination

OPTICAL SYSTEM DESIGN/ANALYSIS PROGRAM  -  OPTICSOFT-II (v0701)

keck + mosfire imaging                           keck_mosfire_imaging

SYSTEM CONSTRUCTIONAL PARAMETERS:

SURF    RADIUS     DISTANCE   GLASS   SEMI-DIAM  COATING  SPECIAL ATTRIBUTES

0 ************  1.0000E+10  AIR      *********    0

1 ************      0.0000  AIR       5156.610    0

2  -34973.9961 -15394.9834 -AIR       5156.610    0  STO,FIG,CAO,

3   -4737.9160  15394.9834  AIR        634.229    0  FIG,

4 ************   1904.8470  AIR        215.552    0

5 ************     34.9250  FQTZ_I     163.898    0

6 ************    228.6000  AIR        163.245    0

7 ************     12.7000  V76FQTZ    157.046    0

8 ************    333.3750  AIR        156.808    0

9   -2123.5732    580.0000  AIR        147.936    0

10 ************      0.0000  AIR        172.441    0

11     547.5947     47.0000  V76CAF2    173.624    0

12   -4438.9512    489.2183  AIR        172.553    0

13 ************      0.0000  AIR      *********    0  TIL,HID,

14 ************      0.0000 -AIR        116.529    0

15 ************   -441.7242 -AIR      *********    0  TIL,HID,

16     291.6063     -8.0000 -V76CAF2     64.846    0

17 ************     -9.4373 -AIR         64.659    0

18    -150.9766     -8.0000 -V76ZNSE     64.246    0

19    -137.6240    -40.1156 -AIR         61.929    0

20      92.6744     -8.0000 -V76FQTZ     61.975    0

21     130.6669     -8.8248 -AIR         65.738    0

22     128.3060    -20.0000 -V76BAF2     66.860    0

23     109.8851   -123.3233 -AIR         70.583    0

24    2500.0000    -35.6164 -V76CAF2     78.650    0

25     239.7622    -90.0000 -AIR         79.757    0

26 ************    -10.0000 -V76FQTZ     71.316    0

27 ************    -80.0000 -AIR         70.857    0

28 ************    -40.0000 -AIR         65.609    0

29 ************   -328.0000 -AIR         66.265    0

30 ************      0.0000 -AIR      *********    0  TIL,HID,

31 ************      0.0000  AIR         93.004    0

32 ************    300.0000  AIR      *********    0  TIL,HID,

33     435.5621     31.0000  V76CAF2    135.000    0

34    1812.7576      4.0000  AIR        118.538    0

35     302.6267     75.0000  V76CAF2    118.221    0

36    -319.5752     13.2101  AIR        114.207    0

37    -276.5860      7.0000  V76FTM1    109.380    0

38     576.7587      4.0000  AIR        106.869    0

39     210.7587     38.0000  V76FPL51   107.062    0

40     454.7188    142.4862  AIR        103.292    0

41     135.3481     48.0000  V76BAF2     78.717    0

42    6620.2021     24.3616  AIR         71.264    0

43     108.5229     50.0000  V76BAF2     52.559    0

44     221.4500     11.0000  AIR         32.540    0

45    -175.2402      4.0000  V76FTM1     27.763    0

46     134.6514     13.4401  AIR         25.000    0

47 ************      0.0000  AIR         25.000    0  CAO,

48 ************      0.0000  AIR         25.000    0

MEDIUM #    GLASS                  REFRACTIVE INDEX                V-VALUE

1750.0nm      968.0nm     2400.0nm

0       AIR            1.000000     1.000000     1.000000    *********

1       AIR            1.000000     1.000000     1.000000    *********

2      -AIR           -1.000000    -1.000000    -1.000000    *********

3       AIR            1.000000     1.000000     1.000000    *********

4       AIR            1.000000     1.000000     1.000000    *********

5       FQTZ_I         1.441660     1.450924     1.431756       23.042

6       AIR            1.000000     1.000000     1.000000    *********

7       V76FQTZ        1.440754     1.449987     1.430891       23.082

8       AIR            1.000000     1.000000     1.000000    *********

9       AIR            1.000000     1.000000     1.000000    *********

10       AIR            1.000000     1.000000     1.000000    *********

11       V76CAF2        1.427323     1.431343     1.423918       57.555

12       AIR            1.000000     1.000000     1.000000    *********

13       AIR            1.000000     1.000000     1.000000    *********

14      -AIR           -1.000000    -1.000000    -1.000000    *********

15      -AIR           -1.000000    -1.000000    -1.000000    *********

16      -V76CAF2       -1.427323    -1.431343    -1.423918       57.555

17      -AIR           -1.000000    -1.000000    -1.000000    *********

18      -V76ZNSE       -2.438282    -2.479140    -2.429707       29.096

19      -AIR           -1.000000    -1.000000    -1.000000    *********

20      -V76FQTZ       -1.440754    -1.449987    -1.430891       23.082

21      -AIR           -1.000000    -1.000000    -1.000000    *********

22      -V76BAF2       -1.468621    -1.472017    -1.466537       85.510

23      -AIR           -1.000000    -1.000000    -1.000000    *********

24      -V76CAF2       -1.427323    -1.431343    -1.423918       57.555

25      -AIR           -1.000000    -1.000000    -1.000000    *********

26      -V76FQTZ       -1.440754    -1.449987    -1.430891       23.082

27      -AIR           -1.000000    -1.000000    -1.000000    *********

28      -AIR           -1.000000    -1.000000    -1.000000    *********

29      -AIR           -1.000000    -1.000000    -1.000000    *********

30      -AIR           -1.000000    -1.000000    -1.000000    *********

31       AIR            1.000000     1.000000     1.000000    *********

32       AIR            1.000000     1.000000     1.000000    *********

33       V76CAF2        1.427323     1.431343     1.423918       57.555

34       AIR            1.000000     1.000000     1.000000    *********

35       V76CAF2        1.427323     1.431343     1.423918       57.555

36       AIR            1.000000     1.000000     1.000000    *********

37       V76FTM1        1.565632     1.577632     1.555743       25.841

38       AIR            1.000000     1.000000     1.000000    *********

39       V76FPL51       1.486369     1.492460     1.480693       41.332

40       AIR            1.000000     1.000000     1.000000    *********

41       V76BAF2        1.468621     1.472017     1.466537       85.510

42       AIR            1.000000     1.000000     1.000000    *********

43       V76BAF2        1.468621     1.472017     1.466537       85.510

44       AIR            1.000000     1.000000     1.000000    *********

45       V76FTM1        1.565632     1.577632     1.555743       25.841

46       AIR            1.000000     1.000000     1.000000    *********

47       AIR            1.000000     1.000000     1.000000    *********

CLEAR APERTURE AND CENTRAL OBSTRUCTION SURFACE DATA:

OUTER                       INNER                 OFF-SET

SURF     RHO      X        Y        RHO       X        Y         X        Y

2     0.000    0.000    0.000  1289.152    0.000    0.000     0.000    0.000

47    25.000   18.000   18.000     0.000    0.000    0.000     0.000    0.000

FIGURED SURFACE DATA:                                                    NORM

SURF    Q           A1          A2          A3          A4          A5     Y/N

2  -1.0037E+00  0.0000E+00  0.0000E+00  0.0000E+00  0.0000E+00  0.0000E+00 Y

3  -1.6443E+00  0.0000E+00  0.0000E+00  0.0000E+00  0.0000E+00  0.0000E+00 Y

TILT DATA:

SURF     X          Y          Z       ORDER

13    17.5000     0.0000     0.0000     1

15    17.5000     0.0000     0.0000     1

30     0.0000    20.0000     0.0000     1

32     0.0000    20.0000     0.0000     1

# MECHANICAL ISSUES

## General Description

A study of the NFIRAOS to MOSFIRE/IRMS interface has been undertaken, in which it has been assumed that MOSFIRE/IRMS will occupy the lateral instrument port of NFIRAOS. This is proposed because the Keck MOSFIRE has been designed to operate in a variable gravity vector environment and includes a rotator bearing. Since NFIRAOS does not offer field de-rotation nor can support the mass of an instrument on its lateral port, MOSFIRE/IRMS may be a logical choice for the lateral position.

A Generic NFIRAOS Instrument Envelope was defined for CoDR, having dimensions 4.0m length and 2.0m diameter. The mechanical interface was defined to be one end of this cylinder, and NFIRAOS delivers a 750mm back-focal distance relative to this interface plane. In addition, NFIRAOS permits a 250mm long instrument “snout” to protrude into NFIRAOS.

Figure 4-21 illustrates the Generic Instrument Envelope superimposed on MOSFIRE/IRMS. MOSFIRE/IRMS’s focal plane is taken to be the Configurable Slitmask Unit, and this has been situated accordingly within the envelope. The portion of MOSFIRE/IRMS that can be seen to extend outside the generic envelope is its Cassegrain Module Assembly, which incorporates the rotator bearing and cable wraps, and may therefore be retained to serve as the MOSFIRE/IRMS structural interface to the TMT Nasmyth structure. MOSFIRE/IRMS’s Dust Cover Assembly has been omitted, as it would be supplanted by a dynamic lip-seal between MOSFIRE/IRMS and the NFIRAOS enclosure. The Guider and Dust Cover Assemblies have also been suppressed in this view, for reasons explained below.

## Interface Modifications

MOSFIRE/IRMS will be supported by a dedicated intermediate structure which likely will utilize all or part of the existing Cassegrain Module Assembly and which will rest on hard points provided by the TMT structure. Although MOSFIRE/IRMS will abut NFIRAOS, no cantilevered loads will be supported by NFIRAOS.

On-instrument wavefront sensors will need to be provided by MOSFIRE/IRMS. These will likely be located inside its main Dewar, and will supersede the Guider system present in the Keck MOSFIRE design. For this reason, the Guider Assembly has not been depicted in Figure 4-21. Similarly, the Dust Cover Assembly has been omitted, since the optical interface will be protected by NFIRAOS and a simple, manual cover will suffice for transportation and occasional operational usage, if required.

Generic EnvelopeØ 2.0m

Generic Envelope Length 4.0m

MOSFIRE/IRMS

Rotator Bearing + Cable Wrap

Figure 4-21	MOSFIRE/IRMS with Generic NFIRAOS Instrument Envelope Superimposed (in red)

NFIRAOS

MOSFIRE/IRMS

NFIRAOS Length 9.2m

Figure 4-22	MOSFIRE/IRMS Situated on Lateral Instrument Port of NFIRAOS

Figure 4-22 depicts MOSFIRE/IRMS situated at the lateral instrument port of NFIRAOS. Note that NFIRAOS is shown “inverted” relative to its configuration presented at CoDR. This configuration has been proposed to lessen the wind cross section of NFIRAOS to alleviate telescope flushing concerns.

The focal plane delivered by NFIRAOS at CoDR was 264mm in diameter. The configurable slitmask unit of MOSFIRE, however, is 262mm square. The implications of a recent amendment to the TMT prescription will need to be evaluated to ascertain whether the amended focal plane diameter will fit within the slitmask window. A section view of this portion of MOSFIRE/IRMS is presented in Figure 4-23, which also shows the expanding beam after focus incident upon the field lens. The input beamprint is seen to be accommodated by the existing aperture of the field lens.

Focal Plane

(CSU)

Field Lens

Figure 4-23	Section View showing Science Beam and MOSFIRE/IRMS Input Optics

## Impact of modifying pupil location

The pupil location predicted by the TMT+NFIRAOS+MOSFIRE/IRMS prescription is displaced from its location in the original MOSFIRE design (Keck+MOSFIRE), and falls within the space between the final collimator lens and the filter mechanism. To re-optimize image quality, the pupil stop (Lyot stop) would need to be placed at the new pupil location. The available space is approximately 38mm long, and by inspection of Figure 4-24 it appears unlikely that the Lyot stop mechanism can be accommodated there. The large size of the filter wheel mechanism precludes its position being altered to increase the distance from the collimator lens. Furthermore, vignetting has been noticed at the grating-mirror surface in the optical ray trace diagrams of this configuration. Significant re-engineering, including flexural analysis, would be required to alleviate these shortcomings.

A preferable approach may be to substitute a field lens having a prescription designed to deliver the pupil to the existing Lyot stop location. The axial position of the field lens is constrained by its distance from the adjacent folded beam, however some freedom may exist in its placement to optimize optical performance.

Figure 4-24	View of MOSFIRE’s Optical System Components (from MOSFIRE PDR Report)

## Constraints affecting field lens position

The internal structure of MOSFIRE/IRMS consists of three bulkheads which are arranged axially by spacer tubes. The optical path entering MOSFIRE/IRMS passes through central holes in the largest two of these bulkheads, these holes being connected by a internal structure stiffening tube.

The field lens cell is mounted into the stiffening tube, as shown in Figure 4-25. In this view, the focal plane is near the left coinciding with the configurable slitmask unit (in grey). It can be seen that there is no opportunity to displace a replacement field lens to the right of its current position since it is very tightly constrained by the collimator lens assembly.

A field lens optimized for TMT could be positioned slightly to the left of its current position within the space envelope of the existing cell. There are, however, a number of mechanical and structural parts which would need to be altered to accommodate. A slight further gain in the space envelope available in this direction might be available by re-designing the internal structure stiffening tube. The constraints affecting this option are the existence of a lateral reinforcement forming part of the bulkhead weldment (in light green). If re-engineering of the bulkhead weldment is considered to be within scope, a further case can be envisaged where the ultimate constraint on field lens position would be the filter wheel mechanism enclosure. The current close proximity between these parts can be appreciated in Figure 4-26.

Table 4-2 summarizes the approximate range of field lens positions for the re-engineering cases described. Dimensions in this table are given relative to the current lens position.

D

A

C

B

Figure 4-25	Section View showing Field Lens (A), Internal Structure Stiffening Tube (B), Collimator Lenses (C), and Internal Bulkhead A (D)

Figure 4-26	Section through Internal Structure showing Filter Wheel (grey) Clearance to Stiffening Tube (purple)

| Design Case | Relative Axial Position Envelope [mm] |  |
| --- | --- | --- |
|  | Toward CSU | Toward Mirror |
| Modify Cell | 18 | 0 |
| Modify Stiffening Tube | 25 (43) | 0 |
| Re-design Stiffening Tube | 47 (65) | 0 |
| Re-design Bulkhead weldment and stiffening tube. | 155 (173) | 0 |

Table 4-2	Field Lens Position Approximate Space Envelope for Alternative Re-engineering Cases.  Parenthetical Values indicate combination of each case with the Modified Cell

## Recommended Solution:  Displaced CSU

To instate the required optical path length between the TMT-optimized Field Lens and the focal plane, it has been proposed to displace the focal plane, and hence the Configurable Slitmask Unit (CSU), ~20mm closer to the entrance window of MOSFIRE/IRMS. This will have the effect of positioning MOSFIRE/IRMS further from NFIRAOS by the same amount, when the science focus is superimposed onto the CSU plane.

The mechanical design of this proposal is likely to be much simpler in implementation than re-positioning of the Field Lens described in paragraph 4.4.

The CSU is supported by four interface columns, which are attached to a main structural feature called Internal Bulkhead A. To reposition the CSU, these interface columns should be lengthened by the required amount, approximately 20mm. There is a risk that this will cause the CSU to interfere with the Inner Window Rear Baffle. The baffle may need to be re-designed to remain compatible. The Inner Window Assembly which incorporates the baffle is supported by four support columns which are also attached to Internal Bulkhead A near the CSU interface columns. These support columns may also need to be adapted.

The Guider system will likely be replaced by on-instrument wavefront sensors which, in order to be effective at the desired wavelengths, will need to be cooled. It will also be desired that these wavefront sensors will patrol the field of view with minimal potential vignetting of candidate science objects. For these reasons, it is expected that the wavefront sensors will occupy space within the front portion of the MOSFIRE/IRMS Dewar. Although there is presently some unallocated volume in this region of the instrument, the wavefront sensors will likely require the creation of additional space.

Additional space could be created by extending the Dewar toward NFIRAOS probably though the use of a spacer ring. Figure 4-21 suggests this is possible while remaining within the volume of the Generic Instrument Envelope. Increasing the MOSFIRE/IRMS internal volume by this approach also provides the opportunity to address the details of repositioning the CSU and possibly the Inner Window Assembly in a coherent and optimal design.

# Recommendations

The above optical and mechanical analyses indicate that adapting MOSFIRE for use with TMT/NFIRAOS should be relatively straightforward. The recommended solution involves exchanging the field lens and moving the configurable slit unit 20mm further away from the field lens. This implies redesign of the front end of MOSFIRE but this portion of the instrument will have to be modified in any case in order to interface with NFIRAOS and also to incorporate TMT wavefront sensors.