## Slide 1: ZTF Image Quality (etc)
Roger Smith
Caltech

## Slide 2: Delivered Image Quality
Static shims
Open loop map (boustrophedonic vs. random sampled)
Closed loop correction
Example of good
Example of bad
Ambiguity between tilt and translation
Current work
  Direct measurement of flexure primary, ring girder instrument
  Faster TTF sampling rate

## Slide 3: Focus algorithm validation (R. Dekany)
We validated algorithms using artificial data to test immunity to various biases. The wavefront propagation model includes the effects of diffraction and optical aberations.
Inner CCD
Corner CCD (~4 deg off-axis)
Our test case
A ‘step-staircase’ where every CCD was pistoned in 25 micron height increments. We also simulated the effect of not having a trim plate during our May 2017 engineering run.
Conclusions
Calibrated: Vignetting of the pupil due to beam walk off of primary mirror biases sextractor image width determination (critical to our focus algorithms accuracy.)
Confirmed: Presence of spherical aberration during our engineering run did not bias our image width results, except in the presence of primary mirror vignetting.
Confirmed: Effect of moving spider and hub shadow does not bias image width determination.
~1.8 mm defocus (w/ 7 λ spherical aberration)
~0.4 mm defocus
(no spherical aberration)
Thanks to Aaron Roodman, LSST, for extra-focal imaging help.

## Slide 4: Image quality, can be sub-Nyquist !
Median FWHM per quadrant (arcsec)

## Slide 5: Tilt and focus variability
To be removed by auto-focus…. under development
Airmass
Black=min-max tilt
Red=median
Same field

## Slide 6: Static map of flexure induced focus and tilt
After larges we must use static map of flexure induced focus and tilt.
Focus CCDs continuously refine map and apply feedback
Need several steps to beat down noise  lag

## Slide 7: Why is focus tilt so “noisy” ?
Many improvements in algorithm … now the measurement is stable for constant pointing.
There appears to be something moving when we slew:
  Mirror position in cell has been measured
  More instrumentation to break degeneracy between X-Y motion and tip-tilt
    Position of ring girder
    Instrument with respect to ring girder
Tilt and XY shift are degenerate.
Pointing errors (telescope w.r.t. sky) mask instrument motion w.r.t. telescope.

## Slide 8: Pointing issues  masking flexure?
Polar axis mis-alignment (or error in pointing model)
Also there is Dec encoder drift.

## Slide 9: Scattered light
Extensive new baffles
Recent cleaning of corrector; monitor via wings of profiles
Scattering by field flattener frames
Dust
Ghosts

## Slide 10: Ghost Mitigation
Trim plate could have been made a meniscus but wasn’t 
Trim plate and corrector could have been tilted, but weren’t 
Invert corrector doublet, maybe ? – a meniscus !
 Software is the only remedy without down time.

## Slide 11: New baffles
Extensive baffles on tube interior.
No single bounce to primary.

## Slide 12: New baffles
Extensive baffles on tube interior.
No single bounce to primary.
Closely spaced to block scattering onto corrector

## Slide 13: Paint, flocking
Surfaces subject to wear and dust accumulation:
  Avian-D black paint
  <3% reflectivity,
  Much better than average at low incidence
  good wear resistance
Protected surfaces (facing primary)
  “Protostar” Flocking
  <1% reflectivity
  8x better even than Avian D at low incidence
  Robust. Negligible shedding of lint.
Black Annodize
Avian-D

## Slide 14: Contamination monitoring
Track cleanliness of corrector:
Fraction of light in wings of bright stars successfully used by Thomas Kupfer to trace cryopumping.
Should work for corrector cleanliness too. Angie Van Sistine is repeating after first cleaning of trim plate.
Dust spot tracking:
useful to understand flat field stability and dewar cleanliness.

## Slide 15: Flat fields
Illumination/vignetting pattern
Shutter timing error map
Filter leaks
Illuminator
  How we acquire flats
  Post factor color balance
  Effects of screen non-uniformity
  Requirement on stability of intensity distribution
Star-flats
  dominated by scattering from field flattener frames
  Compare to DES
CCD sensitivity dependence on wavelength
  mid scales
  PRNU
  fringing

## Slide 16: Raw R flat
Dust
on field flattener
How stable?
not real?
Scattering off frames holding field flatteners
Vignetting: beam walks of primary
Rings due to CCD QE (thickness?)
Quadrant gains adjusted for continuity across boundaries by Nadia Blagorodnova

## Slide 17
The bright initial line has been fixed … check this!

## Slide 18
Scattering off frames is an increase in background.
It is not a sensitivity enhancement so must be removed from flats (about 6% correction)

## Slide 19: Star Flats
From report by Matteo Giomi
Star flats have been fat fielded and should not show…
CCD QE structure
CCD channel gains
Scattering off field flatteners
Note that large scale illumination structure has been removed quite well by dome flats

## Slide 20: DES star-flats …for comparison
From Robert Lupton

## Slide 21: Shutter timing error
Shutter open

## Slide 22: Shutter timing error
Shadows of blades cross obstructions at different times for different field points

## Slide 23: Timing error map
Seconds
Dec (pixels)
HA (pixels)
Only 0.2% error
but easily corrected.
Not present in dome flats since LED is pulsed when shutter is fully open.
Single exposure:
Multi-exposure:
Effective exposure time error:
Multiply signal by:

## Slide 24: Flat field illuminator
Telescope docks to screen mounted on tower, not dome.
Makes light seal. Baffled to reduce scattering from corrector.

## Slide 26: Flat Field Illuminator boards
Ring of 24 LEDs per color: individually addressable on
8 boards in ring; daisy chained; I2C serial bus
15 color options: 3 per filter passband plus some out-of-band.
Independent intensity control, per color, per board.
Read back of board temperature and LED forward voltage to allow correction for efficiency changes with LED junction temperature.
Independent photometer with calibrated photodiode to record long term intensity changes.
Teensy
Controller
LED switches
15 wavelengths, 3 LEDs each
Voltage & Temp RO
Programmable Current Source

## Slide 27: Beam footprint
Lies within shadowing at edges
…check this !

## Slide 28: Intensity Uniformity
Uniformity measured with RGB CMOS camera,
calibrate by moving closer to screen
Mask off pattern beyond FoV of telescope
For each color channel, plot:
  Image/mean(Image)-1
RGB analysis is:
  (Image_R + Image_G + Image_B) / mean(Image_R + Image_G + Image_B)
  - 1

## Slide 29: Line Cuts

## Slide 30: LED Wavelengths
300
400
500
600
700
800
900
1000
Wavelength (nm)
100
80
60
40
0
20
G-Band
R-Band
I-Band
% Intensity

## Slide 31: What spectral weighting?
For each filter{
Select Filter
  For each wavelength in passband {
  Open shutter
  Pulse LED
  Close shutter
  Read out }
  Sum flats for passband. }
LED current is set for similar signal in each color.
Currently uniform weighting is used but this is clearly not optimal as shown at left.

## Slide 32: Alternate LED weighting
Example only.
Probably still not color balanced.

## Slide 33: Flat field flip-book – I band example
Need to a set of full-mosiac reference flats
All wavelengths for all filters gives us:
  Flat field variation with wavelength for CCDs
  Out-of-band rejection by filters
  Leaks around edges of filters, and at focus CCDs.
  Bandpass uniformity

## Slide 34: Example out-of-band flat (R filter)
e.g. Before frames redesigned…. Much better now

## Slide 35: PRNU – low to mid frequencies suppressed
The following images are from WaSP
Each image has been divided by a box car filtered version of itself (~20x20 pixels) to leave only high patial frequencies.

## Slide 36: Flats vs wavelength
FLAT
Laser annealing pattern fro Boron implant at backside
AR coating defect or surface contamination
from WaSP

## Slide 37
FLAT

## Slide 38
FLAT

## Slide 39
FLAT

## Slide 40
FLAT

## Slide 41
FLAT

## Slide 45
FLAT

## Slide 46
FLAT

## Slide 47
FLAT

## Slide 48
FLAT

## Slide 49: Non-uniformity vs wavelength
Non-uniformity (%)
Wavelength (nm)
Stddev/mean in large region without dust spots,
after dividing by 200x200 box car to remove mild illumination gradient
Probably pixel area, not sensitivity variation
Laser annealing pattern,
for Boron implant at back surface
Sum of 10 flats  ~2E6 e-/pixel  shot noise ~0.07%