## Slide 1: SEDM Enhancement Feb 09 2019
Don Neill
Michael Feeney
Richard Walters
David Hover

## Slide 2: Facility instrument on P60

## Slide 3: Hyperspectral Imaging Spectrograph

## Slide 4: Purpose for Enhancement: Reliability
Cable Management
  Existing design had poor cable organization
  Airmass limit of 1.8 due to intermittent connectivity
Electronics Management
  Modules were attached via zip ties
  Poor serviceability and reliability
Instrument Enclosure
  Poor sealing yields condensation issues
  Foam core paneling shed particulate and is a poor material choice
  Difficult to access instrument (non-captive thumbscrews)
Optical Mount Issues
  Collimator and Camera Barrels have a band-aid mounting solution
  Not all mounts reference the optical bench (datum surface)

## Slide 5: Purpose for Enhancement: Throughput
  Clean optics
  Replace pick-off prism
  Re-align optics

## Slide 6: Enhancement Overview
1. Optical Bench Assembly
2. Enclosure Assembly
3. IO Panel Assembly
4. Instrument Cart

## Slide 7: Optical Bench Assembly Details
Hoist Rings (3X)
Enclosure Latch Brackets (6X)
IFU Camera + Mount
Camera Barrel + Mount
Collimator Barrel + Mount
Shutter Drivers + Enclosure
RC Camera + Mount

## Slide 8: Optical Bench Assembly Details: Camera Barrel + Mount
Clamps (2X)
(Lined with Compliant Rubber)
V-Block
Newport 1 Axis Stage
Camera Barrel
OLD
Band-aid Clamps
Flexure Base
Actuator Pushes on Flexure
NEW

## Slide 9: Optical Bench Assembly Details: Collimator Barrel + Mount
Clamps (2X)
(Lined with Compliant Rubber)
V-Block
Collimator Barrel
OLD
Band-aid Clamps
NEW
Pin from Old Buckle Design
New MLA Cell
(Installed Already)
Limited Clocking Adjustment

## Slide 10: Optical Bench Assembly Details: Rainbow/Guide Camera + Mount
OLD
NEW
Gusset Style
(Makes Enclosure Complicated)
Rainbow/Guide Camera
Camera Mount Plate
Top Standoff Plate
Standoff (4X)
Enclosure Panel
Bottom Standoff Plate + Clamps
4X Filter + Pickoff Assembly
Shutter Mount Already Installed

## Slide 11: Optical Bench Assembly Details: Rainbow/Guide Camera + Mount
Cable Passthroughs (2X)
Weather Stripping for Sealing

## Slide 12: Enclosure Assembly Details
Main Features:
  Quick release clamps
  Aluminum Extrusion Framework
  Opaque Black Acrylic Paneling
  T-Slot rubber weather-stripping for sealing
  Easy/fast removal (no unplugging cables)
  Full access to optical bench

## Slide 13: IO Panel Assembly Details
Newport Drivers
Glycol Supply/Return Manifolds
Dry Air Supply
Camera PSU’s
Shutter Driver PSU
Newport PSU
NPS’s
Camera USB Feedthrough (2X)
Newport Driver Feedthrough
Camera PSU Feedthrough
TP-Link 5 Port Switch
IOLAN Serial-Ethernet

## Slide 14: IO Panel Assembly Details
Standoffs with Sandwich Plates
¼-20 Holes for Cable Management

## Slide 15: IO and Utilities Layout

## Slide 16: Handling Cart Assembly
Compact Gas Springs (4X)
Engagement Holes in Optical Bench (4X)
Aluminum Extrusion Framework
Full Swivel with Break

## Slide 17: Handling Cart Assembly
Manual Forklift (48” Travel)
Manual Forklift Engagement
Forklift Profile Representation

## Slide 18: Enhancement Installation
SEDM removed January 5th, 2019
SEDM+ installed February 8th, 2019
Michael Feeney, engineer & instrument master
Re-plumbed P60:
  Dry air
  Glycol
  Network

## Slide 21: Preliminary Results
Feb 26, 2019 / Nov 7, 2018
Nov 7 2018:
  New Moon
  No Clouds
  Photometric?
Feb 26 2019:
  Last Quarter Moon
  Partial Night: Clouds
  Not Photometric

## Slide 22: Mirror Reflectances
Pickoff Mirror
Fold Mirror

## Slide 23: Mirror Reflectances
Acton #4100 (green)
ordered for Pickoff prism & Fold mirror
March 26th
No downtime

## Slide 24: Illumination Uniformity
OLD: 20181107
NEW: 20190209
Narrower scale

## Slide 25: Illumination Uniformity
Hg master: 30s: -60 – 15,000 DN
OLD: 20181107
NEW: 20190209

## Slide 26: Still going faint: g ~ 19.5

## Slide 27: Still going faint: g ~ 19.5

## Slide 28: Still going faint: g ~20

## Slide 29: Still going faint: g ~ 20

## Slide 30: Future Work
Photometric Night for more precise throughput comparison
Calculate absolute throughput (total and instrumental)
Early tests show we can remove airmass limit (done)
Improve wavelength coverage:
  More expensive dielectric coating for fold mirror and pickoff mirror
More baffling for fold mirror

## Slide 31: SEDM @ ZTF
Nadia’s SEDM statistics (next)
Mickael’s talk on pysedm
Friday workshop on custom extractions with pysedm
Please cite:
  Blagorodnova, Neill, Walters et al. 2018 (SEDM instrument)
  Rigault, Neill, Blagorodnova et al. 2019 (pysedm pipeline)
Acknowledge:
  SED Machine is based upon work supported by the National Science Foundation under Grant No. 1106171

## Slide 32: P60 Status
pharos.caltech.edu/monitor

## Slide 33: Web pages
P60 Status (NEW):
  pharos.caltech.edu/monitor
Documentation:
  www.astro.caltech.edu/sedm
Data access:
  pharos.caltech.edu
  Account required (rsw@astro.caltech.edu)
Twiki:
  http://www.oir.caltech.edu/twiki_ptf/bin/view/ZTF/SEDMOperations